## Results
Hidex Automatic Gamma Counter
Counting 18F
Run
Template | Counting 18F
Run status | Finished
Run started | 2021-09-01 18:09
Run type | Count only, single rack
Software version | 1.8.0.0
AMG serial number | 2210227
Firmware versions | 0: 50, 1: 50, F: 100
Counting time | 60
Result normalization | 2021-09-01 11:48
Comment
Counting windows
Window start | Window end | Additional windows | Isotope | Counting efficiency | Half-life
400 | 650 |  | ¹⁸F | 1 | 1.82951 h
Racks
Rack | Time | Barcode | Slot count | Latest QC check | QC status
1 | 2021-09-01 18:09 | 006 | 10 | 2021-09-01 10:39 | Calibration suggested
Results
Rack | Vial | Time | Counted time (s) | Dead time factor | ¹⁸F (counts) | ¹⁸F (CPM) | Normalized ¹⁸F (CPM) | Normalized ¹⁸F (Bq) | Tags
1 | 1 | 2021-09-01 18:11 | 60 | 1.000 | 46 | 46 | 515 | 9
1 | 2 | 2021-09-01 18:12 | 60 | 1.462 | 1560521 | 2281669 | 25796654 | 429944
1 | 3 | 2021-09-01 18:14 | 60 | 1.457 | 1550410 | 2258323 | 25774145 | 429569
1 | 4 | 2021-09-01 18:15 | 60 | 1.450 | 1533509 | 2223906 | 25621026 | 427017
1 | 5 | 2021-09-01 18:17 | 60 | 1.451 | 1530040 | 2220002 | 25817524 | 430292
1 | 6 | 2021-09-01 18:18 | 60 | 1.156 | 682452 | 788767 | 9258866 | 154314
1 | 7 | 2021-09-01 18:20 | 60 | 1.151 | 659611 | 759152 | 8994688 | 149911
1 | 8 | 2021-09-01 18:21 | 60 | 1.149 | 649680 | 746138 | 8923359 | 148723
1 | 9 | 2021-09-01 18:22 | 60 | 1.146 | 644448 | 738581 | 8916645 | 148611
1 | 10 | 2021-09-01 18:24 | 60 | 1.000 | 56 | 56 | 682 | 11

## Spectra
Hidex Automatic Gamma Counter
Counting 18F
Run
Template | Counting 18F
Run status | Finished
Run started | 2021-09-01 18:09
Run type | Count only, single rack
Software version | 1.8.0.0
AMG serial number | 2210227
Firmware versions | 0: 50, 1: 50, F: 100
Counting time | 60
Result normalization | 2021-09-01 11:48
Comment
Spectra
Channel
Rack | Vial | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66 | 67 | 68 | 69 | 70 | 71 | 72 | 73 | 74 | 75 | 76 | 77 | 78 | 79 | 80 | 81 | 82 | 83 | 84 | 85 | 86 | 87 | 88 | 89 | 90 | 91 | 92 | 93 | 94 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 117 | 118 | 119 | 120 | 121 | 122 | 123 | 124 | 125 | 126 | 127 | 128 | 129 | 130 | 131 | 132 | 133 | 134 | 135 | 136 | 137 | 138 | 139 | 140 | 141 | 142 | 143 | 144 | 145 | 146 | 147 | 148 | 149 | 150 | 151 | 152 | 153 | 154 | 155 | 156 | 157 | 158 | 159 | 160 | 161 | 162 | 163 | 164 | 165 | 166 | 167 | 168 | 169 | 170 | 171 | 172 | 173 | 174 | 175 | 176 | 177 | 178 | 179 | 180 | 181 | 182 | 183 | 184 | 185 | 186 | 187 | 188 | 189 | 190 | 191 | 192 | 193 | 194 | 195 | 196 | 197 | 198 | 199 | 200 | 201 | 202 | 203 | 204 | 205 | 206 | 207 | 208 | 209 | 210 | 211 | 212 | 213 | 214 | 215 | 216 | 217 | 218 | 219 | 220 | 221 | 222 | 223 | 224 | 225 | 226 | 227 | 228 | 229 | 230 | 231 | 232 | 233 | 234 | 235 | 236 | 237 | 238 | 239 | 240 | 241 | 242 | 243 | 244 | 245 | 246 | 247 | 248 | 249 | 250 | 251 | 252 | 253 | 254 | 255 | 256 | 257 | 258 | 259 | 260 | 261 | 262 | 263 | 264 | 265 | 266 | 267 | 268 | 269 | 270 | 271 | 272 | 273 | 274 | 275 | 276 | 277 | 278 | 279 | 280 | 281 | 282 | 283 | 284 | 285 | 286 | 287 | 288 | 289 | 290 | 291 | 292 | 293 | 294 | 295 | 296 | 297 | 298 | 299 | 300 | 301 | 302 | 303 | 304 | 305 | 306 | 307 | 308 | 309 | 310 | 311 | 312 | 313 | 314 | 315 | 316 | 317 | 318 | 319 | 320 | 321 | 322 | 323 | 324 | 325 | 326 | 327 | 328 | 329 | 330 | 331 | 332 | 333 | 334 | 335 | 336 | 337 | 338 | 339 | 340 | 341 | 342 | 343 | 344 | 345 | 346 | 347 | 348 | 349 | 350 | 351 | 352 | 353 | 354 | 355 | 356 | 357 | 358 | 359 | 360 | 361 | 362 | 363 | 364 | 365 | 366 | 367 | 368 | 369 | 370 | 371 | 372 | 373 | 374 | 375 | 376 | 377 | 378 | 379 | 380 | 381 | 382 | 383 | 384 | 385 | 386 | 387 | 388 | 389 | 390 | 391 | 392 | 393 | 394 | 395 | 396 | 397 | 398 | 399 | 400 | 401 | 402 | 403 | 404 | 405 | 406 | 407 | 408 | 409 | 410 | 411 | 412 | 413 | 414 | 415 | 416 | 417 | 418 | 419 | 420 | 421 | 422 | 423 | 424 | 425 | 426 | 427 | 428 | 429 | 430 | 431 | 432 | 433 | 434 | 435 | 436 | 437 | 438 | 439 | 440 | 441 | 442 | 443 | 444 | 445 | 446 | 447 | 448 | 449 | 450 | 451 | 452 | 453 | 454 | 455 | 456 | 457 | 458 | 459 | 460 | 461 | 462 | 463 | 464 | 465 | 466 | 467 | 468 | 469 | 470 | 471 | 472 | 473 | 474 | 475 | 476 | 477 | 478 | 479 | 480 | 481 | 482 | 483 | 484 | 485 | 486 | 487 | 488 | 489 | 490 | 491 | 492 | 493 | 494 | 495 | 496 | 497 | 498 | 499 | 500 | 501 | 502 | 503 | 504 | 505 | 506 | 507 | 508 | 509 | 510 | 511 | 512 | 513 | 514 | 515 | 516 | 517 | 518 | 519 | 520 | 521 | 522 | 523 | 524 | 525 | 526 | 527 | 528 | 529 | 530 | 531 | 532 | 533 | 534 | 535 | 536 | 537 | 538 | 539 | 540 | 541 | 542 | 543 | 544 | 545 | 546 | 547 | 548 | 549 | 550 | 551 | 552 | 553 | 554 | 555 | 556 | 557 | 558 | 559 | 560 | 561 | 562 | 563 | 564 | 565 | 566 | 567 | 568 | 569 | 570 | 571 | 572 | 573 | 574 | 575 | 576 | 577 | 578 | 579 | 580 | 581 | 582 | 583 | 584 | 585 | 586 | 587 | 588 | 589 | 590 | 591 | 592 | 593 | 594 | 595 | 596 | 597 | 598 | 599 | 600 | 601 | 602 | 603 | 604 | 605 | 606 | 607 | 608 | 609 | 610 | 611 | 612 | 613 | 614 | 615 | 616 | 617 | 618 | 619 | 620 | 621 | 622 | 623 | 624 | 625 | 626 | 627 | 628 | 629 | 630 | 631 | 632 | 633 | 634 | 635 | 636 | 637 | 638 | 639 | 640 | 641 | 642 | 643 | 644 | 645 | 646 | 647 | 648 | 649 | 650 | 651 | 652 | 653 | 654 | 655 | 656 | 657 | 658 | 659 | 660 | 661 | 662 | 663 | 664 | 665 | 666 | 667 | 668 | 669 | 670 | 671 | 672 | 673 | 674 | 675 | 676 | 677 | 678 | 679 | 680 | 681 | 682 | 683 | 684 | 685 | 686 | 687 | 688 | 689 | 690 | 691 | 692 | 693 | 694 | 695 | 696 | 697 | 698 | 699 | 700 | 701 | 702 | 703 | 704 | 705 | 706 | 707 | 708 | 709 | 710 | 711 | 712 | 713 | 714 | 715 | 716 | 717 | 718 | 719 | 720 | 721 | 722 | 723 | 724 | 725 | 726 | 727 | 728 | 729 | 730 | 731 | 732 | 733 | 734 | 735 | 736 | 737 | 738 | 739 | 740 | 741 | 742 | 743 | 744 | 745 | 746 | 747 | 748 | 749 | 750 | 751 | 752 | 753 | 754 | 755 | 756 | 757 | 758 | 759 | 760 | 761 | 762 | 763 | 764 | 765 | 766 | 767 | 768 | 769 | 770 | 771 | 772 | 773 | 774 | 775 | 776 | 777 | 778 | 779 | 780 | 781 | 782 | 783 | 784 | 785 | 786 | 787 | 788 | 789 | 790 | 791 | 792 | 793 | 794 | 795 | 796 | 797 | 798 | 799 | 800 | 801 | 802 | 803 | 804 | 805 | 806 | 807 | 808 | 809 | 810 | 811 | 812 | 813 | 814 | 815 | 816 | 817 | 818 | 819 | 820 | 821 | 822 | 823 | 824 | 825 | 826 | 827 | 828 | 829 | 830 | 831 | 832 | 833 | 834 | 835 | 836 | 837 | 838 | 839 | 840 | 841 | 842 | 843 | 844 | 845 | 846 | 847 | 848 | 849 | 850 | 851 | 852 | 853 | 854 | 855 | 856 | 857 | 858 | 859 | 860 | 861 | 862 | 863 | 864 | 865 | 866 | 867 | 868 | 869 | 870 | 871 | 872 | 873 | 874 | 875 | 876 | 877 | 878 | 879 | 880 | 881 | 882 | 883 | 884 | 885 | 886 | 887 | 888 | 889 | 890 | 891 | 892 | 893 | 894 | 895 | 896 | 897 | 898 | 899 | 900 | 901 | 902 | 903 | 904 | 905 | 906 | 907 | 908 | 909 | 910 | 911 | 912 | 913 | 914 | 915 | 916 | 917 | 918 | 919 | 920 | 921 | 922 | 923 | 924 | 925 | 926 | 927 | 928 | 929 | 930 | 931 | 932 | 933 | 934 | 935 | 936 | 937 | 938 | 939 | 940 | 941 | 942 | 943 | 944 | 945 | 946 | 947 | 948 | 949 | 950 | 951 | 952 | 953 | 954 | 955 | 956 | 957 | 958 | 959 | 960 | 961 | 962 | 963 | 964 | 965 | 966 | 967 | 968 | 969 | 970 | 971 | 972 | 973 | 974 | 975 | 976 | 977 | 978 | 979 | 980 | 981 | 982 | 983 | 984 | 985 | 986 | 987 | 988 | 989 | 990 | 991 | 992 | 993 | 994 | 995 | 996 | 997 | 998 | 999 | 1000 | 1001 | 1002 | 1003 | 1004 | 1005 | 1006 | 1007 | 1008 | 1009 | 1010 | 1011 | 1012 | 1013 | 1014 | 1015 | 1016 | 1017 | 1018 | 1019 | 1020 | 1021 | 1022 | 1023 | 1024 | 1025 | 1026 | 1027 | 1028 | 1029 | 1030 | 1031 | 1032 | 1033 | 1034 | 1035 | 1036 | 1037 | 1038 | 1039 | 1040 | 1041 | 1042 | 1043 | 1044 | 1045 | 1046 | 1047 | 1048 | 1049 | 1050 | 1051 | 1052 | 1053 | 1054 | 1055 | 1056 | 1057 | 1058 | 1059 | 1060 | 1061 | 1062 | 1063 | 1064 | 1065 | 1066 | 1067 | 1068 | 1069 | 1070 | 1071 | 1072 | 1073 | 1074 | 1075 | 1076 | 1077 | 1078 | 1079 | 1080 | 1081 | 1082 | 1083 | 1084 | 1085 | 1086 | 1087 | 1088 | 1089 | 1090 | 1091 | 1092 | 1093 | 1094 | 1095 | 1096 | 1097 | 1098 | 1099 | 1100 | 1101 | 1102 | 1103 | 1104 | 1105 | 1106 | 1107 | 1108 | 1109 | 1110 | 1111 | 1112 | 1113 | 1114 | 1115 | 1116 | 1117 | 1118 | 1119 | 1120 | 1121 | 1122 | 1123 | 1124 | 1125 | 1126 | 1127 | 1128 | 1129 | 1130 | 1131 | 1132 | 1133 | 1134 | 1135 | 1136 | 1137 | 1138 | 1139 | 1140 | 1141 | 1142 | 1143 | 1144 | 1145 | 1146 | 1147 | 1148 | 1149 | 1150 | 1151 | 1152 | 1153 | 1154 | 1155 | 1156 | 1157 | 1158 | 1159 | 1160 | 1161 | 1162 | 1163 | 1164 | 1165 | 1166 | 1167 | 1168 | 1169 | 1170 | 1171 | 1172 | 1173 | 1174 | 1175 | 1176 | 1177 | 1178 | 1179 | 1180 | 1181 | 1182 | 1183 | 1184 | 1185 | 1186 | 1187 | 1188 | 1189 | 1190 | 1191 | 1192 | 1193 | 1194 | 1195 | 1196 | 1197 | 1198 | 1199 | 1200 | 1201 | 1202 | 1203 | 1204 | 1205 | 1206 | 1207 | 1208 | 1209 | 1210 | 1211 | 1212 | 1213 | 1214 | 1215 | 1216 | 1217 | 1218 | 1219 | 1220 | 1221 | 1222 | 1223 | 1224 | 1225 | 1226 | 1227 | 1228 | 1229 | 1230 | 1231 | 1232 | 1233 | 1234 | 1235 | 1236 | 1237 | 1238 | 1239 | 1240 | 1241 | 1242 | 1243 | 1244 | 1245 | 1246 | 1247 | 1248 | 1249 | 1250 | 1251 | 1252 | 1253 | 1254 | 1255 | 1256 | 1257 | 1258 | 1259 | 1260 | 1261 | 1262 | 1263 | 1264 | 1265 | 1266 | 1267 | 1268 | 1269 | 1270 | 1271 | 1272 | 1273 | 1274 | 1275 | 1276 | 1277 | 1278 | 1279 | 1280 | 1281 | 1282 | 1283 | 1284 | 1285 | 1286 | 1287 | 1288 | 1289 | 1290 | 1291 | 1292 | 1293 | 1294 | 1295 | 1296 | 1297 | 1298 | 1299 | 1300 | 1301 | 1302 | 1303 | 1304 | 1305 | 1306 | 1307 | 1308 | 1309 | 1310 | 1311 | 1312 | 1313 | 1314 | 1315 | 1316 | 1317 | 1318 | 1319 | 1320 | 1321 | 1322 | 1323 | 1324 | 1325 | 1326 | 1327 | 1328 | 1329 | 1330 | 1331 | 1332 | 1333 | 1334 | 1335 | 1336 | 1337 | 1338 | 1339 | 1340 | 1341 | 1342 | 1343 | 1344 | 1345 | 1346 | 1347 | 1348 | 1349 | 1350 | 1351 | 1352 | 1353 | 1354 | 1355 | 1356 | 1357 | 1358 | 1359 | 1360 | 1361 | 1362 | 1363 | 1364 | 1365 | 1366 | 1367 | 1368 | 1369 | 1370 | 1371 | 1372 | 1373 | 1374 | 1375 | 1376 | 1377 | 1378 | 1379 | 1380 | 1381 | 1382 | 1383 | 1384 | 1385 | 1386 | 1387 | 1388 | 1389 | 1390 | 1391 | 1392 | 1393 | 1394 | 1395 | 1396 | 1397 | 1398 | 1399 | 1400 | 1401 | 1402 | 1403 | 1404 | 1405 | 1406 | 1407 | 1408 | 1409 | 1410 | 1411 | 1412 | 1413 | 1414 | 1415 | 1416 | 1417 | 1418 | 1419 | 1420 | 1421 | 1422 | 1423 | 1424 | 1425 | 1426 | 1427 | 1428 | 1429 | 1430 | 1431 | 1432 | 1433 | 1434 | 1435 | 1436 | 1437 | 1438 | 1439 | 1440 | 1441 | 1442 | 1443 | 1444 | 1445 | 1446 | 1447 | 1448 | 1449 | 1450 | 1451 | 1452 | 1453 | 1454 | 1455 | 1456 | 1457 | 1458 | 1459 | 1460 | 1461 | 1462 | 1463 | 1464 | 1465 | 1466 | 1467 | 1468 | 1469 | 1470 | 1471 | 1472 | 1473 | 1474 | 1475 | 1476 | 1477 | 1478 | 1479 | 1480 | 1481 | 1482 | 1483 | 1484 | 1485 | 1486 | 1487 | 1488 | 1489 | 1490 | 1491 | 1492 | 1493 | 1494 | 1495 | 1496 | 1497 | 1498 | 1499 | 1500 | 1501 | 1502 | 1503 | 1504 | 1505 | 1506 | 1507 | 1508 | 1509 | 1510 | 1511 | 1512 | 1513 | 1514 | 1515 | 1516 | 1517 | 1518 | 1519 | 1520 | 1521 | 1522 | 1523 | 1524 | 1525 | 1526 | 1527 | 1528 | 1529 | 1530 | 1531 | 1532 | 1533 | 1534 | 1535 | 1536 | 1537 | 1538 | 1539 | 1540 | 1541 | 1542 | 1543 | 1544 | 1545 | 1546 | 1547 | 1548 | 1549 | 1550 | 1551 | 1552 | 1553 | 1554 | 1555 | 1556 | 1557 | 1558 | 1559 | 1560 | 1561 | 1562 | 1563 | 1564 | 1565 | 1566 | 1567 | 1568 | 1569 | 1570 | 1571 | 1572 | 1573 | 1574 | 1575 | 1576 | 1577 | 1578 | 1579 | 1580 | 1581 | 1582 | 1583 | 1584 | 1585 | 1586 | 1587 | 1588 | 1589 | 1590 | 1591 | 1592 | 1593 | 1594 | 1595 | 1596 | 1597 | 1598 | 1599 | 1600 | 1601 | 1602 | 1603 | 1604 | 1605 | 1606 | 1607 | 1608 | 1609 | 1610 | 1611 | 1612 | 1613 | 1614 | 1615 | 1616 | 1617 | 1618 | 1619 | 1620 | 1621 | 1622 | 1623 | 1624 | 1625 | 1626 | 1627 | 1628 | 1629 | 1630 | 1631 | 1632 | 1633 | 1634 | 1635 | 1636 | 1637 | 1638 | 1639 | 1640 | 1641 | 1642 | 1643 | 1644 | 1645 | 1646 | 1647 | 1648 | 1649 | 1650 | 1651 | 1652 | 1653 | 1654 | 1655 | 1656 | 1657 | 1658 | 1659 | 1660 | 1661 | 1662 | 1663 | 1664 | 1665 | 1666 | 1667 | 1668 | 1669 | 1670 | 1671 | 1672 | 1673 | 1674 | 1675 | 1676 | 1677 | 1678 | 1679 | 1680 | 1681 | 1682 | 1683 | 1684 | 1685 | 1686 | 1687 | 1688 | 1689 | 1690 | 1691 | 1692 | 1693 | 1694 | 1695 | 1696 | 1697 | 1698 | 1699 | 1700 | 1701 | 1702 | 1703 | 1704 | 1705 | 1706 | 1707 | 1708 | 1709 | 1710 | 1711 | 1712 | 1713 | 1714 | 1715 | 1716 | 1717 | 1718 | 1719 | 1720 | 1721 | 1722 | 1723 | 1724 | 1725 | 1726 | 1727 | 1728 | 1729 | 1730 | 1731 | 1732 | 1733 | 1734 | 1735 | 1736 | 1737 | 1738 | 1739 | 1740 | 1741 | 1742 | 1743 | 1744 | 1745 | 1746 | 1747 | 1748 | 1749 | 1750 | 1751 | 1752 | 1753 | 1754 | 1755 | 1756 | 1757 | 1758 | 1759 | 1760 | 1761 | 1762 | 1763 | 1764 | 1765 | 1766 | 1767 | 1768 | 1769 | 1770 | 1771 | 1772 | 1773 | 1774 | 1775 | 1776 | 1777 | 1778 | 1779 | 1780 | 1781 | 1782 | 1783 | 1784 | 1785 | 1786 | 1787 | 1788 | 1789 | 1790 | 1791 | 1792 | 1793 | 1794 | 1795 | 1796 | 1797 | 1798 | 1799 | 1800 | 1801 | 1802 | 1803 | 1804 | 1805 | 1806 | 1807 | 1808 | 1809 | 1810 | 1811 | 1812 | 1813 | 1814 | 1815 | 1816 | 1817 | 1818 | 1819 | 1820 | 1821 | 1822 | 1823 | 1824 | 1825 | 1826 | 1827 | 1828 | 1829 | 1830 | 1831 | 1832 | 1833 | 1834 | 1835 | 1836 | 1837 | 1838 | 1839 | 1840 | 1841 | 1842 | 1843 | 1844 | 1845 | 1846 | 1847 | 1848 | 1849 | 1850 | 1851 | 1852 | 1853 | 1854 | 1855 | 1856 | 1857 | 1858 | 1859 | 1860 | 1861 | 1862 | 1863 | 1864 | 1865 | 1866 | 1867 | 1868 | 1869 | 1870 | 1871 | 1872 | 1873 | 1874 | 1875 | 1876 | 1877 | 1878 | 1879 | 1880 | 1881 | 1882 | 1883 | 1884 | 1885 | 1886 | 1887 | 1888 | 1889 | 1890 | 1891 | 1892 | 1893 | 1894 | 1895 | 1896 | 1897 | 1898 | 1899 | 1900 | 1901 | 1902 | 1903 | 1904 | 1905 | 1906 | 1907 | 1908 | 1909 | 1910 | 1911 | 1912 | 1913 | 1914 | 1915 | 1916 | 1917 | 1918 | 1919 | 1920 | 1921 | 1922 | 1923 | 1924 | 1925 | 1926 | 1927 | 1928 | 1929 | 1930 | 1931 | 1932 | 1933 | 1934 | 1935 | 1936 | 1937 | 1938 | 1939 | 1940 | 1941 | 1942 | 1943 | 1944 | 1945 | 1946 | 1947 | 1948 | 1949 | 1950 | 1951 | 1952 | 1953 | 1954 | 1955 | 1956 | 1957 | 1958 | 1959 | 1960 | 1961 | 1962 | 1963 | 1964 | 1965 | 1966 | 1967 | 1968 | 1969 | 1970 | 1971 | 1972 | 1973 | 1974 | 1975 | 1976 | 1977 | 1978 | 1979 | 1980 | 1981 | 1982 | 1983 | 1984 | 1985 | 1986 | 1987 | 1988 | 1989 | 1990 | 1991 | 1992 | 1993 | 1994 | 1995 | 1996 | 1997 | 1998 | 1999 | 2000 | 2001 | 2002 | 2003 | 2004 | 2005 | 2006 | 2007 | 2008 | 2009 | 2010 | 2011 | 2012 | 2013 | 2014 | 2015 | 2016 | 2017 | 2018 | 2019 | 2020 | 2021 | 2022 | 2023 | 2024 | 2025 | 2026 | 2027 | 2028 | 2029 | 2030 | 2031 | 2032 | 2033 | 2034 | 2035 | 2036 | 2037 | 2038 | 2039 | 2040 | 2041 | 2042 | 2043 | 2044 | 2045 | 2046 | 2047
1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 2 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 3 | 3 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 2 | 1 | 1 | 2 | 1 | 0 | 1 | 0 | 2 | 1 | 1 | 1 | 1 | 1 | 1 | 2 | 1 | 2 | 1 | 2 | 1 | 2 | 2 | 2 | 4 | 3 | 4 | 5 | 4 | 7 | 3 | 2 | 3 | 1 | 2 | 3 | 2 | 1 | 4 | 2 | 2 | 2 | 1 | 0 | 1 | 1 | 0 | 2 | 2 | 1 | 1 | 2 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 2 | 1 | 1 | 2 | 2 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 2 | 2 | 0 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 2 | 1 | 1 | 3 | 1 | 0 | 0 | 2 | 0 | 0 | 1 | 1 | 1 | 1 | 2 | 1 | 0 | 1 | 2 | 0 | 1 | 0 | 0 | 2 | 2 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 1 | 1 | 1 | 2 | 1 | 1 | 0 | 0 | 1 | 3 | 1 | 2 | 2 | 1 | 1 | 1 | 1 | 2 | 2 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 2 | 0 | 1 | 0 | 0 | 2 | 2 | 1 | 0 | 0 | 1 | 0 | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 2 | 0 | 0 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 2 | 2 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 2 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 2 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 2 | 3122 | 268 | 273 | 290 | 302 | 308 | 345 | 357 | 411 | 443 | 480 | 534 | 573 | 605 | 634 | 696 | 761 | 822 | 921 | 1021 | 1002 | 1120 | 1244 | 1351 | 1405 | 1559 | 1763 | 1821 | 1908 | 2170 | 2369 | 2328 | 2469 | 2639 | 2781 | 2951 | 3134 | 3251 | 3337 | 3434 | 3619 | 3703 | 3664 | 3770 | 3806 | 3863 | 3891 | 3903 | 3894 | 3893 | 3955 | 4078 | 4076 | 4011 | 3941 | 3902 | 3971 | 3987 | 3951 | 3963 | 3994 | 4156 | 4100 | 3993 | 4070 | 4047 | 4013 | 4031 | 4047 | 4074 | 4011 | 4028 | 4199 | 4135 | 4030 | 4079 | 4099 | 4028 | 4073 | 4118 | 4092 | 4007 | 4036 | 4223 | 4106 | 4066 | 4124 | 4184 | 4068 | 4196 | 4147 | 4011 | 4081 | 4302 | 4262 | 4078 | 4029 | 3989 | 3918 | 3923 | 3909 | 3896 | 3901 | 3921 | 4125 | 3970 | 3888 | 3877 | 3899 | 3908 | 3881 | 3870 | 3842 | 3806 | 3791 | 3680 | 3487 | 3426 | 3477 | 3512 | 3518 | 3512 | 3514 | 3508 | 3530 | 3519 | 3443 | 3468 | 3488 | 3523 | 3492 | 3490 | 3539 | 3572 | 3603 | 3602 | 3570 | 3544 | 3553 | 3579 | 3599 | 3570 | 3537 | 3579 | 3628 | 3649 | 3666 | 3674 | 3663 | 3662 | 3723 | 3681 | 3770 | 3796 | 3744 | 3747 | 3842 | 3908 | 3895 | 3891 | 3979 | 3986 | 4011 | 4060 | 4135 | 4160 | 4223 | 4265 | 4293 | 4357 | 4411 | 4474 | 4546 | 4581 | 4658 | 4699 | 4771 | 4808 | 4901 | 4982 | 4948 | 4932 | 5002 | 5093 | 5172 | 5214 | 5246 | 5256 | 5196 | 5157 | 5179 | 5237 | 5264 | 5210 | 5165 | 5108 | 5039 | 5026 | 5059 | 5070 | 5062 | 4942 | 4855 | 4824 | 4832 | 4832 | 4795 | 4686 | 4591 | 4583 | 4600 | 4529 | 4510 | 4395 | 4483 | 4370 | 4350 | 4351 | 4090 | 4190 | 4232 | 4074 | 4097 | 4092 | 4050 | 3991 | 3895 | 3999 | 3958 | 3857 | 3775 | 3814 | 3786 | 3646 | 3626 | 3638 | 3639 | 3574 | 3577 | 3574 | 3588 | 3579 | 3560 | 3557 | 3501 | 3418 | 3359 | 3365 | 3365 | 3344 | 3359 | 3343 | 3259 | 3219 | 3278 | 3218 | 3168 | 3161 | 3177 | 3056 | 3040 | 3074 | 3052 | 3007 | 3039 | 3036 | 3027 | 2899 | 2870 | 2967 | 3002 | 2984 | 2923 | 2913 | 2868 | 2853 | 2848 | 2811 | 2831 | 2755 | 2759 | 2792 | 2710 | 2692 | 2661 | 2653 | 2664 | 2611 | 2634 | 2655 | 2618 | 2550 | 2552 | 2610 | 2618 | 2596 | 2560 | 2551 | 2546 | 2528 | 2525 | 2492 | 2465 | 2456 | 2473 | 2480 | 2444 | 2402 | 2376 | 2381 | 2387 | 2392 | 2359 | 2310 | 2302 | 2318 | 2344 | 2286 | 2209 | 2291 | 2303 | 2223 | 2247 | 2183 | 2212 | 2218 | 2194 | 2134 | 2129 | 2041 | 2005 | 2019 | 1977 | 1980 | 1963 | 1949 | 1985 | 1917 | 1931 | 1924 | 1891 | 1883 | 1866 | 1841 | 1862 | 1872 | 1825 | 1777 | 1746 | 1748 | 1746 | 1750 | 1749 | 1726 | 1697 | 1684 | 1680 | 1651 | 1607 | 1596 | 1634 | 1610 | 1626 | 1654 | 1645 | 1596 | 1594 | 1620 | 1608 | 1582 | 1539 | 1551 | 1577 | 1568 | 1562 | 1583 | 1531 | 1453 | 1478 | 1485 | 1466 | 1484 | 1518 | 1515 | 1496 | 1490 | 1469 | 1463 | 1427 | 1402 | 1424 | 1444 | 1457 | 1447 | 1446 | 1425 | 1391 | 1436 | 1468 | 1436 | 1398 | 1367 | 1369 | 1390 | 1427 | 1400 | 1393 | 1423 | 1377 | 1337 | 1390 | 1422 | 1421 | 1357 | 1325 | 1357 | 1363 | 1374 | 1384 | 1399 | 1390 | 1398 | 1401 | 1387 | 1434 | 1426 | 1458 | 1406 | 1380 | 1396 | 1455 | 1473 | 1503 | 1546 | 1568 | 1556 | 1591 | 1630 | 1687 | 1717 | 1729 | 1790 | 1831 | 1897 | 2011 | 2055 | 2118 | 2229 | 2354 | 2459 | 2535 | 2678 | 2841 | 2983 | 3121 | 3321 | 3524 | 3665 | 3886 | 4112 | 4442 | 4773 | 5103 | 5382 | 5800 | 6105 | 6589 | 6961 | 7366 | 7861 | 8338 | 8892 | 9405 | 10096 | 10802 | 11290 | 12050 | 12656 | 13166 | 13830 | 14820 | 15586 | 16167 | 16983 | 17738 | 18401 | 19011 | 19665 | 20363 | 21221 | 21834 | 22518 | 23204 | 23564 | 24168 | 24816 | 25146 | 25427 | 25780 | 26009 | 26164 | 26362 | 26521 | 26526 | 26336 | 26094 | 25649 | 25407 | 25203 | 24633 | 24035 | 23608 | 22901 | 22146 | 21617 | 20671 | 19855 | 18832 | 18198 | 17415 | 16420 | 15573 | 14761 | 13881 | 13062 | 12449 | 11565 | 10712 | 9912 | 9074 | 8664 | 7812 | 7316 | 6706 | 6010 | 5668 | 5217 | 4722 | 4421 | 4132 | 3832 | 3604 | 3308 | 3054 | 2912 | 2758 | 2606 | 2528 | 2431 | 2316 | 2262 | 2197 | 2103 | 2018 | 1965 | 1935 | 1906 | 1897 | 1894 | 1881 | 1838 | 1808 | 1815 | 1830 | 1853 | 1822 | 1786 | 1769 | 1786 | 1809 | 1799 | 1832 | 1811 | 1782 | 1783 | 1786 | 1762 | 1767 | 1755 | 1773 | 1766 | 1783 | 1697 | 1703 | 1706 | 1668 | 1728 | 1737 | 1689 | 1671 | 1695 | 1677 | 1736 | 1757 | 1725 | 1746 | 1678 | 1669 | 1680 | 1651 | 1615 | 1639 | 1692 | 1686 | 1655 | 1649 | 1630 | 1613 | 1609 | 1597 | 1617 | 1620 | 1599 | 1597 | 1611 | 1564 | 1563 | 1606 | 1629 | 1583 | 1552 | 1557 | 1544 | 1538 | 1536 | 1591 | 1585 | 1548 | 1617 | 1574 | 1484 | 1472 | 1517 | 1533 | 1551 | 1563 | 1535 | 1522 | 1530 | 1549 | 1586 | 1595 | 1564 | 1583 | 1553 | 1546 | 1553 | 1603 | 1708 | 1599 | 1632 | 1580 | 1635 | 1659 | 1693 | 1653 | 1648 | 1670 | 1686 | 1675 | 1661 | 1707 | 1728 | 1724 | 1746 | 1720 | 1771 | 1854 | 1973 | 1867 | 1856 | 1795 | 1864 | 1861 | 1851 | 1866 | 1880 | 1902 | 1927 | 1944 | 1943 | 1968 | 1988 | 1984 | 1978 | 1952 | 1985 | 2016 | 1987 | 2003 | 1980 | 1977 | 2000 | 1974 | 1963 | 1943 | 1929 | 1904 | 1890 | 1916 | 1900 | 1918 | 1925 | 1928 | 1888 | 1831 | 1870 | 1913 | 1789 | 1772 | 1767 | 1826 | 1778 | 1737 | 1752 | 1744 | 1717 | 1709 | 1728 | 1679 | 1646 | 1609 | 1545 | 1552 | 1531 | 1541 | 1589 | 1551 | 1450 | 1494 | 1447 | 1412 | 1364 | 1350 | 1391 | 1419 | 1413 | 1400 | 1385 | 1354 | 1318 | 1287 | 1257 | 1241 | 1256 | 1280 | 1273 | 1238 | 1219 | 1240 | 1228 | 1206 | 1187 | 1170 | 1148 | 1133 | 1129 | 1128 | 1138 | 1114 | 1079 | 1076 | 1044 | 1034 | 1037 | 1078 | 1086 | 1039 | 1059 | 1060 | 974 | 982 | 993 | 1007 | 1005 | 992 | 967 | 956 | 968 | 946 | 917 | 946 | 906 | 950 | 885 | 925 | 969 | 878 | 887 | 906 | 864 | 859 | 879 | 859 | 831 | 828 | 834 | 850 | 857 | 836 | 834 | 851 | 818 | 797 | 823 | 826 | 783 | 790 | 772 | 735 | 711 | 727 | 735 | 730 | 729 | 741 | 735 | 718 | 734 | 749 | 707 | 695 | 680 | 688 | 706 | 725 | 684 | 674 | 672 | 654 | 628 | 619 | 635 | 605 | 569 | 582 | 603 | 613 | 613 | 604 | 610 | 587 | 576 | 595 | 616 | 550 | 595 | 569 | 544 | 553 | 535 | 552 | 584 | 555 | 532 | 537 | 522 | 506 | 511 | 502 | 480 | 491 | 514 | 521 | 521 | 512 | 491 | 481 | 461 | 473 | 478 | 482 | 480 | 459 | 444 | 446 | 457 | 470 | 483 | 469 | 465 | 462 | 464 | 473 | 454 | 432 | 446 | 448 | 453 | 455 | 453 | 443 | 435 | 445 | 439 | 432 | 440 | 417 | 432 | 455 | 412 | 430 | 380 | 453 | 432 | 413 | 410 | 420 | 398 | 424 | 435 | 427 | 426 | 418 | 427 | 444 | 427 | 420 | 445 | 454 | 415 | 428 | 424 | 415 | 425 | 395 | 442 | 464 | 454 | 421 | 425 | 429 | 446 | 464 | 466 | 463 | 457 | 445 | 461 | 487 | 495 | 495 | 502 | 522 | 543 | 536 | 510 | 515 | 546 | 564 | 574 | 588 | 604 | 614 | 646 | 700 | 696 | 687 | 741 | 751 | 751 | 793 | 875 | 909 | 872 | 931 | 964 | 993 | 1017 | 1093 | 1130 | 1139 | 1190 | 1254 | 1324 | 1392 | 1447 | 1496 | 1499 | 1528 | 1586 | 1675 | 1807 | 1823 | 1845 | 1930 | 2009 | 2091 | 2158 | 2200 | 2258 | 2329 | 2379 | 2415 | 2493 | 2575 | 2591 | 2598 | 2674 | 2732 | 2756 | 2917 | 2931 | 2904 | 2982 | 2989 | 3020 | 3040 | 3015 | 2998 | 3021 | 3041 | 3030 | 3027 | 3014 | 2971 | 3000 | 3044 | 2987 | 2891 | 2813 | 2889 | 2751 | 2732 | 2732 | 2659 | 2576 | 2500 | 2393 | 2339 | 2308 | 2271 | 2201 | 2098 | 2016 | 1895 | 1876 | 1832 | 1705 | 1618 | 1624 | 1565 | 1448 | 1369 | 1315 | 1273 | 1168 | 1054 | 979 | 953 | 918 | 839 | 782 | 748 | 712 | 675 | 615 | 599 | 572 | 498 | 480 | 483 | 442 | 407 | 384 | 352 | 342 | 314 | 299 | 304 | 287 | 259 | 241 | 235 | 240 | 218 | 194 | 190 | 187 | 178 | 166 | 156 | 162 | 157 | 155 | 156 | 158 | 148 | 144 | 143 | 132 | 133 | 142 | 136 | 142 | 154 | 137 | 141 | 145 | 143 | 127 | 118 | 127 | 129 | 118 | 120 | 135 | 134 | 130 | 119 | 116 | 122 | 122 | 121 | 117 | 112 | 116 | 132 | 131 | 114 | 117 | 113 | 112 | 119 | 119 | 121 | 112 | 118 | 127 | 122 | 119 | 120 | 120 | 117 | 107 | 111 | 113 | 111 | 108 | 114 | 113 | 114 | 121 | 117 | 117 | 115 | 114 | 112 | 118 | 120 | 117 | 123 | 127 | 121 | 108 | 117 | 108 | 118 | 113 | 127 | 110 | 98 | 100 | 101 | 118 | 115 | 106 | 109 | 102 | 97 | 105 | 110 | 111 | 112 | 107 | 106 | 109 | 106 | 114 | 123 | 107 | 100 | 109 | 119 | 117 | 117 | 108 | 102 | 104 | 106 | 106 | 104 | 106 | 118 | 126 | 116 | 106 | 97 | 106 | 110 | 104 | 97 | 99 | 106 | 102 | 98 | 94 | 90 | 92 | 96 | 95 | 94 | 97 | 91 | 98 | 99 | 100 | 90 | 103 | 92 | 105 | 92 | 84 | 87 | 79 | 82 | 97 | 99 | 83 | 70 | 71 | 82 | 81 | 81 | 89 | 75 | 72 | 86 | 81 | 91 | 94 | 82 | 83 | 80 | 73 | 73 | 76 | 78 | 74 | 76 | 82 | 80 | 81 | 81 | 73 | 79 | 76 | 67 | 70 | 74 | 76 | 77 | 72 | 74 | 75 | 73 | 69 | 65 | 65 | 64 | 66 | 71 | 67 | 66 | 61 | 72 | 67 | 63 | 61 | 73 | 63 | 65 | 61 | 60 | 65 | 59 | 50 | 46 | 56 | 64 | 59 | 58 | 61 | 58 | 51 | 52 | 65 | 66 | 55 | 52 | 55 | 54 | 52 | 56 | 62 | 60 | 55 | 55 | 55 | 53 | 49 | 45 | 46 | 55 | 55 | 54 | 50 | 46 | 49 | 52 | 56 | 56 | 51 | 51 | 53 | 53 | 53 | 49 | 49 | 55 | 53 | 53 | 52 | 53 | 47 | 42 | 49 | 51 | 53 | 51 | 45 | 48 | 47 | 47 | 47 | 43 | 40 | 42 | 48 | 52 | 48 | 54 | 55 | 54 | 51 | 48 | 61 | 52 | 53 | 47 | 47 | 48 | 45 | 46 | 43 | 47 | 47 | 46 | 48 | 42 | 38 | 41 | 42 | 40 | 46 | 57 | 51 | 46 | 44 | 42 | 45 | 46 | 49 | 49 | 47 | 47 | 48 | 50 | 45 | 41 | 41 | 51 | 51 | 46 | 43 | 52 | 42 | 48 | 43 | 41 | 44 | 45 | 50 | 49 | 42 | 39 | 46 | 48 | 44 | 44 | 47 | 50 | 40 | 32 | 47 | 52 | 55 | 47 | 45 | 39 | 46 | 39 | 43 | 38 | 43 | 37 | 38 | 50 | 48 | 40 | 45 | 52 | 52 | 51 | 47 | 50 | 51 | 51 | 52 | 55 | 53 | 54 | 54 | 57 | 55 | 51 | 51 | 56 | 61 | 58 | 55 | 54 | 55 | 49 | 49 | 50 | 43 | 54 | 56 | 52 | 60 | 56 | 60 | 61 | 48 | 61 | 59 | 67 | 42 | 55 | 72 | 61 | 68 | 55 | 56 | 56 | 44 | 56 | 70 | 67 | 65 | 67 | 67 | 64 | 63 | 60 | 52 | 62 | 53 | 56 | 72 | 65 | 75 | 60 | 55 | 54 | 60 | 64 | 58 | 67 | 65 | 57 | 60 | 65 | 71 | 73 | 65 | 67 | 59 | 61 | 64 | 68 | 68 | 68 | 62 | 67 | 52 | 60 | 58 | 52 | 42 | 56 | 63 | 54 | 57 | 45 | 43 | 53 | 42 | 45 | 46 | 45 | 41 | 41 | 32 | 40 | 41 | 47 | 46 | 35 | 28 | 33 | 40 | 31 | 38 | 32 | 26 | 28 | 35 | 36 | 24 | 17 | 19 | 21 | 21 | 21 | 23 | 26 | 27 | 26 | 20 | 19 | 21 | 20 | 19 | 20 | 21 | 24 | 24 | 25 | 20 | 14 | 15 | 15 | 16 | 19 | 22 | 16 | 15 | 13 | 14 | 18 | 16 | 16 | 9 | 11 | 16 | 16 | 18 | 15 | 12 | 15 | 13 | 11 | 11 | 12 | 16 | 9 | 16 | 16 | 16 | 16 | 9 | 16 | 14 | 18 | 8 | 12 | 13 | 13 | 6 | 9 | 11 | 10 | 14 | 16 | 8 | 9 | 10 | 9 | 10 | 14 | 11 | 12 | 12 | 11 | 12 | 8 | 9 | 11 | 13 | 10 | 9 | 15 | 18 | 16 | 13 | 15 | 15 | 11 | 11 | 12 | 11 | 10 | 18 | 13 | 9 | 14 | 13 | 13 | 9 | 8 | 6 | 9 | 15 | 11 | 16 | 13 | 11 | 8 | 13 | 7 | 7 | 13 | 12 | 11 | 12 | 13 | 7 | 7 | 16 | 11 | 8 | 13 | 12 | 13 | 10 | 9 | 11 | 5 | 5 | 6 | 9 | 10 | 10 | 7 | 8 | 9 | 8 | 7 | 8 | 5 | 8 | 11 | 10 | 7 | 10 | 12 | 9 | 9 | 9 | 7 | 10 | 10 | 8 | 10 | 7 | 5 | 9 | 11 | 8 | 7 | 11 | 9 | 6 | 2 | 10 | 13 | 5 | 9 | 10 | 16 | 5 | 4 | 8 | 2 | 13 | 8 | 9 | 11 | 8 | 5 | 9 | 6 | 3 | 5 | 9 | 5 | 6 | 5 | 5 | 7 | 5 | 7 | 9 | 8 | 4 | 6 | 6 | 7 | 7 | 5 | 8 | 8 | 7 | 5 | 4 | 5 | 7 | 6 | 6 | 7 | 6 | 8 | 9 | 8 | 7 | 5 | 4 | 4 | 4 | 3 | 4 | 3 | 2 | 4 | 4 | 5 | 6 | 10 | 3 | 5 | 4 | 3 | 2 | 7 | 7 | 4 | 8 | 7 | 6 | 2 | 2 | 5 | 6 | 6 | 3 | 4 | 6 | 3 | 4 | 6 | 5 | 6 | 4 | 3 | 5 | 6 | 6 | 9 | 3 | 0 | 1 | 3 | 4 | 3 | 5 | 7 | 6 | 7 | 5 | 1 | 2 | 5 | 4 | 4 | 5 | 4 | 5 | 3 | 3 | 4 | 6 | 5 | 4 | 3 | 3 | 4 | 5 | 3 | 3 | 3 | 2 | 1 | 3 | 1 | 2 | 1 | 2 | 1 | 3 | 2 | 2 | 3 | 1 | 1 | 5 | 6 | 4 | 7 | 5 | 2 | 1 | 4 | 3 | 5 | 3 | 2 | 4 | 6 | 5 | 3 | 1 | 3 | 4 | 3 | 3 | 4 | 3 | 4 | 4 | 2 | 3 | 5 | 3 | 3 | 3 | 6 | 6 | 5 | 4 | 4 | 4 | 3 | 3 | 2 | 1 | 5 | 4 | 3 | 2 | 5 | 3 | 3 | 2 | 5 | 4 | 4 | 0 | 4 | 2 | 5 | 4 | 3 | 5 | 4 | 5 | 5 | 2 | 3 | 1 | 1 | 5 | 6 | 3 | 4 | 2 | 3 | 4 | 5 | 4 | 4 | 5 | 4 | 3 | 3 | 4 | 5 | 4 | 5 | 3 | 3 | 2 | 2 | 2 | 5 | 5 | 2 | 2 | 2 | 3 | 3 | 4 | 4 | 3 | 3 | 4 | 2 | 3 | 5 | 2 | 5 | 7 | 2 | 8 | 6 | 7 | 4 | 1 | 2 | 8 | 4 | 6 | 2 | 10 | 6 | 7 | 4 | 6 | 4 | 3 | 7 | 2 | 5 | 3 | 1 | 4 | 4 | 3 | 5 | 3 | 2 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 3 | 3017 | 220 | 238 | 268 | 293 | 300 | 317 | 348 | 388 | 419 | 471 | 516 | 549 | 577 | 626 | 688 | 767 | 794 | 901 | 1008 | 1017 | 1133 | 1240 | 1329 | 1436 | 1570 | 1674 | 1780 | 1935 | 2115 | 2246 | 2314 | 2547 | 2699 | 2806 | 2905 | 3034 | 3229 | 3273 | 3421 | 3582 | 3653 | 3643 | 3721 | 3765 | 3767 | 3800 | 3861 | 3896 | 3915 | 4019 | 4075 | 3931 | 3931 | 3962 | 3988 | 3946 | 3917 | 3929 | 3983 | 4009 | 4090 | 4112 | 4023 | 4048 | 4076 | 4045 | 3996 | 4015 | 4015 | 4013 | 4048 | 4203 | 4013 | 3944 | 4025 | 4007 | 4023 | 4021 | 4042 | 4018 | 4008 | 4171 | 4251 | 4111 | 4012 | 4053 | 4024 | 4066 | 4161 | 4071 | 4054 | 4141 | 4184 | 4136 | 4003 | 3988 | 3966 | 3958 | 3857 | 3841 | 3946 | 3878 | 3890 | 4021 | 3907 | 3887 | 3833 | 3833 | 3866 | 3857 | 3820 | 3823 | 3816 | 3842 | 3697 | 3421 | 3478 | 3420 | 3443 | 3514 | 3469 | 3506 | 3523 | 3505 | 3451 | 3453 | 3512 | 3486 | 3492 | 3516 | 3512 | 3494 | 3497 | 3507 | 3524 | 3517 | 3485 | 3494 | 3515 | 3575 | 3646 | 3659 | 3624 | 3586 | 3625 | 3646 | 3664 | 3712 | 3674 | 3692 | 3766 | 3730 | 3781 | 3728 | 3787 | 3911 | 3893 | 3831 | 3771 | 3904 | 4030 | 4051 | 4038 | 4078 | 4153 | 4189 | 4226 | 4259 | 4316 | 4484 | 4443 | 4466 | 4510 | 4596 | 4628 | 4673 | 4784 | 4841 | 4838 | 4912 | 4978 | 4978 | 5027 | 5060 | 5104 | 5121 | 5134 | 5171 | 5236 | 5258 | 5173 | 5117 | 5147 | 5159 | 5132 | 5082 | 5074 | 5101 | 5027 | 4946 | 4908 | 4934 | 4901 | 4895 | 4839 | 4734 | 4703 | 4638 | 4517 | 4498 | 4532 | 4436 | 4409 | 4515 | 4371 | 4355 | 4217 | 4282 | 4078 | 4274 | 4060 | 3995 | 3902 | 3954 | 3990 | 3918 | 3919 | 3921 | 3868 | 3738 | 3738 | 3767 | 3718 | 3705 | 3724 | 3698 | 3603 | 3581 | 3587 | 3540 | 3453 | 3391 | 3429 | 3462 | 3406 | 3342 | 3323 | 3347 | 3353 | 3329 | 3305 | 3270 | 3255 | 3204 | 3187 | 3176 | 3199 | 3200 | 3112 | 3096 | 3039 | 3053 | 3123 | 3087 | 3072 | 3024 | 2969 | 3024 | 2947 | 2856 | 2881 | 2922 | 2860 | 2917 | 2858 | 2794 | 2794 | 2745 | 2696 | 2728 | 2715 | 2735 | 2704 | 2721 | 2654 | 2644 | 2638 | 2596 | 2576 | 2579 | 2586 | 2553 | 2549 | 2574 | 2546 | 2534 | 2553 | 2507 | 2482 | 2489 | 2430 | 2446 | 2517 | 2501 | 2461 | 2439 | 2362 | 2267 | 2310 | 2344 | 2370 | 2363 | 2308 | 2298 | 2319 | 2323 | 2266 | 2244 | 2305 | 2304 | 2206 | 2151 | 2153 | 2144 | 2139 | 2139 | 2126 | 2069 | 2084 | 2077 | 1990 | 2027 | 1978 | 1950 | 1981 | 1982 | 1957 | 1924 | 1892 | 1880 | 1862 | 1838 | 1857 | 1861 | 1814 | 1793 | 1802 | 1769 | 1761 | 1775 | 1757 | 1734 | 1709 | 1664 | 1663 | 1627 | 1642 | 1636 | 1655 | 1645 | 1640 | 1635 | 1639 | 1612 | 1590 | 1547 | 1582 | 1578 | 1559 | 1554 | 1538 | 1523 | 1518 | 1531 | 1560 | 1537 | 1521 | 1494 | 1503 | 1470 | 1467 | 1481 | 1470 | 1456 | 1451 | 1438 | 1427 | 1435 | 1425 | 1422 | 1425 | 1429 | 1433 | 1399 | 1377 | 1387 | 1385 | 1362 | 1370 | 1418 | 1424 | 1413 | 1408 | 1393 | 1392 | 1401 | 1379 | 1379 | 1395 | 1397 | 1403 | 1357 | 1318 | 1360 | 1365 | 1391 | 1372 | 1356 | 1386 | 1368 | 1330 | 1377 | 1404 | 1416 | 1443 | 1390 | 1405 | 1467 | 1423 | 1439 | 1452 | 1508 | 1550 | 1582 | 1561 | 1572 | 1637 | 1668 | 1705 | 1751 | 1808 | 1867 | 1945 | 2008 | 2055 | 2105 | 2163 | 2262 | 2383 | 2481 | 2602 | 2753 | 2870 | 2983 | 3193 | 3442 | 3665 | 3885 | 4150 | 4390 | 4681 | 5001 | 5309 | 5586 | 6050 | 6423 | 6815 | 7231 | 7877 | 8265 | 8781 | 9381 | 9878 | 10794 | 11422 | 11802 | 12484 | 13181 | 13775 | 14633 | 15365 | 16036 | 16782 | 17513 | 18204 | 18728 | 19652 | 20368 | 21033 | 21676 | 22400 | 23080 | 23579 | 24028 | 24659 | 25149 | 25399 | 25610 | 25926 | 26210 | 26289 | 26246 | 26206 | 26030 | 25823 | 25568 | 25271 | 24922 | 24415 | 23963 | 23510 | 22879 | 22252 | 21605 | 20632 | 19802 | 19054 | 18104 | 17156 | 16239 | 15487 | 14615 | 13897 | 13005 | 12069 | 11327 | 10475 | 9875 | 9041 | 8561 | 7745 | 7305 | 6656 | 6158 | 5712 | 5125 | 4849 | 4478 | 4033 | 3787 | 3515 | 3268 | 3077 | 2924 | 2778 | 2637 | 2510 | 2401 | 2279 | 2188 | 2153 | 2099 | 2042 | 1984 | 1912 | 1894 | 1901 | 1872 | 1850 | 1854 | 1837 | 1801 | 1786 | 1818 | 1839 | 1790 | 1762 | 1804 | 1797 | 1778 | 1768 | 1788 | 1820 | 1791 | 1808 | 1813 | 1794 | 1749 | 1756 | 1710 | 1732 | 1744 | 1692 | 1742 | 1728 | 1714 | 1707 | 1690 | 1701 | 1690 | 1710 | 1654 | 1654 | 1674 | 1694 | 1659 | 1657 | 1650 | 1661 | 1653 | 1616 | 1604 | 1593 | 1622 | 1655 | 1645 | 1628 | 1648 | 1635 | 1606 | 1582 | 1582 | 1567 | 1596 | 1616 | 1608 | 1567 | 1567 | 1591 | 1608 | 1600 | 1557 | 1545 | 1561 | 1576 | 1571 | 1584 | 1588 | 1582 | 1511 | 1487 | 1489 | 1454 | 1480 | 1505 | 1534 | 1566 | 1566 | 1550 | 1544 | 1538 | 1532 | 1499 | 1493 | 1523 | 1565 | 1601 | 1646 | 1602 | 1575 | 1529 | 1553 | 1602 | 1627 | 1637 | 1631 | 1662 | 1690 | 1701 | 1721 | 1704 | 1710 | 1774 | 1712 | 1777 | 1830 | 1791 | 1822 | 1808 | 1739 | 1777 | 1780 | 1820 | 1889 | 1934 | 1906 | 1855 | 1863 | 1893 | 1899 | 1925 | 1949 | 1927 | 1930 | 1921 | 1951 | 2015 | 2035 | 1956 | 1983 | 1965 | 1924 | 1939 | 1965 | 1921 | 1894 | 1919 | 1928 | 1889 | 1877 | 1881 | 1871 | 1871 | 1868 | 1829 | 1854 | 1850 | 1838 | 1811 | 1808 | 1803 | 1745 | 1724 | 1731 | 1698 | 1674 | 1658 | 1645 | 1652 | 1620 | 1562 | 1543 | 1542 | 1479 | 1562 | 1544 | 1584 | 1485 | 1507 | 1465 | 1408 | 1394 | 1366 | 1356 | 1341 | 1328 | 1320 | 1324 | 1330 | 1308 | 1265 | 1242 | 1269 | 1269 | 1258 | 1244 | 1212 | 1148 | 1165 | 1163 | 1135 | 1143 | 1127 | 1103 | 1096 | 1112 | 1138 | 1142 | 1114 | 1101 | 1110 | 1089 | 1050 | 1031 | 1056 | 1077 | 1023 | 1008 | 1022 | 1026 | 1018 | 1007 | 974 | 941 | 948 | 958 | 931 | 928 | 953 | 940 | 959 | 940 | 931 | 935 | 899 | 961 | 905 | 885 | 824 | 868 | 899 | 876 | 847 | 855 | 843 | 810 | 814 | 828 | 817 | 820 | 798 | 769 | 774 | 753 | 771 | 797 | 745 | 777 | 757 | 759 | 756 | 748 | 745 | 741 | 725 | 727 | 730 | 713 | 695 | 699 | 716 | 716 | 695 | 662 | 659 | 655 | 649 | 655 | 665 | 644 | 640 | 660 | 656 | 632 | 622 | 625 | 608 | 586 | 591 | 591 | 586 | 582 | 598 | 583 | 606 | 595 | 610 | 591 | 563 | 547 | 543 | 543 | 539 | 538 | 546 | 555 | 540 | 522 | 511 | 478 | 448 | 494 | 513 | 511 | 530 | 495 | 509 | 528 | 478 | 468 | 471 | 461 | 447 | 457 | 469 | 461 | 443 | 437 | 462 | 460 | 425 | 441 | 457 | 462 | 448 | 452 | 443 | 432 | 430 | 428 | 438 | 430 | 431 | 447 | 442 | 427 | 445 | 433 | 419 | 418 | 403 | 432 | 449 | 408 | 408 | 403 | 435 | 418 | 405 | 410 | 409 | 428 | 431 | 432 | 446 | 431 | 415 | 389 | 387 | 419 | 439 | 425 | 442 | 412 | 424 | 446 | 445 | 434 | 418 | 427 | 412 | 419 | 446 | 447 | 452 | 462 | 449 | 464 | 485 | 482 | 484 | 516 | 511 | 502 | 524 | 515 | 538 | 553 | 553 | 581 | 603 | 608 | 621 | 635 | 661 | 685 | 721 | 751 | 747 | 767 | 804 | 810 | 859 | 878 | 901 | 970 | 1017 | 1048 | 1085 | 1143 | 1190 | 1223 | 1270 | 1320 | 1359 | 1403 | 1443 | 1505 | 1532 | 1631 | 1646 | 1902 | 1847 | 1871 | 1936 | 1934 | 2075 | 2038 | 2099 | 2214 | 2279 | 2333 | 2356 | 2391 | 2462 | 2532 | 2655 | 2671 | 2696 | 2757 | 2903 | 2988 | 2973 | 3010 | 3017 | 2945 | 3047 | 3075 | 3022 | 3023 | 3021 | 3017 | 3022 | 2996 | 2999 | 3002 | 2966 | 2928 | 2891 | 2876 | 2895 | 2809 | 2697 | 2663 | 2581 | 2570 | 2529 | 2434 | 2339 | 2275 | 2200 | 2101 | 2043 | 1998 | 1907 | 1814 | 1762 | 1678 | 1578 | 1534 | 1555 | 1367 | 1268 | 1310 | 1206 | 1104 | 1113 | 1032 | 958 | 901 | 860 | 832 | 770 | 691 | 670 | 640 | 593 | 542 | 514 | 483 | 446 | 392 | 396 | 380 | 331 | 304 | 302 | 298 | 282 | 277 | 265 | 240 | 228 | 219 | 205 | 194 | 181 | 185 | 180 | 185 | 192 | 171 | 157 | 151 | 172 | 177 | 149 | 138 | 148 | 143 | 132 | 135 | 136 | 134 | 140 | 138 | 139 | 138 | 145 | 140 | 119 | 124 | 126 | 128 | 115 | 114 | 115 | 132 | 124 | 120 | 125 | 116 | 127 | 112 | 109 | 119 | 108 | 107 | 110 | 126 | 122 | 117 | 127 | 129 | 126 | 125 | 126 | 126 | 127 | 118 | 112 | 112 | 114 | 122 | 115 | 115 | 124 | 126 | 126 | 124 | 116 | 109 | 110 | 117 | 124 | 119 | 112 | 117 | 126 | 128 | 118 | 105 | 101 | 110 | 114 | 123 | 123 | 117 | 113 | 132 | 115 | 103 | 104 | 110 | 111 | 117 | 108 | 103 | 108 | 112 | 122 | 124 | 107 | 103 | 101 | 113 | 121 | 113 | 98 | 100 | 95 | 106 | 100 | 110 | 104 | 104 | 107 | 103 | 102 | 105 | 107 | 108 | 112 | 112 | 103 | 93 | 104 | 107 | 96 | 110 | 109 | 99 | 99 | 98 | 99 | 104 | 102 | 97 | 98 | 102 | 99 | 89 | 90 | 95 | 91 | 92 | 104 | 92 | 84 | 96 | 89 | 87 | 80 | 80 | 84 | 80 | 77 | 81 | 92 | 88 | 77 | 80 | 73 | 84 | 84 | 80 | 95 | 84 | 84 | 71 | 83 | 86 | 77 | 75 | 78 | 78 | 76 | 74 | 71 | 68 | 71 | 73 | 72 | 67 | 72 | 74 | 78 | 76 | 69 | 58 | 69 | 67 | 70 | 72 | 70 | 67 | 69 | 75 | 73 | 66 | 65 | 64 | 65 | 65 | 63 | 64 | 62 | 65 | 60 | 59 | 63 | 67 | 71 | 71 | 70 | 75 | 75 | 65 | 61 | 66 | 65 | 60 | 62 | 53 | 61 | 53 | 58 | 59 | 46 | 49 | 64 | 59 | 67 | 61 | 57 | 58 | 60 | 57 | 51 | 50 | 50 | 56 | 55 | 56 | 48 | 53 | 52 | 55 | 55 | 46 | 54 | 56 | 52 | 52 | 51 | 50 | 51 | 52 | 49 | 50 | 59 | 56 | 44 | 43 | 48 | 53 | 56 | 50 | 51 | 53 | 47 | 43 | 40 | 42 | 47 | 47 | 42 | 44 | 47 | 46 | 43 | 46 | 41 | 49 | 52 | 48 | 51 | 46 | 46 | 46 | 58 | 48 | 40 | 51 | 46 | 46 | 57 | 52 | 49 | 43 | 43 | 50 | 45 | 50 | 42 | 37 | 44 | 45 | 49 | 42 | 41 | 45 | 47 | 43 | 36 | 37 | 41 | 41 | 43 | 50 | 48 | 43 | 43 | 43 | 46 | 43 | 44 | 51 | 49 | 40 | 41 | 42 | 43 | 45 | 46 | 44 | 42 | 48 | 49 | 41 | 39 | 40 | 35 | 43 | 54 | 46 | 44 | 50 | 47 | 45 | 45 | 33 | 42 | 52 | 41 | 36 | 37 | 45 | 35 | 54 | 40 | 53 | 54 | 52 | 52 | 49 | 44 | 39 | 46 | 47 | 45 | 45 | 49 | 44 | 56 | 51 | 40 | 51 | 51 | 53 | 54 | 49 | 52 | 55 | 61 | 60 | 52 | 55 | 57 | 64 | 56 | 51 | 53 | 49 | 51 | 52 | 49 | 61 | 67 | 59 | 56 | 53 | 56 | 52 | 50 | 53 | 49 | 55 | 67 | 68 | 74 | 62 | 71 | 70 | 71 | 67 | 62 | 61 | 68 | 61 | 73 | 68 | 74 | 59 | 52 | 71 | 74 | 72 | 59 | 55 | 63 | 68 | 62 | 64 | 70 | 56 | 52 | 51 | 61 | 63 | 60 | 65 | 62 | 56 | 55 | 59 | 54 | 58 | 51 | 47 | 55 | 61 | 59 | 53 | 50 | 62 | 56 | 41 | 40 | 38 | 43 | 43 | 45 | 32 | 33 | 41 | 39 | 39 | 37 | 42 | 37 | 38 | 23 | 24 | 29 | 33 | 22 | 24 | 18 | 24 | 34 | 30 | 28 | 17 | 29 | 32 | 20 | 29 | 23 | 21 | 14 | 16 | 27 | 20 | 16 | 21 | 17 | 26 | 20 | 27 | 21 | 17 | 14 | 15 | 16 | 18 | 15 | 16 | 14 | 14 | 15 | 15 | 19 | 17 | 10 | 9 | 14 | 15 | 11 | 9 | 14 | 15 | 11 | 9 | 11 | 14 | 16 | 12 | 12 | 13 | 15 | 12 | 12 | 9 | 14 | 15 | 12 | 12 | 9 | 15 | 5 | 13 | 10 | 10 | 12 | 13 | 12 | 13 | 15 | 8 | 13 | 6 | 10 | 6 | 12 | 14 | 16 | 13 | 9 | 4 | 9 | 10 | 7 | 9 | 16 | 9 | 8 | 10 | 9 | 12 | 16 | 14 | 10 | 15 | 13 | 9 | 11 | 12 | 11 | 10 | 11 | 10 | 7 | 10 | 9 | 9 | 10 | 15 | 14 | 14 | 15 | 12 | 9 | 8 | 9 | 11 | 12 | 14 | 16 | 10 | 13 | 8 | 12 | 11 | 7 | 11 | 11 | 8 | 14 | 11 | 3 | 11 | 5 | 10 | 16 | 7 | 9 | 11 | 11 | 12 | 11 | 10 | 9 | 6 | 10 | 4 | 9 | 9 | 5 | 9 | 10 | 12 | 10 | 8 | 9 | 9 | 11 | 13 | 9 | 9 | 7 | 7 | 9 | 10 | 10 | 8 | 10 | 9 | 5 | 5 | 7 | 10 | 10 | 8 | 6 | 6 | 7 | 5 | 9 | 9 | 10 | 6 | 9 | 8 | 6 | 10 | 7 | 6 | 6 | 10 | 6 | 8 | 5 | 13 | 3 | 6 | 6 | 6 | 3 | 2 | 4 | 2 | 4 | 5 | 2 | 6 | 8 | 9 | 3 | 5 | 8 | 9 | 4 | 5 | 7 | 5 | 7 | 3 | 4 | 2 | 3 | 2 | 4 | 6 | 4 | 2 | 3 | 4 | 3 | 3 | 4 | 2 | 2 | 3 | 2 | 3 | 4 | 5 | 3 | 3 | 3 | 4 | 3 | 2 | 2 | 4 | 3 | 3 | 4 | 5 | 6 | 4 | 4 | 1 | 3 | 4 | 5 | 3 | 4 | 6 | 6 | 6 | 2 | 6 | 1 | 6 | 2 | 5 | 1 | 4 | 3 | 3 | 6 | 7 | 5 | 6 | 2 | 3 | 4 | 4 | 3 | 4 | 5 | 4 | 3 | 3 | 2 | 4 | 4 | 2 | 2 | 4 | 2 | 3 | 4 | 4 | 4 | 3 | 3 | 2 | 2 | 3 | 3 | 3 | 4 | 3 | 5 | 6 | 2 | 2 | 6 | 5 | 4 | 3 | 3 | 0 | 4 | 0 | 3 | 1 | 4 | 2 | 6 | 8 | 2 | 3 | 1 | 1 | 1 | 3 | 2 | 6 | 3 | 2 | 2 | 2 | 1 | 3 | 4 | 3 | 4 | 5 | 3 | 3 | 2 | 2 | 2 | 2 | 1 | 5 | 6 | 2 | 3 | 4 | 4 | 4 | 3 | 3 | 3 | 4 | 3 | 3 | 3 | 2 | 2 | 5 | 3 | 2 | 2 | 4 | 4 | 5 | 3 | 3 | 8 | 4 | 4 | 6 | 3 | 6 | 4 | 2 | 0 | 5 | 0 | 3 | 4 | 1 | 4 | 2 | 0 | 4 | 3 | 7 | 3 | 6 | 2 | 4 | 3 | 5 | 3 | 3 | 5 | 4 | 3 | 3 | 4 | 3 | 5 | 5 | 3 | 2 | 4 | 4 | 2 | 4 | 4 | 3 | 6 | 4 | 2 | 4 | 3 | 4 | 4 | 3 | 3 | 3 | 3 | 3 | 1 | 4 | 5 | 1 | 2 | 2 | 2 | 4 | 6 | 6 | 3 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 4 | 2901 | 226 | 236 | 257 | 271 | 291 | 294 | 307 | 373 | 433 | 454 | 469 | 495 | 553 | 605 | 619 | 672 | 757 | 796 | 916 | 935 | 1017 | 1154 | 1258 | 1378 | 1468 | 1572 | 1691 | 1795 | 2040 | 2232 | 2300 | 2428 | 2542 | 2626 | 2843 | 2975 | 3073 | 3198 | 3317 | 3485 | 3544 | 3559 | 3607 | 3690 | 3765 | 3783 | 3771 | 3779 | 3838 | 3945 | 4118 | 3960 | 3910 | 3964 | 3943 | 3886 | 3887 | 3914 | 3886 | 3846 | 4021 | 4131 | 3919 | 3907 | 3956 | 3939 | 3952 | 3949 | 3894 | 3888 | 3907 | 4067 | 4041 | 3944 | 3966 | 3947 | 3932 | 3915 | 3919 | 3958 | 3989 | 4056 | 4225 | 4091 | 4029 | 3992 | 3979 | 3970 | 3965 | 3974 | 3942 | 3963 | 4136 | 4063 | 4015 | 3924 | 3969 | 3974 | 3859 | 3885 | 4006 | 3894 | 3810 | 4095 | 3856 | 3869 | 3831 | 3787 | 3764 | 3757 | 3742 | 3762 | 3762 | 3755 | 3690 | 3416 | 3412 | 3390 | 3418 | 3410 | 3362 | 3298 | 3325 | 3401 | 3466 | 3523 | 3490 | 3456 | 3461 | 3455 | 3438 | 3415 | 3465 | 3513 | 3487 | 3440 | 3443 | 3532 | 3537 | 3479 | 3483 | 3483 | 3497 | 3559 | 3594 | 3603 | 3603 | 3623 | 3644 | 3690 | 3710 | 3752 | 3733 | 3711 | 3752 | 3712 | 3755 | 3791 | 3804 | 3807 | 3865 | 4012 | 3955 | 3962 | 4039 | 4193 | 4160 | 4170 | 4260 | 4244 | 4309 | 4367 | 4516 | 4520 | 4589 | 4623 | 4702 | 4714 | 4698 | 4764 | 4843 | 4955 | 4975 | 4992 | 5041 | 5081 | 5107 | 5103 | 5172 | 5226 | 5155 | 5138 | 5228 | 5246 | 5148 | 5030 | 4984 | 5027 | 5038 | 4961 | 4901 | 4865 | 4855 | 4809 | 4729 | 4652 | 4554 | 4549 | 4693 | 4557 | 4537 | 4480 | 4372 | 4405 | 4322 | 4277 | 4169 | 4059 | 4068 | 4177 | 4064 | 4004 | 3977 | 3967 | 3968 | 3904 | 3890 | 3753 | 3712 | 3771 | 3714 | 3729 | 3697 | 3677 | 3629 | 3641 | 3629 | 3568 | 3509 | 3419 | 3434 | 3471 | 3396 | 3333 | 3354 | 3363 | 3334 | 3324 | 3322 | 3280 | 3217 | 3204 | 3223 | 3203 | 3094 | 3014 | 3067 | 3113 | 3079 | 3074 | 3042 | 2980 | 2990 | 2944 | 2939 | 2928 | 2920 | 2905 | 2856 | 2955 | 2926 | 2916 | 2825 | 2806 | 2804 | 2805 | 2798 | 2741 | 2677 | 2687 | 2645 | 2633 | 2684 | 2683 | 2637 | 2620 | 2601 | 2594 | 2581 | 2604 | 2611 | 2532 | 2475 | 2494 | 2533 | 2479 | 2465 | 2462 | 2418 | 2406 | 2416 | 2449 | 2399 | 2395 | 2389 | 2403 | 2429 | 2374 | 2365 | 2320 | 2249 | 2263 | 2256 | 2238 | 2247 | 2272 | 2260 | 2229 | 2239 | 2255 | 2174 | 2144 | 2138 | 2098 | 2118 | 2069 | 2026 | 2077 | 2067 | 2041 | 2042 | 2020 | 1985 | 1977 | 1976 | 1952 | 1940 | 1914 | 1908 | 1901 | 1852 | 1858 | 1850 | 1792 | 1752 | 1754 | 1734 | 1704 | 1720 | 1695 | 1682 | 1723 | 1740 | 1695 | 1634 | 1685 | 1660 | 1637 | 1586 | 1630 | 1659 | 1601 | 1613 | 1575 | 1630 | 1571 | 1569 | 1569 | 1543 | 1510 | 1570 | 1610 | 1529 | 1487 | 1493 | 1526 | 1521 | 1509 | 1493 | 1488 | 1457 | 1458 | 1468 | 1446 | 1414 | 1411 | 1429 | 1446 | 1438 | 1384 | 1389 | 1419 | 1407 | 1389 | 1387 | 1405 | 1397 | 1389 | 1401 | 1403 | 1379 | 1349 | 1357 | 1386 | 1383 | 1371 | 1372 | 1334 | 1348 | 1348 | 1358 | 1321 | 1343 | 1356 | 1325 | 1356 | 1370 | 1289 | 1340 | 1338 | 1323 | 1333 | 1346 | 1398 | 1369 | 1360 | 1360 | 1394 | 1416 | 1434 | 1421 | 1395 | 1452 | 1495 | 1502 | 1516 | 1520 | 1540 | 1652 | 1683 | 1730 | 1822 | 1863 | 1849 | 1887 | 2011 | 2152 | 2264 | 2355 | 2454 | 2529 | 2633 | 2818 | 2995 | 3134 | 3304 | 3500 | 3681 | 3893 | 4211 | 4533 | 4802 | 5138 | 5516 | 5907 | 6292 | 6707 | 7199 | 7635 | 7948 | 8559 | 9092 | 9674 | 10348 | 10905 | 11426 | 12294 | 13029 | 13570 | 14322 | 14886 | 15584 | 16481 | 17339 | 18038 | 18624 | 19503 | 20021 | 20932 | 21585 | 22158 | 22809 | 23273 | 23754 | 24174 | 24586 | 24904 | 25414 | 25691 | 25714 | 25901 | 25900 | 25960 | 26036 | 25749 | 25491 | 25201 | 24845 | 24345 | 23707 | 23265 | 22677 | 22107 | 21544 | 20555 | 19701 | 18928 | 18062 | 17306 | 16466 | 15572 | 14644 | 13770 | 12980 | 12218 | 11407 | 10599 | 9877 | 9023 | 8370 | 7779 | 7180 | 6673 | 6094 | 5659 | 5150 | 4852 | 4458 | 4034 | 3780 | 3571 | 3262 | 3075 | 2965 | 2752 | 2601 | 2473 | 2393 | 2329 | 2206 | 2158 | 2102 | 2048 | 1992 | 1956 | 1931 | 1862 | 1831 | 1814 | 1818 | 1784 | 1804 | 1828 | 1814 | 1797 | 1832 | 1786 | 1768 | 1809 | 1809 | 1805 | 1795 | 1767 | 1792 | 1819 | 1774 | 1733 | 1729 | 1739 | 1752 | 1761 | 1719 | 1734 | 1736 | 1650 | 1675 | 1746 | 1718 | 1691 | 1654 | 1636 | 1641 | 1650 | 1648 | 1680 | 1673 | 1620 | 1636 | 1691 | 1673 | 1633 | 1658 | 1685 | 1644 | 1608 | 1635 | 1639 | 1618 | 1638 | 1634 | 1601 | 1562 | 1510 | 1567 | 1582 | 1589 | 1589 | 1538 | 1585 | 1615 | 1576 | 1557 | 1529 | 1580 | 1530 | 1517 | 1580 | 1606 | 1573 | 1534 | 1471 | 1535 | 1529 | 1514 | 1522 | 1519 | 1523 | 1528 | 1521 | 1501 | 1492 | 1513 | 1534 | 1526 | 1517 | 1538 | 1576 | 1629 | 1607 | 1544 | 1548 | 1595 | 1617 | 1645 | 1654 | 1646 | 1620 | 1639 | 1693 | 1693 | 1689 | 1683 | 1686 | 1716 | 1694 | 1745 | 1769 | 1757 | 1832 | 1889 | 1822 | 1796 | 1865 | 1897 | 1876 | 1872 | 1891 | 1925 | 1919 | 1930 | 1928 | 1885 | 1925 | 1925 | 1947 | 1925 | 1933 | 2028 | 1933 | 1874 | 1935 | 1903 | 1916 | 1950 | 1977 | 1939 | 1884 | 1884 | 1920 | 1893 | 1839 | 1866 | 1916 | 1907 | 1854 | 1837 | 1857 | 1810 | 1746 | 1763 | 1775 | 1774 | 1734 | 1691 | 1682 | 1678 | 1653 | 1629 | 1588 | 1565 | 1577 | 1549 | 1526 | 1556 | 1559 | 1555 | 1585 | 1499 | 1429 | 1450 | 1446 | 1428 | 1387 | 1358 | 1311 | 1301 | 1332 | 1339 | 1314 | 1272 | 1254 | 1305 | 1238 | 1213 | 1196 | 1215 | 1301 | 1224 | 1161 | 1153 | 1149 | 1167 | 1162 | 1160 | 1137 | 1111 | 1117 | 1106 | 1090 | 1102 | 1076 | 1050 | 1044 | 1035 | 1076 | 1076 | 1053 | 1035 | 1004 | 1001 | 981 | 973 | 970 | 957 | 956 | 968 | 969 | 980 | 987 | 968 | 923 | 891 | 900 | 903 | 932 | 903 | 853 | 874 | 879 | 860 | 853 | 835 | 850 | 842 | 804 | 807 | 821 | 801 | 787 | 786 | 787 | 788 | 802 | 758 | 768 | 864 | 748 | 778 | 780 | 730 | 747 | 746 | 721 | 729 | 740 | 739 | 722 | 690 | 689 | 699 | 663 | 664 | 661 | 686 | 713 | 687 | 667 | 637 | 634 | 634 | 628 | 628 | 633 | 629 | 623 | 617 | 617 | 611 | 602 | 593 | 583 | 572 | 562 | 585 | 584 | 573 | 585 | 568 | 562 | 537 | 514 | 513 | 527 | 525 | 515 | 505 | 495 | 509 | 502 | 499 | 512 | 470 | 481 | 486 | 506 | 522 | 461 | 479 | 474 | 487 | 457 | 479 | 482 | 458 | 441 | 459 | 462 | 427 | 447 | 465 | 453 | 458 | 419 | 433 | 425 | 428 | 436 | 439 | 449 | 430 | 410 | 414 | 425 | 418 | 407 | 404 | 408 | 421 | 418 | 416 | 434 | 411 | 427 | 454 | 428 | 419 | 430 | 411 | 396 | 402 | 403 | 407 | 411 | 401 | 397 | 418 | 429 | 406 | 387 | 397 | 405 | 387 | 418 | 442 | 414 | 430 | 430 | 421 | 419 | 438 | 431 | 444 | 441 | 424 | 432 | 444 | 449 | 468 | 466 | 456 | 497 | 473 | 523 | 502 | 539 | 550 | 561 | 551 | 568 | 579 | 574 | 568 | 598 | 619 | 644 | 687 | 699 | 708 | 725 | 767 | 807 | 845 | 882 | 897 | 894 | 935 | 991 | 1024 | 1071 | 1119 | 1168 | 1192 | 1232 | 1304 | 1356 | 1381 | 1429 | 1483 | 1554 | 1604 | 1670 | 1740 | 1790 | 1838 | 1899 | 1969 | 2050 | 2112 | 2135 | 2229 | 2259 | 2309 | 2409 | 2435 | 2477 | 2559 | 2594 | 2559 | 2684 | 2790 | 2844 | 2990 | 2878 | 3011 | 2980 | 2970 | 2982 | 2985 | 3037 | 3066 | 3056 | 2994 | 2993 | 3019 | 2973 | 3002 | 3050 | 2986 | 2912 | 2859 | 2840 | 2889 | 2699 | 2695 | 2621 | 2496 | 2416 | 2370 | 2335 | 2307 | 2236 | 2166 | 2098 | 2016 | 1905 | 1780 | 1731 | 1737 | 1675 | 1587 | 1513 | 1425 | 1359 | 1256 | 1177 | 1148 | 1089 | 1062 | 1019 | 951 | 879 | 806 | 750 | 699 | 653 | 631 | 602 | 575 | 546 | 466 | 482 | 418 | 406 | 369 | 347 | 314 | 299 | 279 | 258 | 257 | 250 | 236 | 224 | 217 | 212 | 186 | 172 | 192 | 184 | 169 | 157 | 167 | 148 | 163 | 162 | 148 | 137 | 140 | 143 | 142 | 135 | 132 | 137 | 139 | 130 | 116 | 115 | 125 | 123 | 124 | 130 | 131 | 133 | 124 | 120 | 120 | 124 | 120 | 111 | 114 | 117 | 114 | 117 | 129 | 121 | 119 | 121 | 117 | 122 | 118 | 120 | 121 | 129 | 124 | 130 | 125 | 113 | 112 | 123 | 125 | 114 | 105 | 102 | 115 | 126 | 114 | 122 | 112 | 124 | 101 | 106 | 107 | 109 | 119 | 115 | 119 | 127 | 127 | 115 | 110 | 117 | 121 | 114 | 102 | 96 | 104 | 110 | 109 | 113 | 115 | 123 | 120 | 102 | 98 | 94 | 97 | 105 | 99 | 94 | 95 | 96 | 102 | 110 | 112 | 107 | 107 | 103 | 102 | 100 | 107 | 109 | 102 | 104 | 110 | 102 | 108 | 108 | 100 | 100 | 101 | 98 | 97 | 96 | 99 | 109 | 96 | 91 | 111 | 101 | 91 | 95 | 103 | 108 | 89 | 90 | 93 | 96 | 91 | 86 | 91 | 93 | 93 | 99 | 97 | 96 | 99 | 86 | 94 | 103 | 97 | 85 | 83 | 84 | 88 | 88 | 88 | 89 | 89 | 82 | 80 | 83 | 90 | 86 | 76 | 78 | 81 | 80 | 82 | 75 | 72 | 73 | 67 | 80 | 81 | 82 | 81 | 75 | 75 | 75 | 73 | 82 | 86 | 65 | 58 | 65 | 76 | 71 | 81 | 58 | 66 | 73 | 79 | 83 | 72 | 57 | 58 | 63 | 72 | 75 | 66 | 60 | 59 | 64 | 64 | 57 | 60 | 72 | 65 | 56 | 52 | 65 | 72 | 63 | 57 | 54 | 52 | 54 | 56 | 56 | 58 | 65 | 67 | 62 | 56 | 57 | 63 | 65 | 64 | 58 | 55 | 60 | 58 | 58 | 57 | 59 | 65 | 61 | 54 | 57 | 56 | 52 | 48 | 44 | 60 | 58 | 57 | 50 | 48 | 42 | 54 | 54 | 46 | 49 | 50 | 52 | 46 | 42 | 46 | 50 | 53 | 49 | 51 | 54 | 43 | 46 | 46 | 50 | 44 | 53 | 51 | 46 | 48 | 51 | 54 | 54 | 50 | 49 | 45 | 49 | 53 | 49 | 46 | 43 | 41 | 47 | 55 | 48 | 40 | 40 | 43 | 41 | 45 | 51 | 46 | 41 | 43 | 43 | 41 | 41 | 39 | 37 | 40 | 39 | 36 | 47 | 39 | 50 | 46 | 34 | 37 | 44 | 44 | 46 | 42 | 36 | 38 | 40 | 37 | 34 | 46 | 54 | 48 | 36 | 46 | 45 | 41 | 35 | 45 | 48 | 43 | 44 | 46 | 56 | 54 | 43 | 35 | 36 | 40 | 42 | 44 | 44 | 45 | 41 | 39 | 44 | 45 | 50 | 51 | 42 | 37 | 41 | 43 | 42 | 45 | 45 | 47 | 46 | 42 | 43 | 49 | 47 | 53 | 51 | 44 | 53 | 60 | 38 | 39 | 43 | 46 | 38 | 36 | 55 | 48 | 44 | 53 | 57 | 61 | 39 | 48 | 52 | 42 | 60 | 50 | 47 | 47 | 42 | 55 | 50 | 54 | 64 | 50 | 54 | 38 | 43 | 69 | 67 | 75 | 61 | 54 | 54 | 65 | 53 | 58 | 65 | 56 | 53 | 61 | 60 | 60 | 46 | 47 | 52 | 57 | 58 | 56 | 63 | 62 | 61 | 53 | 53 | 59 | 56 | 55 | 55 | 69 | 65 | 61 | 57 | 51 | 47 | 48 | 59 | 72 | 63 | 72 | 65 | 57 | 58 | 55 | 51 | 50 | 57 | 61 | 64 | 58 | 62 | 56 | 58 | 48 | 48 | 45 | 46 | 38 | 33 | 42 | 37 | 47 | 61 | 47 | 44 | 46 | 35 | 34 | 41 | 43 | 39 | 36 | 35 | 35 | 32 | 28 | 29 | 26 | 21 | 25 | 29 | 28 | 26 | 26 | 27 | 25 | 22 | 29 | 25 | 17 | 24 | 24 | 20 | 18 | 21 | 22 | 17 | 23 | 22 | 23 | 19 | 24 | 18 | 20 | 19 | 18 | 12 | 16 | 19 | 17 | 19 | 14 | 16 | 8 | 10 | 18 | 22 | 5 | 15 | 11 | 22 | 14 | 11 | 10 | 9 | 16 | 15 | 15 | 10 | 11 | 7 | 16 | 13 | 12 | 13 | 10 | 9 | 14 | 10 | 11 | 13 | 11 | 11 | 11 | 10 | 8 | 9 | 13 | 13 | 14 | 14 | 8 | 8 | 14 | 12 | 6 | 10 | 9 | 8 | 11 | 11 | 7 | 10 | 10 | 15 | 14 | 8 | 6 | 7 | 9 | 14 | 14 | 12 | 9 | 13 | 7 | 17 | 11 | 5 | 10 | 20 | 7 | 11 | 16 | 10 | 9 | 10 | 11 | 9 | 10 | 11 | 10 | 11 | 8 | 12 | 9 | 17 | 10 | 11 | 8 | 7 | 6 | 10 | 10 | 10 | 8 | 10 | 12 | 13 | 12 | 8 | 8 | 11 | 13 | 12 | 8 | 6 | 4 | 8 | 11 | 10 | 11 | 14 | 16 | 14 | 7 | 9 | 12 | 9 | 10 | 9 | 12 | 12 | 11 | 9 | 9 | 12 | 8 | 11 | 7 | 9 | 9 | 4 | 8 | 6 | 10 | 10 | 6 | 6 | 7 | 11 | 7 | 6 | 13 | 10 | 12 | 2 | 6 | 5 | 7 | 7 | 6 | 7 | 7 | 9 | 10 | 8 | 8 | 7 | 7 | 7 | 5 | 5 | 4 | 5 | 7 | 5 | 6 | 6 | 5 | 3 | 6 | 7 | 5 | 7 | 7 | 8 | 6 | 3 | 3 | 6 | 3 | 4 | 4 | 3 | 8 | 5 | 7 | 5 | 11 | 3 | 5 | 7 | 3 | 7 | 6 | 4 | 7 | 6 | 3 | 4 | 5 | 7 | 6 | 3 | 1 | 5 | 6 | 7 | 4 | 3 | 5 | 4 | 3 | 4 | 5 | 4 | 9 | 6 | 2 | 4 | 6 | 5 | 4 | 7 | 5 | 3 | 4 | 5 | 4 | 4 | 4 | 4 | 5 | 4 | 4 | 6 | 5 | 3 | 2 | 3 | 4 | 3 | 3 | 5 | 2 | 3 | 3 | 5 | 2 | 3 | 1 | 4 | 4 | 1 | 3 | 4 | 5 | 3 | 1 | 7 | 6 | 1 | 2 | 1 | 6 | 5 | 4 | 4 | 5 | 5 | 7 | 3 | 3 | 6 | 3 | 3 | 4 | 2 | 5 | 1 | 2 | 5 | 5 | 3 | 2 | 3 | 5 | 4 | 2 | 2 | 3 | 4 | 5 | 6 | 4 | 4 | 3 | 3 | 5 | 6 | 6 | 6 | 4 | 2 | 1 | 3 | 3 | 5 | 4 | 5 | 4 | 4 | 2 | 7 | 5 | 6 | 5 | 2 | 4 | 3 | 2 | 2 | 3 | 1 | 2 | 1 | 5 | 3 | 2 | 2 | 3 | 2 | 5 | 3 | 4 | 4 | 3 | 3 | 5 | 5 | 5 | 2 | 0 | 2 | 5 | 4 | 2 | 4 | 4 | 4 | 6 | 2 | 1 | 2 | 1 | 2 | 3 | 1 | 1 | 4 | 3 | 3 | 2 | 2 | 4 | 5 | 5 | 4 | 3 | 1 | 2 | 6 | 4 | 4 | 3 | 4 | 4 | 2 | 4 | 3 | 6 | 4 | 5 | 6 | 2 | 4 | 0 | 5 | 2 | 2 | 1 | 2 | 2 | 0 | 5 | 2 | 4 | 7 | 5 | 4 | 6 | 4 | 4 | 4 | 2 | 2 | 5 | 2 | 4 | 2 | 5 | 7 | 4 | 3 | 5 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 5 | 2906 | 242 | 247 | 239 | 272 | 307 | 323 | 332 | 358 | 402 | 425 | 478 | 521 | 543 | 588 | 649 | 712 | 755 | 866 | 970 | 993 | 1046 | 1152 | 1274 | 1385 | 1473 | 1604 | 1737 | 1833 | 1976 | 2184 | 2264 | 2411 | 2582 | 2700 | 2807 | 2986 | 3077 | 3305 | 3370 | 3533 | 3507 | 3540 | 3669 | 3683 | 3781 | 3847 | 3852 | 3831 | 3724 | 3735 | 3975 | 3854 | 3912 | 3922 | 3891 | 3874 | 3845 | 3845 | 3882 | 3928 | 4045 | 4111 | 3979 | 3967 | 3894 | 3890 | 3901 | 3930 | 3971 | 3986 | 3966 | 4017 | 4009 | 3959 | 3918 | 3892 | 3904 | 3876 | 3887 | 3860 | 3946 | 4120 | 4190 | 3981 | 3975 | 4025 | 4051 | 4051 | 3994 | 4012 | 3999 | 3994 | 4080 | 4211 | 4060 | 3961 | 3881 | 3890 | 3901 | 3964 | 3993 | 3845 | 3780 | 4045 | 3805 | 3841 | 3848 | 3819 | 3818 | 3762 | 3754 | 3786 | 3818 | 3775 | 3695 | 3378 | 3408 | 3435 | 3424 | 3451 | 3457 | 3456 | 3429 | 3436 | 3419 | 3447 | 3477 | 3480 | 3452 | 3477 | 3438 | 3413 | 3468 | 3479 | 3477 | 3548 | 3606 | 3604 | 3560 | 3551 | 3553 | 3547 | 3547 | 3497 | 3491 | 3524 | 3544 | 3579 | 3574 | 3630 | 3641 | 3654 | 3658 | 3595 | 3629 | 3751 | 3785 | 3846 | 3870 | 3900 | 3877 | 3883 | 4013 | 4036 | 4019 | 4118 | 4149 | 4138 | 4280 | 4275 | 4323 | 4396 | 4554 | 4594 | 4640 | 4698 | 4733 | 4806 | 4834 | 4867 | 4917 | 4999 | 5001 | 5054 | 5080 | 5049 | 5076 | 5127 | 5144 | 5134 | 5095 | 5086 | 5116 | 5081 | 5022 | 5049 | 5094 | 5063 | 4940 | 4863 | 4884 | 4862 | 4845 | 4833 | 4660 | 4572 | 4570 | 4550 | 4548 | 4485 | 4464 | 4388 | 4366 | 4318 | 4299 | 4347 | 4285 | 4227 | 4082 | 4024 | 4011 | 4062 | 3913 | 3993 | 3952 | 3847 | 3834 | 3858 | 3827 | 3868 | 3757 | 3625 | 3659 | 3676 | 3645 | 3522 | 3496 | 3600 | 3580 | 3546 | 3514 | 3432 | 3372 | 3323 | 3305 | 3305 | 3293 | 3283 | 3272 | 3261 | 3215 | 3184 | 3160 | 3137 | 3168 | 3201 | 3156 | 3066 | 3040 | 3030 | 3049 | 3029 | 2986 | 2988 | 2920 | 2836 | 2899 | 2910 | 2936 | 2898 | 2830 | 2802 | 2733 | 2749 | 2728 | 2739 | 2761 | 2714 | 2726 | 2783 | 2762 | 2714 | 2651 | 2667 | 2602 | 2603 | 2675 | 2639 | 2634 | 2603 | 2536 | 2538 | 2598 | 2550 | 2494 | 2505 | 2547 | 2517 | 2473 | 2450 | 2456 | 2453 | 2414 | 2413 | 2450 | 2420 | 2345 | 2347 | 2359 | 2291 | 2323 | 2310 | 2256 | 2317 | 2293 | 2281 | 2249 | 2179 | 2213 | 2204 | 2163 | 2185 | 2185 | 2164 | 2162 | 2180 | 2141 | 2088 | 2096 | 2080 | 2025 | 1998 | 2015 | 2010 | 1960 | 1937 | 1946 | 1949 | 1903 | 1863 | 1864 | 1870 | 1838 | 1816 | 1796 | 1785 | 1754 | 1712 | 1744 | 1743 | 1702 | 1694 | 1660 | 1620 | 1670 | 1628 | 1612 | 1660 | 1627 | 1602 | 1544 | 1527 | 1567 | 1592 | 1598 | 1549 | 1592 | 1529 | 1525 | 1470 | 1460 | 1516 | 1476 | 1463 | 1503 | 1482 | 1446 | 1458 | 1415 | 1405 | 1431 | 1430 | 1458 | 1485 | 1478 | 1435 | 1415 | 1416 | 1409 | 1402 | 1388 | 1378 | 1394 | 1429 | 1448 | 1429 | 1400 | 1403 | 1407 | 1381 | 1372 | 1390 | 1389 | 1401 | 1387 | 1343 | 1347 | 1352 | 1365 | 1356 | 1351 | 1381 | 1378 | 1332 | 1382 | 1339 | 1385 | 1356 | 1402 | 1334 | 1321 | 1355 | 1417 | 1409 | 1318 | 1436 | 1406 | 1358 | 1415 | 1423 | 1419 | 1474 | 1467 | 1477 | 1543 | 1519 | 1541 | 1566 | 1645 | 1695 | 1731 | 1811 | 1882 | 1922 | 1932 | 1987 | 2088 | 2215 | 2356 | 2443 | 2554 | 2693 | 2805 | 2920 | 3064 | 3267 | 3562 | 3821 | 4005 | 4211 | 4486 | 4730 | 5052 | 5373 | 5792 | 6246 | 6663 | 7101 | 7511 | 8014 | 8585 | 9034 | 9719 | 10334 | 10910 | 11458 | 12174 | 12713 | 13541 | 14206 | 15016 | 15641 | 16298 | 17074 | 18131 | 18520 | 19522 | 20160 | 20789 | 21367 | 22020 | 22792 | 23190 | 23790 | 24461 | 24821 | 25118 | 25357 | 25556 | 25762 | 25857 | 25825 | 25936 | 25945 | 25635 | 25447 | 25267 | 24859 | 24251 | 23619 | 23147 | 22722 | 22204 | 21455 | 20588 | 19743 | 18970 | 18026 | 17185 | 16306 | 15378 | 14664 | 13830 | 12874 | 12057 | 11277 | 10618 | 9923 | 9114 | 8301 | 7725 | 7246 | 6713 | 6148 | 5649 | 5165 | 4766 | 4453 | 4066 | 3720 | 3481 | 3270 | 3087 | 2918 | 2735 | 2577 | 2483 | 2391 | 2302 | 2225 | 2144 | 2095 | 2027 | 1967 | 1871 | 1848 | 1870 | 1858 | 1856 | 1841 | 1774 | 1794 | 1845 | 1828 | 1830 | 1799 | 1835 | 1804 | 1793 | 1716 | 1742 | 1764 | 1753 | 1772 | 1723 | 1696 | 1742 | 1775 | 1761 | 1765 | 1722 | 1711 | 1708 | 1719 | 1713 | 1683 | 1702 | 1696 | 1684 | 1712 | 1721 | 1664 | 1648 | 1645 | 1633 | 1627 | 1669 | 1671 | 1641 | 1627 | 1628 | 1633 | 1596 | 1584 | 1623 | 1629 | 1577 | 1512 | 1525 | 1574 | 1565 | 1518 | 1521 | 1535 | 1511 | 1523 | 1569 | 1588 | 1644 | 1567 | 1573 | 1582 | 1560 | 1578 | 1524 | 1533 | 1519 | 1565 | 1491 | 1417 | 1468 | 1565 | 1542 | 1498 | 1506 | 1478 | 1500 | 1522 | 1506 | 1485 | 1502 | 1534 | 1532 | 1507 | 1484 | 1527 | 1558 | 1588 | 1571 | 1553 | 1613 | 1626 | 1632 | 1650 | 1641 | 1618 | 1647 | 1683 | 1699 | 1710 | 1704 | 1702 | 1724 | 1706 | 1738 | 1711 | 1785 | 1825 | 1771 | 1801 | 1805 | 1875 | 1846 | 1803 | 1846 | 1896 | 1867 | 1840 | 1857 | 1881 | 1942 | 1984 | 1939 | 1931 | 1910 | 1939 | 1904 | 1973 | 1959 | 1874 | 1909 | 1882 | 1928 | 1965 | 1962 | 1926 | 1906 | 1909 | 1871 | 1843 | 1865 | 1855 | 1845 | 1874 | 1860 | 1795 | 1858 | 1861 | 1770 | 1789 | 1763 | 1733 | 1710 | 1655 | 1612 | 1606 | 1607 | 1618 | 1633 | 1588 | 1549 | 1528 | 1547 | 1561 | 1528 | 1552 | 1565 | 1493 | 1404 | 1380 | 1360 | 1395 | 1447 | 1403 | 1368 | 1373 | 1344 | 1287 | 1250 | 1251 | 1270 | 1247 | 1246 | 1217 | 1343 | 1310 | 1313 | 1179 | 1197 | 1176 | 1168 | 1160 | 1144 | 1141 | 1120 | 1090 | 1101 | 1107 | 1065 | 1063 | 1090 | 1039 | 1030 | 1070 | 1035 | 1042 | 1078 | 1076 | 1022 | 1010 | 1022 | 991 | 979 | 974 | 966 | 942 | 936 | 948 | 949 | 941 | 911 | 874 | 904 | 887 | 889 | 938 | 885 | 840 | 859 | 877 | 851 | 845 | 822 | 813 | 837 | 840 | 809 | 791 | 790 | 787 | 798 | 778 | 769 | 782 | 787 | 779 | 777 | 746 | 749 | 722 | 749 | 744 | 740 | 729 | 723 | 711 | 683 | 679 | 699 | 699 | 680 | 667 | 649 | 674 | 717 | 704 | 667 | 667 | 664 | 631 | 614 | 626 | 652 | 647 | 630 | 622 | 617 | 609 | 585 | 574 | 589 | 583 | 576 | 581 | 589 | 558 | 550 | 563 | 548 | 534 | 532 | 522 | 526 | 524 | 524 | 529 | 498 | 488 | 510 | 519 | 496 | 485 | 471 | 507 | 497 | 482 | 485 | 467 | 459 | 473 | 490 | 487 | 477 | 464 | 458 | 462 | 457 | 455 | 446 | 448 | 459 | 414 | 458 | 500 | 436 | 429 | 430 | 433 | 438 | 431 | 412 | 408 | 423 | 426 | 432 | 423 | 403 | 415 | 426 | 410 | 403 | 398 | 421 | 440 | 415 | 418 | 425 | 416 | 394 | 409 | 431 | 416 | 390 | 391 | 413 | 420 | 414 | 417 | 410 | 393 | 397 | 406 | 413 | 433 | 405 | 417 | 436 | 418 | 415 | 434 | 444 | 444 | 452 | 451 | 440 | 448 | 455 | 440 | 469 | 492 | 488 | 450 | 520 | 475 | 490 | 510 | 531 | 540 | 557 | 566 | 566 | 575 | 608 | 622 | 626 | 660 | 694 | 716 | 738 | 763 | 783 | 832 | 881 | 922 | 932 | 934 | 955 | 1006 | 1056 | 1098 | 1136 | 1167 | 1217 | 1301 | 1344 | 1335 | 1404 | 1478 | 1546 | 1591 | 1665 | 1775 | 1790 | 1846 | 1915 | 2006 | 2015 | 2037 | 2080 | 2160 | 2232 | 2292 | 2374 | 2422 | 2440 | 2532 | 2628 | 2629 | 2638 | 2680 | 2747 | 2990 | 2837 | 2895 | 2872 | 2894 | 2947 | 2977 | 3020 | 3003 | 3000 | 3015 | 2966 | 2918 | 2937 | 2942 | 2959 | 2954 | 2920 | 2838 | 2770 | 2800 | 2658 | 2607 | 2554 | 2486 | 2414 | 2421 | 2362 | 2262 | 2212 | 2133 | 2029 | 1964 | 1927 | 1877 | 1810 | 1719 | 1622 | 1534 | 1505 | 1444 | 1327 | 1240 | 1172 | 1092 | 1037 | 994 | 965 | 928 | 858 | 793 | 739 | 692 | 664 | 616 | 569 | 521 | 497 | 462 | 458 | 423 | 370 | 353 | 347 | 325 | 296 | 290 | 268 | 259 | 244 | 213 | 207 | 205 | 185 | 179 | 155 | 171 | 188 | 152 | 171 | 162 | 161 | 146 | 136 | 134 | 131 | 135 | 143 | 141 | 131 | 125 | 132 | 134 | 130 | 131 | 126 | 125 | 124 | 122 | 120 | 123 | 122 | 114 | 113 | 116 | 118 | 121 | 126 | 129 | 123 | 117 | 125 | 119 | 100 | 103 | 115 | 121 | 122 | 114 | 108 | 115 | 128 | 118 | 100 | 109 | 115 | 126 | 124 | 117 | 121 | 119 | 117 | 124 | 119 | 122 | 109 | 98 | 119 | 104 | 99 | 110 | 107 | 124 | 113 | 111 | 113 | 119 | 113 | 105 | 104 | 101 | 91 | 91 | 107 | 119 | 112 | 106 | 106 | 113 | 107 | 101 | 102 | 106 | 109 | 106 | 102 | 104 | 104 | 98 | 97 | 98 | 102 | 108 | 110 | 101 | 101 | 110 | 111 | 94 | 83 | 99 | 100 | 92 | 85 | 94 | 105 | 106 | 104 | 103 | 100 | 101 | 106 | 111 | 102 | 99 | 99 | 99 | 88 | 105 | 82 | 78 | 95 | 94 | 90 | 90 | 99 | 94 | 88 | 91 | 88 | 83 | 84 | 91 | 87 | 80 | 75 | 90 | 96 | 86 | 88 | 93 | 89 | 86 | 89 | 83 | 72 | 72 | 75 | 81 | 87 | 82 | 77 | 77 | 80 | 77 | 73 | 78 | 87 | 79 | 72 | 68 | 71 | 72 | 74 | 79 | 83 | 81 | 73 | 71 | 72 | 70 | 69 | 71 | 69 | 62 | 67 | 75 | 64 | 67 | 68 | 65 | 61 | 70 | 75 | 67 | 64 | 71 | 78 | 65 | 42 | 51 | 70 | 68 | 67 | 61 | 66 | 59 | 65 | 62 | 55 | 57 | 59 | 59 | 62 | 68 | 69 | 60 | 53 | 52 | 48 | 45 | 44 | 45 | 47 | 52 | 61 | 58 | 58 | 59 | 67 | 63 | 51 | 51 | 63 | 68 | 62 | 59 | 54 | 58 | 53 | 45 | 58 | 53 | 51 | 51 | 43 | 56 | 53 | 43 | 48 | 45 | 49 | 53 | 47 | 47 | 52 | 52 | 55 | 49 | 51 | 50 | 46 | 37 | 40 | 44 | 48 | 43 | 55 | 59 | 49 | 43 | 44 | 52 | 53 | 47 | 39 | 35 | 38 | 40 | 46 | 47 | 37 | 43 | 46 | 45 | 44 | 44 | 47 | 46 | 39 | 43 | 48 | 48 | 42 | 40 | 43 | 44 | 46 | 50 | 44 | 46 | 57 | 39 | 42 | 44 | 43 | 40 | 37 | 36 | 44 | 48 | 40 | 33 | 42 | 49 | 43 | 41 | 39 | 36 | 41 | 46 | 47 | 45 | 51 | 49 | 38 | 51 | 48 | 36 | 42 | 43 | 38 | 42 | 45 | 44 | 44 | 38 | 35 | 36 | 38 | 38 | 36 | 40 | 45 | 39 | 43 | 49 | 51 | 53 | 44 | 39 | 46 | 48 | 48 | 49 | 48 | 43 | 42 | 42 | 42 | 47 | 47 | 44 | 49 | 50 | 44 | 48 | 53 | 44 | 43 | 39 | 42 | 48 | 45 | 47 | 45 | 45 | 47 | 56 | 52 | 49 | 45 | 46 | 48 | 59 | 56 | 56 | 49 | 55 | 41 | 51 | 66 | 57 | 68 | 54 | 57 | 61 | 52 | 55 | 50 | 54 | 61 | 54 | 59 | 61 | 70 | 64 | 56 | 61 | 64 | 67 | 57 | 55 | 61 | 66 | 66 | 67 | 53 | 49 | 55 | 58 | 51 | 58 | 60 | 63 | 60 | 65 | 57 | 45 | 54 | 54 | 60 | 59 | 59 | 75 | 71 | 58 | 54 | 50 | 56 | 64 | 63 | 65 | 51 | 51 | 61 | 54 | 41 | 46 | 52 | 47 | 54 | 52 | 47 | 42 | 51 | 47 | 51 | 38 | 45 | 42 | 46 | 45 | 57 | 42 | 37 | 36 | 35 | 35 | 34 | 30 | 28 | 33 | 31 | 26 | 19 | 30 | 27 | 20 | 26 | 30 | 28 | 24 | 23 | 29 | 25 | 20 | 20 | 17 | 15 | 19 | 20 | 14 | 13 | 18 | 19 | 18 | 14 | 17 | 18 | 17 | 19 | 18 | 15 | 19 | 17 | 21 | 15 | 14 | 12 | 9 | 9 | 12 | 13 | 16 | 25 | 14 | 16 | 12 | 10 | 15 | 15 | 11 | 11 | 9 | 12 | 10 | 10 | 9 | 10 | 8 | 12 | 6 | 8 | 11 | 12 | 13 | 19 | 13 | 14 | 13 | 14 | 13 | 10 | 13 | 14 | 12 | 9 | 10 | 11 | 10 | 9 | 11 | 10 | 13 | 13 | 11 | 8 | 7 | 8 | 11 | 9 | 10 | 7 | 5 | 8 | 11 | 9 | 10 | 9 | 11 | 13 | 9 | 8 | 13 | 13 | 11 | 14 | 9 | 7 | 11 | 6 | 9 | 10 | 8 | 14 | 12 | 9 | 10 | 7 | 7 | 10 | 11 | 8 | 9 | 9 | 9 | 9 | 11 | 13 | 11 | 10 | 8 | 11 | 13 | 12 | 10 | 10 | 12 | 14 | 10 | 8 | 8 | 8 | 7 | 9 | 11 | 8 | 9 | 11 | 8 | 9 | 11 | 11 | 17 | 12 | 10 | 8 | 6 | 10 | 11 | 9 | 6 | 7 | 6 | 8 | 12 | 9 | 9 | 4 | 8 | 7 | 5 | 6 | 10 | 10 | 7 | 6 | 8 | 4 | 8 | 7 | 9 | 13 | 10 | 13 | 12 | 5 | 7 | 11 | 4 | 7 | 8 | 11 | 10 | 5 | 3 | 4 | 5 | 7 | 8 | 7 | 7 | 6 | 4 | 3 | 5 | 7 | 6 | 6 | 8 | 10 | 8 | 4 | 4 | 6 | 6 | 5 | 4 | 5 | 5 | 7 | 5 | 4 | 5 | 5 | 7 | 9 | 6 | 5 | 5 | 5 | 5 | 3 | 8 | 5 | 5 | 4 | 9 | 3 | 3 | 3 | 8 | 5 | 6 | 4 | 6 | 7 | 2 | 2 | 3 | 1 | 3 | 3 | 4 | 4 | 5 | 6 | 4 | 6 | 3 | 2 | 5 | 3 | 2 | 4 | 5 | 6 | 8 | 7 | 4 | 4 | 4 | 3 | 3 | 4 | 5 | 3 | 3 | 3 | 6 | 5 | 2 | 3 | 4 | 6 | 7 | 3 | 2 | 4 | 3 | 6 | 4 | 4 | 6 | 3 | 3 | 5 | 6 | 4 | 3 | 4 | 2 | 2 | 1 | 3 | 4 | 5 | 2 | 4 | 1 | 6 | 9 | 4 | 6 | 3 | 5 | 5 | 2 | 4 | 4 | 6 | 6 | 3 | 2 | 2 | 3 | 4 | 4 | 1 | 2 | 3 | 2 | 2 | 4 | 3 | 1 | 3 | 3 | 5 | 4 | 3 | 3 | 2 | 1 | 2 | 3 | 2 | 2 | 1 | 3 | 5 | 1 | 3 | 5 | 5 | 4 | 2 | 2 | 1 | 2 | 0 | 4 | 2 | 3 | 2 | 3 | 2 | 4 | 4 | 5 | 2 | 3 | 1 | 4 | 3 | 5 | 3 | 6 | 2 | 3 | 4 | 3 | 3 | 5 | 5 | 6 | 3 | 2 | 4 | 2 | 2 | 4 | 4 | 6 | 7 | 3 | 2 | 2 | 2 | 4 | 3 | 2 | 4 | 3 | 3 | 2 | 2 | 3 | 5 | 7 | 7 | 6 | 5 | 5 | 4 | 3 | 5 | 8 | 8 | 7 | 2 | 3 | 2 | 4 | 3 | 3 | 1 | 4 | 4 | 1 | 7 | 3 | 4 | 1 | 2 | 4 | 2 | 7 | 3 | 4 | 3 | 4 | 3 | 5 | 4 | 4 | 3 | 4 | 2 | 2 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 6 | 422 | 39 | 35 | 39 | 47 | 43 | 49 | 50 | 59 | 59 | 87 | 102 | 122 | 126 | 142 | 184 | 213 | 249 | 294 | 342 | 386 | 440 | 518 | 613 | 660 | 760 | 863 | 972 | 1083 | 1184 | 1304 | 1395 | 1475 | 1501 | 1516 | 1558 | 1581 | 1616 | 1622 | 1620 | 1606 | 1631 | 1683 | 1738 | 1696 | 1624 | 1670 | 1677 | 1655 | 1644 | 1634 | 1648 | 1680 | 1742 | 1746 | 1676 | 1683 | 1682 | 1667 | 1679 | 1708 | 1705 | 1674 | 1623 | 1693 | 1743 | 1691 | 1670 | 1688 | 1668 | 1667 | 1702 | 1693 | 1678 | 1680 | 1762 | 1795 | 1704 | 1691 | 1743 | 1764 | 1759 | 1769 | 1760 | 1764 | 1778 | 1829 | 1785 | 1758 | 1777 | 1737 | 1700 | 1715 | 1714 | 1681 | 1696 | 1738 | 1757 | 1662 | 1629 | 1626 | 1608 | 1611 | 1659 | 1648 | 1645 | 1603 | 1669 | 1637 | 1627 | 1601 | 1580 | 1582 | 1579 | 1593 | 1584 | 1583 | 1601 | 1571 | 1418 | 1486 | 1464 | 1494 | 1492 | 1510 | 1447 | 1445 | 1454 | 1441 | 1480 | 1463 | 1441 | 1460 | 1432 | 1427 | 1448 | 1468 | 1492 | 1456 | 1484 | 1498 | 1464 | 1486 | 1468 | 1481 | 1528 | 1493 | 1513 | 1537 | 1495 | 1520 | 1527 | 1554 | 1543 | 1534 | 1541 | 1578 | 1603 | 1603 | 1633 | 1657 | 1650 | 1676 | 1762 | 1709 | 1688 | 1787 | 1812 | 1863 | 1846 | 1968 | 1967 | 2038 | 2016 | 2111 | 2130 | 2191 | 2194 | 2213 | 2227 | 2259 | 2237 | 2268 | 2260 | 2278 | 2262 | 2215 | 2207 | 2213 | 2194 | 2198 | 2141 | 2133 | 2119 | 2124 | 2102 | 2074 | 2071 | 2047 | 2058 | 2058 | 1992 | 1946 | 1953 | 1932 | 1856 | 1805 | 1801 | 1816 | 1783 | 1715 | 1727 | 1749 | 1742 | 1728 | 1684 | 1695 | 1698 | 1653 | 1632 | 1634 | 1636 | 1626 | 1597 | 1551 | 1522 | 1520 | 1483 | 1432 | 1443 | 1518 | 1486 | 1421 | 1410 | 1424 | 1384 | 1355 | 1375 | 1396 | 1397 | 1389 | 1370 | 1341 | 1355 | 1374 | 1368 | 1357 | 1330 | 1294 | 1279 | 1285 | 1283 | 1252 | 1234 | 1242 | 1250 | 1246 | 1210 | 1171 | 1162 | 1186 | 1199 | 1184 | 1151 | 1168 | 1204 | 1167 | 1123 | 1138 | 1138 | 1119 | 1081 | 1059 | 1070 | 1112 | 1094 | 1080 | 1056 | 1010 | 1020 | 1055 | 1060 | 1041 | 1024 | 1010 | 1020 | 972 | 997 | 1004 | 989 | 962 | 957 | 954 | 931 | 928 | 960 | 967 | 945 | 925 | 923 | 943 | 934 | 907 | 900 | 907 | 927 | 923 | 890 | 900 | 907 | 902 | 907 | 881 | 846 | 849 | 872 | 836 | 821 | 813 | 806 | 761 | 850 | 851 | 819 | 789 | 783 | 798 | 762 | 744 | 786 | 731 | 765 | 704 | 755 | 740 | 702 | 703 | 710 | 699 | 696 | 691 | 683 | 685 | 689 | 690 | 669 | 687 | 666 | 661 | 644 | 639 | 667 | 669 | 642 | 622 | 634 | 637 | 606 | 596 | 622 | 633 | 639 | 613 | 614 | 613 | 617 | 624 | 580 | 600 | 602 | 618 | 600 | 615 | 590 | 597 | 619 | 612 | 580 | 566 | 586 | 555 | 570 | 568 | 584 | 574 | 572 | 577 | 557 | 519 | 526 | 533 | 548 | 548 | 542 | 541 | 553 | 551 | 542 | 530 | 516 | 511 | 501 | 511 | 520 | 517 | 517 | 499 | 491 | 498 | 501 | 503 | 507 | 525 | 534 | 518 | 522 | 542 | 536 | 527 | 511 | 502 | 515 | 500 | 501 | 523 | 516 | 516 | 508 | 494 | 472 | 485 | 522 | 549 | 542 | 513 | 508 | 526 | 538 | 548 | 559 | 580 | 615 | 632 | 634 | 635 | 638 | 686 | 732 | 737 | 766 | 816 | 873 | 920 | 986 | 1082 | 1172 | 1256 | 1358 | 1478 | 1597 | 1725 | 1893 | 2104 | 2276 | 2492 | 2707 | 2918 | 3169 | 3447 | 3781 | 4150 | 4402 | 4703 | 5030 | 5397 | 5778 | 6209 | 6623 | 7143 | 7522 | 7968 | 8616 | 9053 | 9361 | 9858 | 10376 | 10717 | 11035 | 11224 | 11672 | 12105 | 12369 | 12706 | 12725 | 12986 | 13084 | 13219 | 13285 | 13276 | 13195 | 12940 | 12805 | 12701 | 12490 | 12160 | 11897 | 11525 | 11011 | 10598 | 10236 | 9719 | 9167 | 8741 | 8186 | 7642 | 7244 | 6730 | 6166 | 5765 | 5385 | 4970 | 4561 | 4189 | 3755 | 3432 | 3171 | 2966 | 2760 | 2441 | 2254 | 2026 | 1898 | 1679 | 1517 | 1489 | 1297 | 1327 | 1290 | 1161 | 1149 | 1035 | 1029 | 988 | 968 | 956 | 982 | 941 | 887 | 870 | 908 | 867 | 877 | 882 | 864 | 854 | 874 | 857 | 832 | 840 | 849 | 895 | 905 | 870 | 848 | 840 | 869 | 842 | 823 | 840 | 842 | 855 | 852 | 840 | 852 | 827 | 820 | 825 | 843 | 823 | 818 | 802 | 778 | 809 | 827 | 837 | 801 | 810 | 816 | 794 | 803 | 822 | 801 | 766 | 786 | 779 | 760 | 786 | 780 | 780 | 765 | 743 | 771 | 774 | 749 | 757 | 772 | 772 | 768 | 761 | 745 | 731 | 731 | 742 | 739 | 723 | 720 | 726 | 724 | 723 | 733 | 728 | 712 | 710 | 714 | 712 | 727 | 730 | 739 | 762 | 757 | 743 | 710 | 710 | 750 | 761 | 758 | 739 | 742 | 744 | 746 | 766 | 763 | 772 | 770 | 724 | 691 | 721 | 766 | 772 | 770 | 788 | 785 | 767 | 756 | 774 | 780 | 802 | 805 | 832 | 818 | 813 | 836 | 815 | 810 | 848 | 886 | 864 | 843 | 854 | 870 | 901 | 899 | 893 | 892 | 902 | 910 | 945 | 1011 | 910 | 965 | 959 | 943 | 970 | 983 | 947 | 941 | 953 | 949 | 959 | 978 | 981 | 982 | 956 | 927 | 885 | 896 | 920 | 961 | 937 | 877 | 887 | 858 | 844 | 877 | 858 | 834 | 855 | 849 | 821 | 813 | 801 | 810 | 766 | 762 | 783 | 692 | 774 | 745 | 687 | 716 | 722 | 695 | 655 | 716 | 698 | 686 | 683 | 662 | 650 | 637 | 605 | 597 | 611 | 626 | 591 | 582 | 597 | 605 | 550 | 549 | 557 | 577 | 541 | 548 | 545 | 546 | 562 | 548 | 507 | 497 | 509 | 493 | 485 | 489 | 491 | 485 | 490 | 501 | 479 | 471 | 491 | 488 | 469 | 480 | 459 | 444 | 435 | 429 | 426 | 433 | 441 | 432 | 422 | 424 | 435 | 443 | 421 | 399 | 378 | 385 | 390 | 411 | 385 | 373 | 378 | 375 | 372 | 360 | 360 | 374 | 376 | 368 | 366 | 363 | 357 | 348 | 334 | 340 | 335 | 337 | 338 | 330 | 331 | 343 | 334 | 321 | 313 | 308 | 302 | 296 | 285 | 293 | 295 | 290 | 282 | 284 | 290 | 291 | 279 | 268 | 273 | 268 | 261 | 273 | 275 | 267 | 263 | 250 | 246 | 263 | 267 | 244 | 236 | 229 | 231 | 248 | 240 | 225 | 226 | 226 | 229 | 237 | 230 | 226 | 217 | 203 | 200 | 210 | 212 | 203 | 206 | 213 | 213 | 194 | 175 | 200 | 205 | 192 | 201 | 201 | 180 | 186 | 180 | 178 | 175 | 174 | 177 | 183 | 186 | 171 | 177 | 187 | 179 | 172 | 160 | 182 | 176 | 186 | 192 | 165 | 147 | 163 | 165 | 162 | 164 | 162 | 156 | 152 | 161 | 167 | 141 | 139 | 142 | 155 | 144 | 180 | 178 | 174 | 155 | 151 | 159 | 159 | 146 | 151 | 163 | 160 | 150 | 148 | 149 | 137 | 133 | 143 | 137 | 148 | 122 | 150 | 141 | 159 | 146 | 133 | 144 | 140 | 147 | 155 | 152 | 143 | 139 | 140 | 139 | 148 | 173 | 173 | 147 | 149 | 154 | 160 | 158 | 175 | 164 | 176 | 186 | 167 | 184 | 205 | 202 | 195 | 203 | 217 | 213 | 210 | 238 | 250 | 251 | 267 | 262 | 294 | 310 | 324 | 340 | 335 | 378 | 417 | 434 | 441 | 460 | 486 | 510 | 546 | 582 | 609 | 643 | 670 | 713 | 776 | 811 | 827 | 847 | 882 | 935 | 972 | 1028 | 1066 | 1120 | 1166 | 1222 | 1276 | 1313 | 1324 | 1373 | 1442 | 1473 | 1538 | 1611 | 1628 | 1681 | 1707 | 1721 | 1798 | 1842 | 1823 | 1847 | 1878 | 1891 | 1929 | 1926 | 1910 | 1923 | 1920 | 1877 | 1875 | 1866 | 1871 | 1857 | 1817 | 1848 | 1784 | 1704 | 1693 | 1657 | 1613 | 1584 | 1534 | 1485 | 1437 | 1402 | 1361 | 1276 | 1204 | 1168 | 1125 | 1096 | 1043 | 971 | 947 | 933 | 855 | 776 | 743 | 687 | 642 | 589 | 556 | 521 | 479 | 446 | 416 | 398 | 383 | 344 | 322 | 286 | 265 | 248 | 217 | 221 | 197 | 157 | 144 | 129 | 120 | 121 | 111 | 103 | 97 | 96 | 87 | 64 | 56 | 62 | 53 | 58 | 49 | 35 | 43 | 46 | 48 | 37 | 41 | 39 | 39 | 46 | 37 | 30 | 31 | 33 | 32 | 32 | 34 | 31 | 32 | 28 | 29 | 31 | 35 | 25 | 25 | 22 | 19 | 23 | 27 | 31 | 27 | 28 | 27 | 23 | 25 | 31 | 33 | 24 | 21 | 23 | 26 | 29 | 18 | 23 | 21 | 22 | 18 | 17 | 31 | 29 | 26 | 26 | 25 | 22 | 26 | 29 | 23 | 26 | 27 | 24 | 32 | 34 | 27 | 21 | 18 | 22 | 27 | 24 | 22 | 26 | 21 | 20 | 23 | 23 | 21 | 22 | 21 | 18 | 17 | 25 | 29 | 23 | 23 | 25 | 25 | 23 | 25 | 20 | 20 | 23 | 24 | 22 | 19 | 21 | 25 | 23 | 22 | 24 | 23 | 21 | 21 | 20 | 17 | 22 | 23 | 27 | 29 | 27 | 24 | 19 | 18 | 18 | 17 | 20 | 26 | 23 | 16 | 17 | 21 | 22 | 21 | 20 | 21 | 16 | 18 | 25 | 23 | 20 | 20 | 21 | 17 | 13 | 14 | 17 | 16 | 17 | 21 | 18 | 14 | 19 | 19 | 19 | 17 | 18 | 20 | 15 | 17 | 18 | 20 | 20 | 18 | 20 | 25 | 25 | 18 | 15 | 14 | 13 | 17 | 17 | 17 | 15 | 19 | 18 | 20 | 15 | 19 | 16 | 16 | 16 | 12 | 11 | 14 | 12 | 10 | 11 | 11 | 14 | 18 | 14 | 17 | 14 | 11 | 16 | 18 | 18 | 13 | 19 | 20 | 17 | 16 | 15 | 9 | 8 | 11 | 9 | 14 | 13 | 13 | 12 | 11 | 10 | 8 | 11 | 20 | 14 | 11 | 16 | 14 | 16 | 13 | 8 | 7 | 7 | 11 | 16 | 15 | 10 | 7 | 8 | 9 | 14 | 13 | 14 | 5 | 11 | 14 | 12 | 11 | 11 | 12 | 15 | 13 | 12 | 12 | 12 | 14 | 12 | 5 | 7 | 9 | 10 | 9 | 11 | 5 | 8 | 7 | 8 | 9 | 9 | 10 | 12 | 13 | 11 | 10 | 12 | 11 | 7 | 11 | 7 | 8 | 10 | 13 | 13 | 10 | 10 | 8 | 9 | 13 | 10 | 6 | 5 | 7 | 7 | 6 | 6 | 5 | 7 | 10 | 11 | 8 | 8 | 9 | 7 | 8 | 7 | 7 | 6 | 8 | 9 | 10 | 10 | 9 | 10 | 13 | 12 | 10 | 9 | 8 | 8 | 6 | 8 | 12 | 10 | 8 | 10 | 11 | 8 | 5 | 5 | 7 | 9 | 9 | 10 | 8 | 8 | 8 | 9 | 11 | 10 | 8 | 7 | 6 | 9 | 10 | 10 | 9 | 8 | 6 | 6 | 6 | 9 | 11 | 10 | 8 | 8 | 8 | 9 | 8 | 11 | 9 | 8 | 10 | 7 | 6 | 8 | 9 | 7 | 5 | 5 | 6 | 8 | 10 | 8 | 10 | 11 | 10 | 10 | 13 | 6 | 8 | 4 | 13 | 11 | 7 | 10 | 9 | 12 | 6 | 13 | 11 | 15 | 11 | 6 | 11 | 9 | 4 | 9 | 9 | 6 | 13 | 8 | 5 | 4 | 6 | 11 | 11 | 9 | 11 | 6 | 13 | 10 | 10 | 10 | 9 | 8 | 8 | 19 | 12 | 8 | 11 | 14 | 11 | 13 | 12 | 12 | 9 | 11 | 9 | 10 | 20 | 9 | 8 | 20 | 19 | 17 | 14 | 9 | 15 | 12 | 15 | 12 | 12 | 12 | 14 | 13 | 15 | 14 | 13 | 11 | 11 | 12 | 18 | 18 | 12 | 14 | 16 | 9 | 17 | 15 | 15 | 10 | 10 | 7 | 5 | 6 | 8 | 15 | 11 | 9 | 13 | 8 | 8 | 10 | 9 | 5 | 8 | 11 | 9 | 6 | 10 | 9 | 8 | 11 | 8 | 9 | 9 | 9 | 9 | 8 | 9 | 6 | 5 | 5 | 3 | 4 | 6 | 8 | 6 | 2 | 4 | 6 | 7 | 5 | 1 | 5 | 6 | 3 | 4 | 3 | 2 | 2 | 2 | 3 | 4 | 4 | 2 | 3 | 2 | 2 | 3 | 5 | 4 | 2 | 4 | 4 | 2 | 1 | 2 | 1 | 2 | 1 | 2 | 5 | 5 | 2 | 1 | 2 | 5 | 4 | 3 | 2 | 2 | 2 | 2 | 4 | 3 | 2 | 3 | 1 | 0 | 4 | 5 | 3 | 4 | 2 | 1 | 2 | 2 | 4 | 2 | 3 | 6 | 4 | 2 | 4 | 3 | 5 | 5 | 3 | 4 | 2 | 3 | 0 | 2 | 2 | 3 | 1 | 2 | 6 | 2 | 0 | 3 | 1 | 2 | 3 | 2 | 2 | 0 | 5 | 2 | 1 | 1 | 1 | 3 | 2 | 2 | 4 | 2 | 1 | 0 | 1 | 4 | 1 | 3 | 2 | 2 | 0 | 3 | 0 | 1 | 2 | 3 | 2 | 1 | 0 | 0 | 3 | 3 | 3 | 1 | 0 | 2 | 1 | 1 | 2 | 0 | 2 | 2 | 0 | 0 | 1 | 2 | 2 | 4 | 2 | 3 | 3 | 1 | 1 | 2 | 0 | 1 | 1 | 2 | 1 | 2 | 2 | 1 | 1 | 1 | 2 | 2 | 3 | 3 | 3 | 2 | 2 | 3 | 2 | 2 | 1 | 1 | 1 | 0 | 0 | 2 | 2 | 1 | 2 | 1 | 2 | 2 | 2 | 2 | 1 | 1 | 2 | 2 | 1 | 2 | 2 | 1 | 1 | 0 | 1 | 1 | 2 | 2 | 3 | 2 | 1 | 0 | 1 | 2 | 2 | 1 | 1 | 1 | 2 | 2 | 2 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 2 | 1 | 1 | 2 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 2 | 1 | 1 | 2 | 2 | 1 | 1 | 2 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 2 | 1 | 0 | 1 | 0 | 0 | 0 | 3 | 0 | 3 | 2 | 0 | 0 | 2 | 2 | 0 | 2 | 2 | 1 | 0 | 0 | 1 | 0 | 3 | 3 | 2 | 1 | 2 | 1 | 0 | 1 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 2 | 1 | 1 | 0 | 1 | 3 | 0 | 2 | 2 | 1 | 0 | 2 | 1 | 0 | 3 | 2 | 1 | 1 | 2 | 0 | 0 | 2 | 0 | 0 | 2 | 1 | 0 | 1 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 2 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 2 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 2 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 2 | 1 | 0 | 1 | 2 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 2 | 2 | 1 | 0 | 2 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 2 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 2 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 2 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 2 | 0 | 3 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 7 | 478 | 39 | 42 | 48 | 40 | 54 | 57 | 55 | 57 | 68 | 74 | 88 | 108 | 144 | 162 | 172 | 202 | 231 | 257 | 337 | 390 | 442 | 512 | 594 | 653 | 773 | 873 | 963 | 1076 | 1187 | 1262 | 1362 | 1450 | 1495 | 1484 | 1526 | 1567 | 1597 | 1571 | 1542 | 1565 | 1612 | 1645 | 1682 | 1637 | 1608 | 1638 | 1593 | 1562 | 1588 | 1605 | 1623 | 1641 | 1715 | 1713 | 1599 | 1593 | 1600 | 1621 | 1609 | 1593 | 1630 | 1666 | 1684 | 1684 | 1627 | 1611 | 1634 | 1633 | 1636 | 1658 | 1667 | 1621 | 1591 | 1597 | 1695 | 1730 | 1691 | 1683 | 1691 | 1680 | 1665 | 1690 | 1698 | 1698 | 1723 | 1729 | 1694 | 1660 | 1666 | 1649 | 1643 | 1628 | 1611 | 1586 | 1603 | 1641 | 1643 | 1578 | 1534 | 1566 | 1603 | 1600 | 1585 | 1569 | 1546 | 1557 | 1641 | 1581 | 1557 | 1536 | 1532 | 1538 | 1520 | 1525 | 1533 | 1545 | 1548 | 1581 | 1451 | 1409 | 1378 | 1360 | 1374 | 1450 | 1458 | 1415 | 1384 | 1422 | 1402 | 1437 | 1415 | 1411 | 1422 | 1432 | 1461 | 1442 | 1436 | 1440 | 1435 | 1447 | 1447 | 1413 | 1430 | 1456 | 1441 | 1437 | 1420 | 1446 | 1491 | 1458 | 1448 | 1486 | 1522 | 1529 | 1507 | 1527 | 1570 | 1589 | 1613 | 1580 | 1573 | 1595 | 1598 | 1681 | 1719 | 1699 | 1743 | 1864 | 1802 | 1895 | 2016 | 1970 | 1926 | 2000 | 2102 | 2088 | 2059 | 2097 | 2119 | 2209 | 2204 | 2195 | 2164 | 2178 | 2143 | 2126 | 2149 | 2129 | 2069 | 2090 | 2097 | 2060 | 2037 | 2047 | 2042 | 1993 | 1924 | 1886 | 1891 | 1891 | 1867 | 1841 | 1844 | 1863 | 1865 | 1815 | 1743 | 1735 | 1753 | 1693 | 1711 | 1693 | 1683 | 1701 | 1685 | 1634 | 1566 | 1567 | 1565 | 1560 | 1568 | 1548 | 1480 | 1484 | 1532 | 1460 | 1405 | 1445 | 1438 | 1450 | 1447 | 1451 | 1423 | 1390 | 1363 | 1343 | 1341 | 1327 | 1315 | 1315 | 1328 | 1329 | 1322 | 1310 | 1285 | 1255 | 1244 | 1251 | 1258 | 1248 | 1242 | 1239 | 1192 | 1157 | 1163 | 1150 | 1140 | 1160 | 1165 | 1144 | 1146 | 1163 | 1138 | 1102 | 1099 | 1094 | 1112 | 1110 | 1094 | 1076 | 1062 | 1026 | 987 | 1010 | 1059 | 1034 | 1015 | 1045 | 1029 | 994 | 972 | 987 | 990 | 989 | 971 | 977 | 984 | 969 | 960 | 959 | 950 | 934 | 927 | 916 | 915 | 898 | 864 | 854 | 871 | 894 | 907 | 903 | 910 | 917 | 903 | 886 | 873 | 887 | 874 | 850 | 845 | 848 | 854 | 842 | 824 | 828 | 816 | 768 | 786 | 792 | 795 | 785 | 789 | 783 | 790 | 775 | 747 | 710 | 721 | 726 | 698 | 673 | 720 | 717 | 671 | 675 | 683 | 690 | 695 | 686 | 669 | 632 | 658 | 673 | 673 | 649 | 608 | 599 | 618 | 621 | 608 | 619 | 619 | 608 | 624 | 610 | 609 | 620 | 604 | 638 | 605 | 574 | 595 | 584 | 585 | 568 | 574 | 607 | 593 | 568 | 575 | 548 | 553 | 552 | 517 | 576 | 597 | 551 | 530 | 551 | 560 | 575 | 535 | 529 | 536 | 541 | 529 | 535 | 537 | 529 | 517 | 511 | 509 | 496 | 505 | 514 | 512 | 501 | 492 | 504 | 509 | 496 | 490 | 493 | 496 | 500 | 505 | 495 | 487 | 497 | 499 | 500 | 470 | 450 | 466 | 482 | 503 | 513 | 498 | 470 | 473 | 470 | 458 | 440 | 458 | 479 | 488 | 504 | 494 | 513 | 510 | 501 | 482 | 482 | 508 | 517 | 524 | 548 | 542 | 538 | 560 | 585 | 597 | 609 | 634 | 662 | 686 | 708 | 737 | 788 | 839 | 870 | 929 | 1034 | 1141 | 1236 | 1318 | 1395 | 1517 | 1659 | 1785 | 1961 | 2176 | 2392 | 2628 | 2875 | 3095 | 3301 | 3557 | 3917 | 4246 | 4557 | 4917 | 5279 | 5695 | 6105 | 6489 | 6900 | 7308 | 7700 | 8216 | 8762 | 9080 | 9540 | 9907 | 10354 | 10871 | 11234 | 11470 | 11727 | 11997 | 12240 | 12440 | 12579 | 12698 | 12774 | 12788 | 12783 | 12672 | 12543 | 12465 | 12233 | 11939 | 11747 | 11443 | 11033 | 10738 | 10380 | 9866 | 9337 | 8846 | 8424 | 7970 | 7428 | 6869 | 6427 | 5964 | 5618 | 5277 | 4795 | 4362 | 3942 | 3647 | 3373 | 3090 | 2824 | 2583 | 2283 | 2081 | 1964 | 1835 | 1670 | 1528 | 1371 | 1378 | 1284 | 1192 | 1138 | 1044 | 1029 | 962 | 922 | 953 | 904 | 877 | 875 | 885 | 906 | 900 | 849 | 891 | 869 | 830 | 808 | 833 | 861 | 825 | 804 | 830 | 815 | 817 | 847 | 828 | 824 | 839 | 831 | 827 | 823 | 780 | 776 | 816 | 833 | 831 | 813 | 809 | 832 | 793 | 807 | 785 | 805 | 798 | 798 | 763 | 780 | 787 | 778 | 759 | 764 | 765 | 808 | 760 | 733 | 740 | 792 | 792 | 746 | 757 | 743 | 737 | 744 | 737 | 740 | 740 | 734 | 717 | 715 | 722 | 744 | 762 | 731 | 700 | 718 | 714 | 688 | 692 | 711 | 718 | 712 | 699 | 685 | 687 | 703 | 717 | 716 | 718 | 701 | 693 | 704 | 691 | 692 | 730 | 732 | 709 | 723 | 724 | 691 | 722 | 742 | 716 | 709 | 723 | 713 | 718 | 711 | 710 | 730 | 719 | 707 | 733 | 758 | 749 | 737 | 732 | 715 | 730 | 765 | 799 | 797 | 798 | 818 | 796 | 796 | 818 | 839 | 852 | 857 | 871 | 872 | 866 | 879 | 881 | 872 | 864 | 860 | 884 | 913 | 961 | 926 | 895 | 908 | 900 | 919 | 918 | 904 | 925 | 957 | 947 | 930 | 943 | 941 | 922 | 892 | 860 | 873 | 873 | 911 | 922 | 897 | 842 | 830 | 850 | 825 | 843 | 803 | 788 | 779 | 759 | 777 | 802 | 799 | 769 | 750 | 729 | 739 | 762 | 750 | 707 | 684 | 677 | 666 | 674 | 664 | 653 | 635 | 619 | 654 | 667 | 649 | 624 | 586 | 604 | 585 | 560 | 580 | 562 | 562 | 611 | 584 | 548 | 557 | 525 | 528 | 536 | 557 | 531 | 509 | 502 | 491 | 505 | 509 | 483 | 493 | 502 | 487 | 481 | 450 | 469 | 473 | 476 | 486 | 460 | 436 | 462 | 433 | 427 | 441 | 439 | 430 | 428 | 432 | 421 | 408 | 404 | 409 | 423 | 399 | 411 | 387 | 379 | 381 | 397 | 391 | 396 | 364 | 348 | 370 | 370 | 363 | 368 | 357 | 345 | 336 | 318 | 315 | 317 | 330 | 337 | 327 | 321 | 319 | 312 | 303 | 304 | 316 | 301 | 302 | 308 | 300 | 300 | 296 | 287 | 289 | 289 | 281 | 266 | 268 | 265 | 274 | 262 | 275 | 278 | 247 | 231 | 242 | 250 | 243 | 229 | 214 | 219 | 234 | 234 | 235 | 238 | 232 | 216 | 210 | 228 | 238 | 231 | 225 | 218 | 210 | 207 | 209 | 214 | 213 | 208 | 209 | 209 | 209 | 209 | 204 | 202 | 198 | 191 | 190 | 189 | 175 | 180 | 176 | 178 | 196 | 187 | 174 | 175 | 171 | 167 | 174 | 177 | 176 | 173 | 178 | 173 | 154 | 175 | 160 | 169 | 172 | 156 | 151 | 161 | 162 | 160 | 158 | 153 | 145 | 141 | 141 | 147 | 155 | 152 | 144 | 145 | 151 | 137 | 141 | 141 | 146 | 154 | 131 | 125 | 134 | 144 | 148 | 142 | 129 | 129 | 137 | 140 | 154 | 149 | 144 | 133 | 157 | 164 | 142 | 123 | 134 | 132 | 129 | 136 | 142 | 136 | 137 | 140 | 138 | 138 | 136 | 145 | 148 | 137 | 136 | 140 | 169 | 157 | 156 | 180 | 158 | 158 | 174 | 177 | 190 | 186 | 176 | 180 | 191 | 196 | 206 | 227 | 217 | 215 | 242 | 298 | 287 | 285 | 281 | 297 | 306 | 361 | 360 | 374 | 381 | 403 | 425 | 462 | 496 | 521 | 544 | 551 | 605 | 648 | 648 | 709 | 722 | 825 | 868 | 928 | 925 | 990 | 1017 | 1031 | 1100 | 1183 | 1210 | 1230 | 1278 | 1318 | 1357 | 1402 | 1441 | 1503 | 1524 | 1581 | 1621 | 1706 | 1738 | 1683 | 1761 | 1773 | 1787 | 1774 | 1796 | 1861 | 1901 | 1897 | 1887 | 1871 | 1846 | 1840 | 1855 | 1860 | 1833 | 1791 | 1772 | 1786 | 1730 | 1757 | 1698 | 1621 | 1559 | 1508 | 1489 | 1448 | 1394 | 1326 | 1270 | 1225 | 1185 | 1138 | 1084 | 1048 | 997 | 957 | 936 | 841 | 732 | 698 | 656 | 624 | 595 | 550 | 504 | 472 | 445 | 415 | 374 | 343 | 329 | 308 | 282 | 262 | 236 | 210 | 195 | 185 | 159 | 129 | 117 | 119 | 121 | 111 | 96 | 82 | 77 | 81 | 78 | 64 | 59 | 59 | 54 | 52 | 47 | 39 | 43 | 39 | 36 | 34 | 32 | 33 | 31 | 30 | 30 | 30 | 29 | 29 | 34 | 33 | 25 | 27 | 28 | 21 | 27 | 32 | 31 | 28 | 27 | 26 | 24 | 22 | 26 | 22 | 23 | 24 | 22 | 22 | 24 | 24 | 28 | 23 | 24 | 26 | 24 | 21 | 21 | 24 | 25 | 28 | 21 | 24 | 29 | 25 | 23 | 22 | 22 | 23 | 27 | 27 | 25 | 20 | 22 | 29 | 23 | 18 | 27 | 22 | 28 | 19 | 26 | 22 | 24 | 24 | 24 | 24 | 19 | 15 | 19 | 25 | 23 | 19 | 19 | 21 | 26 | 18 | 18 | 17 | 25 | 14 | 22 | 23 | 16 | 14 | 13 | 14 | 17 | 21 | 23 | 23 | 22 | 28 | 19 | 22 | 29 | 20 | 17 | 33 | 31 | 20 | 19 | 12 | 11 | 15 | 17 | 18 | 16 | 15 | 12 | 12 | 19 | 18 | 17 | 17 | 20 | 19 | 16 | 21 | 20 | 21 | 17 | 23 | 24 | 25 | 24 | 20 | 18 | 18 | 20 | 20 | 17 | 16 | 16 | 18 | 17 | 19 | 21 | 19 | 15 | 18 | 21 | 17 | 14 | 16 | 16 | 15 | 16 | 18 | 19 | 17 | 13 | 12 | 14 | 16 | 19 | 16 | 13 | 17 | 18 | 13 | 13 | 15 | 15 | 14 | 13 | 14 | 12 | 13 | 17 | 17 | 12 | 9 | 11 | 13 | 11 | 11 | 14 | 16 | 17 | 12 | 10 | 12 | 14 | 16 | 14 | 9 | 9 | 12 | 12 | 9 | 9 | 8 | 9 | 11 | 13 | 10 | 8 | 7 | 8 | 11 | 9 | 9 | 11 | 13 | 10 | 9 | 13 | 15 | 14 | 13 | 10 | 8 | 6 | 9 | 15 | 15 | 14 | 11 | 12 | 14 | 12 | 11 | 12 | 13 | 12 | 10 | 7 | 9 | 14 | 13 | 11 | 11 | 11 | 13 | 12 | 10 | 8 | 10 | 11 | 9 | 12 | 10 | 9 | 9 | 9 | 7 | 4 | 7 | 9 | 9 | 11 | 8 | 12 | 13 | 11 | 10 | 7 | 12 | 8 | 6 | 4 | 5 | 5 | 6 | 7 | 8 | 11 | 8 | 7 | 7 | 7 | 9 | 8 | 6 | 7 | 5 | 9 | 7 | 10 | 12 | 15 | 11 | 9 | 8 | 7 | 7 | 7 | 9 | 11 | 9 | 11 | 9 | 7 | 8 | 7 | 12 | 10 | 11 | 6 | 10 | 11 | 10 | 9 | 7 | 6 | 7 | 8 | 8 | 10 | 9 | 4 | 7 | 10 | 9 | 12 | 9 | 10 | 10 | 8 | 8 | 6 | 3 | 6 | 11 | 10 | 7 | 6 | 7 | 8 | 7 | 9 | 9 | 7 | 9 | 7 | 6 | 11 | 12 | 7 | 8 | 9 | 10 | 10 | 9 | 8 | 9 | 11 | 9 | 8 | 9 | 8 | 7 | 5 | 9 | 15 | 14 | 13 | 10 | 7 | 9 | 10 | 10 | 10 | 10 | 9 | 9 | 9 | 5 | 6 | 5 | 8 | 12 | 9 | 8 | 9 | 8 | 10 | 13 | 14 | 11 | 8 | 10 | 10 | 8 | 9 | 7 | 9 | 14 | 11 | 9 | 7 | 11 | 11 | 11 | 11 | 9 | 9 | 9 | 12 | 13 | 15 | 18 | 15 | 13 | 13 | 15 | 10 | 7 | 10 | 9 | 12 | 11 | 9 | 9 | 7 | 8 | 11 | 11 | 14 | 16 | 12 | 11 | 22 | 13 | 10 | 18 | 8 | 16 | 7 | 13 | 12 | 8 | 12 | 17 | 8 | 9 | 10 | 17 | 12 | 16 | 13 | 7 | 7 | 10 | 13 | 6 | 1 | 8 | 11 | 4 | 12 | 5 | 6 | 7 | 8 | 9 | 7 | 11 | 5 | 8 | 6 | 10 | 6 | 4 | 6 | 5 | 7 | 3 | 2 | 4 | 6 | 7 | 4 | 3 | 5 | 8 | 5 | 4 | 5 | 2 | 2 | 3 | 6 | 6 | 7 | 2 | 3 | 3 | 5 | 5 | 4 | 5 | 2 | 2 | 3 | 5 | 3 | 1 | 3 | 1 | 2 | 4 | 5 | 3 | 1 | 2 | 1 | 3 | 3 | 2 | 3 | 3 | 1 | 1 | 2 | 1 | 2 | 3 | 0 | 3 | 5 | 3 | 2 | 2 | 3 | 3 | 2 | 1 | 1 | 2 | 3 | 3 | 2 | 2 | 1 | 2 | 2 | 2 | 2 | 2 | 2 | 4 | 2 | 1 | 4 | 3 | 1 | 1 | 3 | 3 | 2 | 1 | 2 | 2 | 2 | 3 | 2 | 3 | 1 | 2 | 1 | 1 | 3 | 2 | 2 | 1 | 1 | 2 | 3 | 2 | 3 | 2 | 0 | 0 | 3 | 3 | 1 | 1 | 2 | 1 | 0 | 1 | 2 | 1 | 2 | 2 | 2 | 2 | 1 | 2 | 2 | 1 | 1 | 1 | 1 | 2 | 4 | 4 | 2 | 2 | 4 | 4 | 1 | 1 | 1 | 2 | 3 | 1 | 2 | 1 | 1 | 3 | 3 | 2 | 1 | 1 | 1 | 2 | 0 | 1 | 0 | 3 | 1 | 2 | 2 | 2 | 1 | 1 | 0 | 3 | 1 | 2 | 1 | 7 | 0 | 0 | 2 | 2 | 0 | 1 | 3 | 2 | 0 | 2 | 1 | 0 | 0 | 4 | 1 | 3 | 2 | 6 | 0 | 0 | 1 | 1 | 0 | 2 | 0 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 3 | 0 | 1 | 0 | 1 | 3 | 1 | 2 | 1 | 1 | 1 | 4 | 1 | 2 | 1 | 2 | 1 | 2 | 0 | 0 | 2 | 1 | 2 | 2 | 0 | 1 | 1 | 3 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 2 | 0 | 1 | 1 | 0 | 1 | 2 | 1 | 1 | 2 | 1 | 2 | 3 | 2 | 1 | 0 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 2 | 2 | 2 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 2 | 1 | 0 | 1 | 0 | 0 | 2 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 2 | 0 | 0 | 3 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 2 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 2 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 2 | 0 | 2 | 2 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 3 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 2 | 2 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 8 | 409 | 27 | 38 | 39 | 46 | 47 | 46 | 61 | 55 | 69 | 67 | 93 | 95 | 104 | 131 | 167 | 206 | 233 | 263 | 300 | 347 | 415 | 536 | 621 | 695 | 776 | 825 | 926 | 1062 | 1171 | 1256 | 1353 | 1454 | 1498 | 1489 | 1479 | 1507 | 1513 | 1526 | 1529 | 1562 | 1601 | 1666 | 1717 | 1634 | 1574 | 1598 | 1606 | 1638 | 1651 | 1656 | 1644 | 1624 | 1630 | 1683 | 1641 | 1621 | 1630 | 1616 | 1610 | 1588 | 1566 | 1598 | 1605 | 1618 | 1641 | 1602 | 1594 | 1595 | 1632 | 1617 | 1592 | 1593 | 1616 | 1620 | 1625 | 1650 | 1641 | 1665 | 1660 | 1642 | 1632 | 1596 | 1598 | 1623 | 1669 | 1748 | 1723 | 1625 | 1568 | 1587 | 1618 | 1643 | 1648 | 1617 | 1607 | 1563 | 1593 | 1544 | 1540 | 1583 | 1587 | 1559 | 1542 | 1533 | 1512 | 1506 | 1597 | 1597 | 1538 | 1524 | 1555 | 1525 | 1511 | 1497 | 1493 | 1550 | 1536 | 1561 | 1431 | 1400 | 1382 | 1396 | 1396 | 1417 | 1395 | 1408 | 1425 | 1404 | 1398 | 1414 | 1385 | 1373 | 1405 | 1426 | 1415 | 1382 | 1427 | 1425 | 1411 | 1393 | 1391 | 1444 | 1480 | 1456 | 1441 | 1432 | 1421 | 1474 | 1518 | 1477 | 1434 | 1468 | 1451 | 1535 | 1513 | 1469 | 1498 | 1498 | 1509 | 1585 | 1593 | 1624 | 1632 | 1642 | 1704 | 1745 | 1749 | 1840 | 1879 | 1932 | 1816 | 1930 | 1876 | 2101 | 2001 | 2059 | 2082 | 2105 | 2026 | 2143 | 2102 | 2087 | 2130 | 2119 | 2093 | 2118 | 2107 | 2067 | 2035 | 2044 | 2095 | 2043 | 1996 | 2015 | 2015 | 1976 | 1950 | 1930 | 1922 | 1914 | 1867 | 1842 | 1805 | 1781 | 1800 | 1790 | 1765 | 1734 | 1703 | 1688 | 1711 | 1661 | 1619 | 1623 | 1574 | 1551 | 1565 | 1569 | 1497 | 1475 | 1470 | 1498 | 1514 | 1459 | 1461 | 1443 | 1388 | 1415 | 1448 | 1465 | 1410 | 1381 | 1324 | 1346 | 1394 | 1367 | 1319 | 1351 | 1355 | 1327 | 1301 | 1256 | 1243 | 1271 | 1291 | 1269 | 1244 | 1254 | 1259 | 1221 | 1195 | 1208 | 1194 | 1172 | 1181 | 1167 | 1137 | 1149 | 1186 | 1177 | 1127 | 1098 | 1088 | 1073 | 1078 | 1102 | 1084 | 1046 | 1025 | 1039 | 1006 | 1028 | 1037 | 1001 | 997 | 1006 | 1037 | 998 | 965 | 1003 | 1004 | 977 | 967 | 967 | 964 | 943 | 946 | 965 | 939 | 920 | 900 | 882 | 889 | 893 | 896 | 884 | 879 | 880 | 870 | 869 | 871 | 872 | 878 | 879 | 861 | 847 | 828 | 842 | 866 | 834 | 826 | 838 | 841 | 823 | 796 | 764 | 792 | 823 | 787 | 768 | 755 | 739 | 745 | 755 | 757 | 751 | 721 | 748 | 738 | 692 | 668 | 674 | 711 | 745 | 696 | 694 | 682 | 663 | 617 | 615 | 649 | 651 | 642 | 642 | 631 | 630 | 642 | 628 | 645 | 668 | 640 | 632 | 642 | 638 | 620 | 584 | 584 | 598 | 586 | 586 | 568 | 576 | 588 | 567 | 554 | 543 | 575 | 577 | 572 | 574 | 552 | 545 | 580 | 611 | 529 | 603 | 579 | 548 | 525 | 522 | 566 | 551 | 552 | 521 | 557 | 506 | 490 | 515 | 531 | 501 | 515 | 509 | 542 | 542 | 514 | 511 | 512 | 503 | 515 | 505 | 503 | 509 | 481 | 469 | 494 | 520 | 517 | 506 | 507 | 496 | 488 | 477 | 489 | 481 | 463 | 448 | 450 | 467 | 473 | 479 | 471 | 515 | 494 | 471 | 500 | 507 | 473 | 473 | 473 | 470 | 476 | 452 | 482 | 509 | 500 | 512 | 525 | 529 | 525 | 540 | 539 | 556 | 580 | 595 | 631 | 665 | 675 | 688 | 728 | 781 | 858 | 947 | 989 | 1016 | 1073 | 1180 | 1294 | 1398 | 1533 | 1678 | 1801 | 1938 | 2138 | 2372 | 2571 | 2769 | 2999 | 3271 | 3587 | 3914 | 4220 | 4571 | 4950 | 5268 | 5682 | 6092 | 6428 | 6887 | 7286 | 7556 | 8098 | 8664 | 9059 | 9456 | 9850 | 10106 | 10457 | 10914 | 11183 | 11591 | 11917 | 12173 | 12329 | 12456 | 12545 | 12624 | 12572 | 12496 | 12470 | 12390 | 12293 | 12124 | 11778 | 11441 | 11164 | 10777 | 10299 | 9918 | 9529 | 9111 | 8755 | 8293 | 7800 | 7283 | 6813 | 6322 | 5808 | 5371 | 4990 | 4648 | 4300 | 3943 | 3552 | 3307 | 3032 | 2715 | 2495 | 2269 | 2085 | 1919 | 1796 | 1579 | 1565 | 1433 | 1318 | 1169 | 1119 | 1119 | 1084 | 1032 | 988 | 933 | 929 | 906 | 891 | 861 | 843 | 849 | 866 | 868 | 854 | 853 | 815 | 781 | 793 | 789 | 797 | 839 | 849 | 814 | 814 | 846 | 836 | 794 | 800 | 815 | 794 | 814 | 801 | 801 | 815 | 821 | 820 | 828 | 803 | 800 | 786 | 782 | 822 | 774 | 759 | 774 | 767 | 753 | 762 | 776 | 742 | 770 | 752 | 781 | 823 | 742 | 751 | 766 | 797 | 740 | 731 | 743 | 715 | 723 | 726 | 717 | 688 | 711 | 731 | 732 | 726 | 738 | 727 | 722 | 704 | 665 | 704 | 747 | 725 | 720 | 708 | 692 | 692 | 689 | 686 | 681 | 709 | 726 | 731 | 726 | 693 | 719 | 722 | 732 | 723 | 706 | 703 | 731 | 739 | 711 | 730 | 725 | 711 | 722 | 706 | 711 | 701 | 700 | 691 | 695 | 716 | 740 | 738 | 720 | 742 | 773 | 754 | 716 | 715 | 746 | 773 | 785 | 772 | 797 | 807 | 816 | 811 | 782 | 794 | 810 | 814 | 827 | 845 | 858 | 883 | 895 | 885 | 891 | 916 | 935 | 928 | 895 | 893 | 909 | 908 | 896 | 901 | 906 | 914 | 919 | 888 | 885 | 912 | 881 | 875 | 897 | 887 | 885 | 892 | 892 | 889 | 870 | 845 | 868 | 852 | 824 | 793 | 784 | 783 | 776 | 786 | 800 | 790 | 775 | 779 | 764 | 780 | 752 | 727 | 716 | 714 | 696 | 657 | 657 | 662 | 654 | 657 | 632 | 629 | 637 | 623 | 634 | 636 | 593 | 594 | 568 | 585 | 567 | 577 | 554 | 615 | 550 | 514 | 508 | 522 | 526 | 538 | 535 | 532 | 517 | 479 | 468 | 499 | 502 | 473 | 475 | 457 | 473 | 452 | 462 | 491 | 463 | 469 | 436 | 418 | 443 | 445 | 436 | 418 | 421 | 418 | 408 | 413 | 412 | 400 | 377 | 389 | 443 | 408 | 388 | 406 | 364 | 386 | 384 | 385 | 380 | 364 | 341 | 338 | 355 | 367 | 361 | 361 | 356 | 333 | 340 | 326 | 342 | 337 | 319 | 329 | 321 | 311 | 294 | 307 | 317 | 327 | 307 | 296 | 305 | 319 | 325 | 310 | 289 | 280 | 269 | 276 | 288 | 294 | 295 | 266 | 274 | 272 | 267 | 252 | 253 | 258 | 256 | 255 | 258 | 260 | 249 | 234 | 240 | 250 | 239 | 233 | 231 | 241 | 226 | 205 | 215 | 218 | 208 | 207 | 213 | 209 | 212 | 219 | 216 | 208 | 195 | 184 | 190 | 191 | 183 | 191 | 194 | 194 | 195 | 192 | 189 | 191 | 194 | 185 | 182 | 188 | 186 | 178 | 171 | 163 | 163 | 172 | 176 | 167 | 168 | 167 | 158 | 171 | 175 | 178 | 156 | 150 | 159 | 158 | 158 | 164 | 160 | 147 | 147 | 152 | 158 | 156 | 142 | 135 | 146 | 164 | 156 | 148 | 140 | 138 | 132 | 135 | 152 | 154 | 146 | 131 | 126 | 129 | 130 | 130 | 129 | 134 | 129 | 133 | 147 | 145 | 130 | 132 | 148 | 139 | 116 | 139 | 126 | 122 | 142 | 144 | 141 | 145 | 144 | 140 | 140 | 143 | 148 | 140 | 145 | 149 | 153 | 132 | 158 | 154 | 155 | 154 | 134 | 156 | 166 | 174 | 188 | 193 | 191 | 195 | 207 | 220 | 219 | 232 | 233 | 230 | 234 | 300 | 288 | 322 | 291 | 340 | 360 | 375 | 395 | 409 | 421 | 449 | 508 | 543 | 539 | 568 | 615 | 656 | 696 | 704 | 738 | 835 | 849 | 917 | 951 | 946 | 992 | 1024 | 1058 | 1097 | 1171 | 1211 | 1262 | 1327 | 1338 | 1388 | 1481 | 1520 | 1518 | 1586 | 1644 | 1777 | 1742 | 1759 | 1771 | 1705 | 1736 | 1794 | 1805 | 1802 | 1813 | 1825 | 1869 | 1868 | 1788 | 1779 | 1820 | 1801 | 1745 | 1692 | 1788 | 1815 | 1669 | 1654 | 1593 | 1585 | 1545 | 1514 | 1469 | 1393 | 1331 | 1293 | 1272 | 1232 | 1165 | 1094 | 1019 | 994 | 974 | 939 | 880 | 840 | 759 | 719 | 651 | 613 | 591 | 540 | 500 | 464 | 443 | 437 | 402 | 342 | 314 | 311 | 294 | 259 | 231 | 224 | 209 | 184 | 159 | 147 | 132 | 119 | 107 | 99 | 94 | 88 | 78 | 71 | 65 | 59 | 59 | 58 | 53 | 51 | 54 | 53 | 47 | 38 | 34 | 36 | 34 | 31 | 33 | 34 | 31 | 29 | 29 | 27 | 22 | 20 | 23 | 20 | 20 | 27 | 31 | 32 | 30 | 26 | 26 | 28 | 26 | 21 | 22 | 20 | 17 | 20 | 25 | 26 | 23 | 26 | 26 | 24 | 22 | 24 | 19 | 21 | 24 | 27 | 24 | 29 | 22 | 21 | 27 | 23 | 25 | 31 | 30 | 23 | 19 | 19 | 22 | 20 | 21 | 24 | 26 | 24 | 26 | 18 | 20 | 21 | 26 | 26 | 26 | 23 | 26 | 23 | 20 | 21 | 17 | 18 | 16 | 23 | 18 | 18 | 26 | 31 | 32 | 23 | 28 | 21 | 18 | 28 | 20 | 18 | 17 | 17 | 21 | 23 | 23 | 20 | 16 | 14 | 19 | 24 | 25 | 21 | 26 | 13 | 15 | 16 | 17 | 19 | 15 | 17 | 20 | 19 | 16 | 16 | 18 | 17 | 16 | 17 | 19 | 20 | 15 | 18 | 22 | 18 | 15 | 16 | 20 | 24 | 19 | 14 | 14 | 17 | 18 | 18 | 16 | 14 | 17 | 19 | 17 | 17 | 16 | 17 | 19 | 15 | 14 | 11 | 12 | 13 | 12 | 10 | 9 | 11 | 15 | 17 | 17 | 17 | 14 | 13 | 11 | 13 | 14 | 15 | 18 | 19 | 15 | 12 | 12 | 14 | 12 | 11 | 12 | 12 | 13 | 10 | 12 | 18 | 18 | 16 | 15 | 12 | 16 | 15 | 10 | 10 | 15 | 15 | 11 | 13 | 15 | 12 | 10 | 10 | 9 | 8 | 10 | 8 | 9 | 12 | 11 | 9 | 10 | 13 | 16 | 14 | 9 | 11 | 14 | 15 | 12 | 9 | 10 | 11 | 13 | 14 | 13 | 12 | 12 | 10 | 11 | 9 | 15 | 12 | 10 | 10 | 8 | 8 | 8 | 9 | 8 | 9 | 9 | 8 | 8 | 9 | 9 | 10 | 10 | 9 | 9 | 13 | 9 | 9 | 10 | 11 | 13 | 14 | 11 | 8 | 6 | 8 | 10 | 10 | 8 | 8 | 10 | 6 | 6 | 5 | 3 | 6 | 10 | 10 | 9 | 6 | 7 | 3 | 7 | 8 | 7 | 7 | 6 | 8 | 11 | 8 | 9 | 14 | 7 | 12 | 14 | 8 | 5 | 11 | 9 | 5 | 7 | 9 | 8 | 5 | 6 | 8 | 7 | 8 | 7 | 7 | 9 | 10 | 6 | 10 | 12 | 6 | 15 | 3 | 5 | 11 | 5 | 5 | 6 | 7 | 5 | 8 | 8 | 4 | 8 | 12 | 11 | 7 | 10 | 8 | 6 | 10 | 8 | 10 | 11 | 12 | 8 | 4 | 3 | 6 | 8 | 7 | 6 | 10 | 10 | 8 | 8 | 10 | 10 | 7 | 7 | 11 | 11 | 7 | 8 | 10 | 9 | 9 | 8 | 7 | 7 | 9 | 8 | 7 | 7 | 8 | 6 | 4 | 7 | 11 | 5 | 8 | 15 | 11 | 8 | 9 | 9 | 9 | 8 | 8 | 10 | 11 | 10 | 7 | 6 | 9 | 8 | 7 | 9 | 11 | 9 | 10 | 9 | 10 | 10 | 7 | 8 | 8 | 10 | 12 | 8 | 8 | 12 | 13 | 10 | 10 | 11 | 11 | 9 | 9 | 11 | 10 | 10 | 13 | 11 | 10 | 11 | 11 | 12 | 15 | 16 | 14 | 9 | 13 | 14 | 14 | 13 | 10 | 13 | 14 | 13 | 12 | 10 | 12 | 12 | 11 | 11 | 15 | 11 | 11 | 13 | 15 | 10 | 10 | 8 | 9 | 11 | 11 | 7 | 15 | 13 | 10 | 11 | 8 | 8 | 11 | 11 | 8 | 9 | 12 | 11 | 9 | 12 | 10 | 7 | 10 | 9 | 6 | 8 | 2 | 5 | 7 | 10 | 7 | 5 | 4 | 7 | 11 | 7 | 8 | 6 | 7 | 2 | 6 | 1 | 2 | 5 | 2 | 3 | 4 | 4 | 2 | 5 | 3 | 4 | 2 | 3 | 5 | 4 | 3 | 1 | 4 | 3 | 13 | 5 | 0 | 2 | 4 | 4 | 3 | 6 | 2 | 4 | 4 | 2 | 0 | 2 | 2 | 1 | 2 | 1 | 3 | 1 | 2 | 3 | 4 | 5 | 4 | 2 | 1 | 4 | 4 | 3 | 2 | 0 | 2 | 4 | 3 | 2 | 2 | 1 | 3 | 2 | 2 | 1 | 3 | 3 | 1 | 3 | 1 | 3 | 4 | 2 | 2 | 1 | 2 | 4 | 3 | 3 | 3 | 5 | 2 | 1 | 1 | 2 | 1 | 1 | 3 | 4 | 3 | 0 | 1 | 2 | 1 | 0 | 1 | 1 | 1 | 2 | 4 | 3 | 2 | 0 | 1 | 2 | 2 | 2 | 2 | 2 | 1 | 3 | 3 | 2 | 1 | 1 | 2 | 4 | 2 | 0 | 1 | 2 | 2 | 3 | 3 | 1 | 2 | 1 | 2 | 3 | 2 | 1 | 1 | 3 | 3 | 3 | 2 | 2 | 1 | 0 | 0 | 1 | 1 | 2 | 1 | 3 | 3 | 1 | 0 | 2 | 2 | 2 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 2 | 2 | 2 | 2 | 1 | 2 | 3 | 4 | 3 | 2 | 2 | 4 | 4 | 2 | 2 | 2 | 1 | 1 | 3 | 1 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 2 | 3 | 3 | 5 | 2 | 0 | 1 | 0 | 0 | 2 | 2 | 0 | 2 | 1 | 0 | 1 | 0 | 0 | 0 | 2 | 2 | 0 | 1 | 2 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 2 | 0 | 0 | 1 | 1 | 0 | 0 | 3 | 1 | 1 | 3 | 2 | 1 | 2 | 1 | 1 | 1 | 0 | 1 | 2 | 0 | 1 | 1 | 1 | 3 | 0 | 0 | 1 | 0 | 1 | 2 | 0 | 2 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 2 | 2 | 0 | 1 | 1 | 0 | 1 | 1 | 4 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 2 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 2 | 2 | 2 | 1 | 1 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 2 | 1 | 0 | 1 | 0 | 0 | 1 | 2 | 1 | 1 | 0 | 1 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 3 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 2 | 1 | 0 | 1 | 0 | 0 | 0 | 2 | 2 | 3 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 2 | 0 | 1 | 1 | 2 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 1 | 0 | 1 | 0 | 0 | 3 | 0 | 2 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 2 | 2 | 0 | 2 | 2 | 0 | 3 | 1 | 1 | 1 | 2 | 0 | 0 | 2 | 2 | 2 | 1 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 9 | 390 | 34 | 40 | 42 | 38 | 40 | 44 | 43 | 59 | 76 | 84 | 106 | 105 | 120 | 135 | 178 | 206 | 232 | 259 | 300 | 356 | 454 | 520 | 556 | 654 | 756 | 858 | 977 | 1039 | 1133 | 1228 | 1308 | 1457 | 1506 | 1457 | 1515 | 1566 | 1553 | 1559 | 1565 | 1541 | 1536 | 1589 | 1654 | 1581 | 1561 | 1596 | 1609 | 1589 | 1563 | 1582 | 1593 | 1597 | 1594 | 1616 | 1600 | 1581 | 1567 | 1574 | 1601 | 1596 | 1550 | 1537 | 1537 | 1621 | 1665 | 1547 | 1528 | 1553 | 1589 | 1638 | 1636 | 1580 | 1547 | 1589 | 1684 | 1667 | 1613 | 1663 | 1674 | 1641 | 1648 | 1678 | 1700 | 1673 | 1677 | 1738 | 1725 | 1675 | 1618 | 1588 | 1564 | 1551 | 1604 | 1654 | 1624 | 1594 | 1581 | 1563 | 1537 | 1521 | 1509 | 1507 | 1515 | 1532 | 1516 | 1518 | 1621 | 1581 | 1495 | 1462 | 1455 | 1479 | 1501 | 1502 | 1492 | 1535 | 1578 | 1606 | 1464 | 1416 | 1340 | 1337 | 1365 | 1341 | 1362 | 1427 | 1364 | 1373 | 1366 | 1328 | 1396 | 1409 | 1396 | 1401 | 1389 | 1399 | 1412 | 1396 | 1403 | 1415 | 1428 | 1393 | 1347 | 1372 | 1391 | 1387 | 1394 | 1425 | 1468 | 1435 | 1418 | 1457 | 1434 | 1451 | 1450 | 1481 | 1512 | 1481 | 1531 | 1518 | 1605 | 1646 | 1603 | 1654 | 1694 | 1703 | 1761 | 1769 | 1813 | 1870 | 1862 | 1928 | 1958 | 1921 | 1970 | 2031 | 2102 | 2116 | 2184 | 2106 | 2155 | 2173 | 2132 | 2126 | 2125 | 2118 | 2120 | 2153 | 2100 | 2046 | 2002 | 1993 | 1998 | 2003 | 2026 | 1953 | 1902 | 1946 | 1930 | 1874 | 1820 | 1782 | 1768 | 1747 | 1755 | 1728 | 1749 | 1749 | 1698 | 1632 | 1659 | 1626 | 1572 | 1589 | 1592 | 1603 | 1561 | 1570 | 1518 | 1518 | 1524 | 1457 | 1476 | 1484 | 1475 | 1503 | 1474 | 1420 | 1423 | 1387 | 1332 | 1359 | 1361 | 1374 | 1363 | 1352 | 1299 | 1285 | 1292 | 1253 | 1274 | 1282 | 1267 | 1249 | 1220 | 1185 | 1191 | 1208 | 1196 | 1173 | 1169 | 1185 | 1180 | 1154 | 1132 | 1138 | 1137 | 1115 | 1116 | 1124 | 1109 | 1088 | 1100 | 1082 | 1063 | 1067 | 1074 | 1057 | 1049 | 1038 | 1028 | 1028 | 1050 | 1050 | 1003 | 1012 | 998 | 976 | 978 | 985 | 981 | 982 | 995 | 958 | 925 | 933 | 946 | 944 | 953 | 922 | 902 | 919 | 930 | 910 | 868 | 862 | 900 | 916 | 894 | 851 | 821 | 835 | 849 | 850 | 863 | 861 | 860 | 870 | 848 | 809 | 817 | 843 | 863 | 841 | 799 | 808 | 779 | 758 | 822 | 827 | 826 | 771 | 735 | 735 | 727 | 757 | 725 | 712 | 699 | 666 | 687 | 719 | 716 | 702 | 676 | 671 | 666 | 682 | 706 | 678 | 621 | 619 | 629 | 643 | 648 | 649 | 645 | 629 | 615 | 596 | 608 | 611 | 569 | 574 | 606 | 605 | 591 | 572 | 572 | 584 | 588 | 574 | 603 | 565 | 545 | 555 | 565 | 565 | 566 | 575 | 564 | 536 | 551 | 621 | 593 | 544 | 551 | 541 | 543 | 535 | 548 | 529 | 497 | 531 | 508 | 525 | 526 | 501 | 536 | 500 | 497 | 496 | 510 | 509 | 505 | 527 | 529 | 513 | 486 | 477 | 470 | 464 | 475 | 470 | 475 | 490 | 491 | 490 | 469 | 452 | 470 | 496 | 480 | 468 | 456 | 479 | 479 | 466 | 485 | 470 | 478 | 488 | 511 | 488 | 487 | 476 | 484 | 478 | 490 | 507 | 497 | 477 | 511 | 472 | 485 | 482 | 512 | 553 | 522 | 516 | 555 | 573 | 601 | 617 | 610 | 632 | 678 | 728 | 740 | 792 | 848 | 896 | 961 | 1039 | 1089 | 1159 | 1274 | 1389 | 1510 | 1662 | 1825 | 1978 | 2145 | 2349 | 2600 | 2812 | 3035 | 3306 | 3546 | 3844 | 4176 | 4552 | 4951 | 5326 | 5691 | 6077 | 6526 | 6839 | 7303 | 7758 | 8210 | 8740 | 9088 | 9444 | 9762 | 10107 | 10506 | 10959 | 11383 | 11716 | 11911 | 12058 | 12225 | 12371 | 12469 | 12567 | 12588 | 12531 | 12441 | 12337 | 12204 | 12004 | 11641 | 11270 | 11004 | 10765 | 10420 | 9929 | 9429 | 8953 | 8425 | 7919 | 7513 | 7070 | 6629 | 6204 | 5726 | 5320 | 4976 | 4561 | 4230 | 3905 | 3480 | 3106 | 2860 | 2639 | 2356 | 2178 | 1990 | 1889 | 1713 | 1547 | 1455 | 1374 | 1245 | 1234 | 1124 | 1034 | 1035 | 1015 | 985 | 913 | 898 | 878 | 879 | 891 | 899 | 860 | 833 | 824 | 828 | 810 | 813 | 810 | 789 | 819 | 819 | 795 | 800 | 802 | 823 | 829 | 801 | 812 | 819 | 810 | 819 | 830 | 794 | 754 | 770 | 814 | 839 | 804 | 782 | 798 | 791 | 802 | 770 | 755 | 785 | 830 | 787 | 777 | 791 | 782 | 766 | 768 | 744 | 774 | 739 | 713 | 671 | 722 | 743 | 731 | 697 | 717 | 717 | 690 | 720 | 720 | 722 | 702 | 692 | 691 | 708 | 724 | 710 | 714 | 716 | 690 | 672 | 678 | 697 | 700 | 694 | 697 | 689 | 681 | 699 | 681 | 662 | 687 | 713 | 726 | 697 | 689 | 690 | 682 | 701 | 679 | 696 | 660 | 679 | 677 | 697 | 698 | 664 | 674 | 672 | 670 | 698 | 723 | 700 | 698 | 717 | 728 | 746 | 761 | 750 | 728 | 739 | 766 | 767 | 780 | 789 | 786 | 763 | 800 | 800 | 788 | 800 | 809 | 828 | 834 | 828 | 828 | 833 | 833 | 849 | 859 | 848 | 848 | 840 | 860 | 890 | 879 | 895 | 927 | 913 | 909 | 915 | 918 | 917 | 917 | 924 | 924 | 905 | 867 | 864 | 902 | 891 | 858 | 841 | 866 | 859 | 828 | 827 | 813 | 819 | 847 | 835 | 816 | 804 | 796 | 783 | 764 | 773 | 784 | 756 | 742 | 741 | 709 | 706 | 688 | 713 | 680 | 682 | 679 | 677 | 653 | 652 | 645 | 633 | 647 | 634 | 609 | 608 | 591 | 587 | 574 | 596 | 594 | 559 | 541 | 542 | 557 | 562 | 550 | 524 | 535 | 542 | 529 | 528 | 511 | 497 | 507 | 494 | 469 | 474 | 476 | 470 | 452 | 465 | 460 | 476 | 430 | 457 | 462 | 431 | 448 | 458 | 433 | 420 | 432 | 445 | 447 | 440 | 423 | 395 | 388 | 391 | 375 | 350 | 388 | 387 | 380 | 351 | 372 | 355 | 337 | 340 | 348 | 354 | 342 | 326 | 324 | 329 | 335 | 337 | 340 | 345 | 318 | 340 | 355 | 342 | 356 | 301 | 297 | 317 | 323 | 295 | 286 | 282 | 281 | 294 | 308 | 298 | 275 | 267 | 268 | 272 | 264 | 263 | 259 | 270 | 276 | 267 | 283 | 271 | 261 | 250 | 244 | 233 | 234 | 244 | 253 | 250 | 232 | 227 | 227 | 232 | 230 | 225 | 228 | 230 | 220 | 207 | 219 | 217 | 218 | 211 | 202 | 206 | 211 | 196 | 188 | 189 | 193 | 197 | 181 | 178 | 176 | 168 | 180 | 194 | 186 | 182 | 184 | 180 | 168 | 166 | 169 | 171 | 179 | 181 | 170 | 160 | 167 | 183 | 179 | 162 | 160 | 164 | 161 | 159 | 166 | 160 | 155 | 151 | 148 | 148 | 148 | 150 | 145 | 148 | 156 | 150 | 141 | 147 | 142 | 129 | 141 | 145 | 154 | 145 | 134 | 131 | 132 | 139 | 135 | 137 | 137 | 137 | 143 | 142 | 135 | 126 | 131 | 144 | 139 | 136 | 147 | 145 | 139 | 146 | 134 | 137 | 139 | 139 | 132 | 134 | 141 | 140 | 135 | 133 | 135 | 136 | 142 | 143 | 147 | 153 | 152 | 146 | 145 | 152 | 151 | 153 | 173 | 183 | 185 | 189 | 182 | 175 | 179 | 194 | 212 | 212 | 212 | 220 | 246 | 247 | 289 | 302 | 300 | 297 | 307 | 343 | 367 | 368 | 375 | 399 | 441 | 471 | 492 | 530 | 569 | 593 | 628 | 667 | 702 | 682 | 793 | 796 | 877 | 887 | 896 | 971 | 1019 | 1052 | 1104 | 1129 | 1168 | 1201 | 1221 | 1276 | 1350 | 1416 | 1495 | 1481 | 1541 | 1615 | 1600 | 1723 | 1691 | 1643 | 1678 | 1696 | 1717 | 1753 | 1771 | 1849 | 1864 | 1828 | 1834 | 1863 | 1842 | 1788 | 1811 | 1820 | 1750 | 1745 | 1761 | 1811 | 1661 | 1602 | 1640 | 1573 | 1490 | 1441 | 1392 | 1351 | 1299 | 1250 | 1219 | 1191 | 1137 | 1093 | 1095 | 1032 | 945 | 899 | 843 | 766 | 694 | 689 | 659 | 629 | 568 | 526 | 489 | 456 | 420 | 386 | 356 | 332 | 315 | 296 | 267 | 241 | 218 | 191 | 181 | 173 | 147 | 146 | 133 | 108 | 90 | 87 | 88 | 82 | 79 | 75 | 70 | 60 | 52 | 49 | 46 | 42 | 43 | 42 | 43 | 38 | 26 | 31 | 34 | 30 | 30 | 32 | 35 | 38 | 33 | 30 | 29 | 25 | 21 | 22 | 23 | 20 | 21 | 25 | 30 | 26 | 23 | 26 | 29 | 31 | 31 | 28 | 24 | 25 | 29 | 28 | 22 | 22 | 29 | 29 | 23 | 22 | 28 | 29 | 20 | 22 | 23 | 24 | 25 | 25 | 19 | 19 | 25 | 24 | 20 | 21 | 21 | 19 | 19 | 21 | 24 | 21 | 26 | 23 | 24 | 18 | 22 | 24 | 18 | 20 | 18 | 22 | 27 | 25 | 27 | 27 | 19 | 15 | 17 | 15 | 18 | 14 | 17 | 18 | 22 | 20 | 18 | 23 | 20 | 19 | 19 | 16 | 18 | 18 | 19 | 16 | 14 | 26 | 25 | 19 | 12 | 12 | 18 | 12 | 16 | 16 | 29 | 20 | 20 | 21 | 15 | 14 | 17 | 23 | 22 | 20 | 20 | 21 | 25 | 19 | 22 | 23 | 19 | 13 | 19 | 25 | 19 | 17 | 14 | 15 | 19 | 20 | 16 | 14 | 17 | 16 | 14 | 16 | 14 | 15 | 18 | 21 | 19 | 13 | 13 | 16 | 17 | 15 | 19 | 12 | 17 | 14 | 10 | 9 | 11 | 15 | 16 | 15 | 15 | 16 | 13 | 14 | 15 | 14 | 18 | 14 | 10 | 9 | 11 | 15 | 15 | 12 | 13 | 14 | 13 | 12 | 11 | 12 | 13 | 14 | 12 | 11 | 14 | 11 | 8 | 11 | 12 | 10 | 11 | 15 | 17 | 15 | 11 | 9 | 11 | 10 | 9 | 14 | 17 | 13 | 10 | 9 | 10 | 11 | 13 | 14 | 11 | 10 | 10 | 9 | 10 | 11 | 10 | 10 | 10 | 9 | 10 | 13 | 12 | 13 | 11 | 10 | 12 | 13 | 11 | 10 | 8 | 11 | 12 | 8 | 7 | 10 | 10 | 7 | 6 | 9 | 11 | 10 | 8 | 7 | 8 | 7 | 11 | 10 | 8 | 9 | 11 | 11 | 12 | 12 | 8 | 5 | 5 | 8 | 11 | 11 | 8 | 8 | 10 | 9 | 10 | 10 | 9 | 7 | 11 | 10 | 9 | 9 | 8 | 7 | 8 | 8 | 8 | 7 | 6 | 8 | 11 | 9 | 10 | 8 | 7 | 7 | 6 | 10 | 8 | 10 | 10 | 6 | 8 | 6 | 8 | 11 | 8 | 3 | 3 | 4 | 4 | 4 | 6 | 9 | 4 | 6 | 7 | 9 | 4 | 3 | 7 | 7 | 11 | 8 | 5 | 5 | 6 | 8 | 9 | 6 | 8 | 8 | 7 | 10 | 8 | 5 | 9 | 8 | 4 | 6 | 8 | 9 | 11 | 9 | 6 | 7 | 8 | 8 | 10 | 10 | 7 | 11 | 12 | 9 | 7 | 6 | 9 | 6 | 8 | 9 | 14 | 8 | 7 | 9 | 9 | 12 | 11 | 9 | 10 | 8 | 8 | 9 | 8 | 6 | 6 | 8 | 11 | 9 | 12 | 7 | 8 | 12 | 11 | 11 | 6 | 7 | 5 | 7 | 8 | 8 | 9 | 10 | 14 | 9 | 6 | 8 | 9 | 9 | 7 | 7 | 10 | 7 | 10 | 12 | 7 | 5 | 6 | 10 | 13 | 9 | 7 | 9 | 9 | 9 | 10 | 11 | 9 | 7 | 7 | 10 | 12 | 11 | 9 | 9 | 12 | 14 | 10 | 11 | 11 | 11 | 12 | 9 | 9 | 13 | 13 | 10 | 14 | 15 | 13 | 13 | 11 | 9 | 11 | 11 | 15 | 21 | 17 | 8 | 13 | 13 | 12 | 13 | 13 | 13 | 8 | 6 | 13 | 11 | 8 | 7 | 12 | 10 | 10 | 9 | 9 | 7 | 9 | 12 | 11 | 9 | 7 | 6 | 8 | 10 | 8 | 7 | 6 | 6 | 5 | 7 | 7 | 5 | 7 | 6 | 6 | 6 | 4 | 4 | 5 | 3 | 5 | 8 | 6 | 6 | 4 | 1 | 4 | 5 | 6 | 3 | 1 | 5 | 5 | 4 | 4 | 6 | 5 | 1 | 7 | 3 | 3 | 2 | 2 | 0 | 0 | 5 | 4 | 0 | 3 | 2 | 3 | 2 | 4 | 2 | 2 | 1 | 1 | 2 | 2 | 4 | 4 | 2 | 0 | 2 | 2 | 1 | 2 | 4 | 0 | 1 | 2 | 3 | 2 | 5 | 0 | 2 | 1 | 2 | 2 | 0 | 3 | 1 | 1 | 2 | 1 | 2 | 0 | 1 | 1 | 2 | 1 | 2 | 2 | 1 | 1 | 0 | 1 | 0 | 1 | 2 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 4 | 1 | 1 | 2 | 2 | 5 | 1 | 2 | 4 | 1 | 2 | 1 | 2 | 2 | 5 | 3 | 2 | 4 | 3 | 3 | 2 | 1 | 2 | 1 | 0 | 2 | 2 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 2 | 2 | 1 | 0 | 2 | 4 | 4 | 2 | 1 | 1 | 3 | 2 | 1 | 2 | 4 | 4 | 2 | 0 | 1 | 1 | 2 | 2 | 3 | 2 | 3 | 2 | 1 | 1 | 0 | 1 | 0 | 1 | 3 | 2 | 3 | 3 | 1 | 2 | 2 | 1 | 0 | 1 | 0 | 1 | 1 | 3 | 2 | 2 | 1 | 2 | 2 | 1 | 1 | 2 | 2 | 1 | 2 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 3 | 1 | 2 | 4 | 2 | 0 | 0 | 1 | 0 | 1 | 2 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 4 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 5 | 2 | 4 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 2 | 1 | 0 | 0 | 1 | 3 | 3 | 0 | 1 | 3 | 2 | 1 | 2 | 0 | 1 | 0 | 0 | 2 | 1 | 1 | 2 | 1 | 3 | 0 | 0 | 0 | 0 | 3 | 0 | 2 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 2 | 2 | 0 | 0 | 0 | 1 | 1 | 2 | 1 | 0 | 1 | 1 | 1 | 1 | 2 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 2 | 2 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 2 | 1 | 2 | 2 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 2 | 0 | 2 | 2 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 1 | 0 | 1 | 2 | 1 | 0 | 0 | 3 | 1 | 0 | 2 | 2 | 1 | 0 | 0 | 0 | 3 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 2 | 0 | 0 | 0 | 2 | 0 | 1 | 2 | 0 | 0 | 2 | 1 | 1 | 0 | 2 | 3 | 0 | 0 | 1 | 1 | 2 | 2 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 3 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 2 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 1 | 2 | 0 | 1 | 0 | 2 | 1 | 1 | 2 | 2 | 1 | 1 | 1 | 3 | 3 | 2 | 1 | 1 | 1 | 2 | 0 | 0 | 1 | 1 | 3 | 2 | 9 | 5 | 6 | 8 | 5 | 4 | 3 | 3 | 3 | 0 | 2 | 4 | 3 | 2 | 1 | 3 | 0 | 1 | 4 | 2 | 0 | 3 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 2 | 0 | 0 | 0 | 2 | 1 | 2 | 1 | 2 | 1 | 1 | 0 | 1 | 3 | 3 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 2 | 0 | 3 | 2 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 2 | 2 | 2 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 3 | 0 | 1 | 0 | 1 | 3 | 2 | 2 | 4 | 2 | 0 | 0 | 4 | 1 | 0 | 3 | 1 | 0 | 0 | 0 | 0 | 2 | 2 | 2 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 2 | 1 | 2 | 2 | 0 | 0 | 1 | 2 | 1 | 2 | 0 | 1 | 0 | 1 | 1 | 2 | 0 | 0 | 0 | 1 | 1 | 2 | 0 | 1 | 3 | 1 | 0 | 2 | 1 | 1 | 0 | 2 | 2 | 2 | 1 | 0 | 0 | 4 | 2 | 0 | 1 | 2 | 1 | 1 | 0 | 1 | 1 | 2 | 2 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 2 | 0 | 1 | 0 | 1 | 0 | 0 | 2 | 0 | 1 | 1 | 1 | 1 | 3 | 0 | 0 | 0 | 0 | 1 | 1 | 2 | 0 | 2 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 2 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 2 | 0 | 1 | 0 | 1 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 1 | 0 | 2 | 0 | 0 | 1 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0