## Results
Hidex Automatic Gamma Counter
Counting 18F
Run
Template | Counting 18F
Run status | Finished
Run started | 2021-09-01 18:31
Run type | Count only, single rack
Software version | 1.8.0.0
AMG serial number | 2210227
Firmware versions | 0: 50, 1: 50, F: 100
Counting time | 60
Result normalization | 2021-09-01 11:48
Comment
Counting windows
Window start | Window end | Additional windows | Isotope | Counting efficiency | Half-life
400 | 650 |  | ¹⁸F | 1 | 1.82951 h
Racks
Rack | Time | Barcode | Slot count | Latest QC check | QC status
1 | 2021-09-01 18:31 | 006 | 10 | 2021-09-01 10:39 | Calibration suggested
Results
Rack | Vial | Time | Counted time (s) | Dead time factor | ¹⁸F (counts) | ¹⁸F (CPM) | Normalized ¹⁸F (CPM) | Normalized ¹⁸F (Bq) | Tags
1 | 1 | 2021-09-01 18:32 | 60 | 1.000 | 36 | 36 | 461 | 8
1 | 2 | 2021-09-01 18:33 | 60 | 1.403 | 1427777 | 2003368 | 25904864 | 431748
1 | 3 | 2021-09-01 18:35 | 60 | 1.401 | 1419800 | 1989532 | 25968469 | 432808
1 | 4 | 2021-09-01 18:36 | 60 | 1.393 | 1398769 | 1948594 | 25673748 | 427896
1 | 5 | 2021-09-01 18:38 | 60 | 1.392 | 1396594 | 1943394 | 25846600 | 430777
1 | 6 | 2021-09-01 18:39 | 60 | 1.138 | 608789 | 692990 | 9302945 | 155049
1 | 7 | 2021-09-01 18:41 | 60 | 1.132 | 587568 | 665222 | 9014670 | 150244
1 | 8 | 2021-09-01 18:42 | 60 | 1.129 | 578175 | 652925 | 8931737 | 148862
1 | 9 | 2021-09-01 18:44 | 60 | 1.128 | 574897 | 648446 | 8954419 | 149240
1 | 10 | 2021-09-01 18:45 | 60 | 1.000 | 40 | 40 | 558 | 9

## Spectra
Hidex Automatic Gamma Counter
Counting 18F
Run
Template | Counting 18F
Run status | Finished
Run started | 2021-09-01 18:31
Run type | Count only, single rack
Software version | 1.8.0.0
AMG serial number | 2210227
Firmware versions | 0: 50, 1: 50, F: 100
Counting time | 60
Result normalization | 2021-09-01 11:48
Comment
Spectra
Channel
Rack | Vial | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66 | 67 | 68 | 69 | 70 | 71 | 72 | 73 | 74 | 75 | 76 | 77 | 78 | 79 | 80 | 81 | 82 | 83 | 84 | 85 | 86 | 87 | 88 | 89 | 90 | 91 | 92 | 93 | 94 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 117 | 118 | 119 | 120 | 121 | 122 | 123 | 124 | 125 | 126 | 127 | 128 | 129 | 130 | 131 | 132 | 133 | 134 | 135 | 136 | 137 | 138 | 139 | 140 | 141 | 142 | 143 | 144 | 145 | 146 | 147 | 148 | 149 | 150 | 151 | 152 | 153 | 154 | 155 | 156 | 157 | 158 | 159 | 160 | 161 | 162 | 163 | 164 | 165 | 166 | 167 | 168 | 169 | 170 | 171 | 172 | 173 | 174 | 175 | 176 | 177 | 178 | 179 | 180 | 181 | 182 | 183 | 184 | 185 | 186 | 187 | 188 | 189 | 190 | 191 | 192 | 193 | 194 | 195 | 196 | 197 | 198 | 199 | 200 | 201 | 202 | 203 | 204 | 205 | 206 | 207 | 208 | 209 | 210 | 211 | 212 | 213 | 214 | 215 | 216 | 217 | 218 | 219 | 220 | 221 | 222 | 223 | 224 | 225 | 226 | 227 | 228 | 229 | 230 | 231 | 232 | 233 | 234 | 235 | 236 | 237 | 238 | 239 | 240 | 241 | 242 | 243 | 244 | 245 | 246 | 247 | 248 | 249 | 250 | 251 | 252 | 253 | 254 | 255 | 256 | 257 | 258 | 259 | 260 | 261 | 262 | 263 | 264 | 265 | 266 | 267 | 268 | 269 | 270 | 271 | 272 | 273 | 274 | 275 | 276 | 277 | 278 | 279 | 280 | 281 | 282 | 283 | 284 | 285 | 286 | 287 | 288 | 289 | 290 | 291 | 292 | 293 | 294 | 295 | 296 | 297 | 298 | 299 | 300 | 301 | 302 | 303 | 304 | 305 | 306 | 307 | 308 | 309 | 310 | 311 | 312 | 313 | 314 | 315 | 316 | 317 | 318 | 319 | 320 | 321 | 322 | 323 | 324 | 325 | 326 | 327 | 328 | 329 | 330 | 331 | 332 | 333 | 334 | 335 | 336 | 337 | 338 | 339 | 340 | 341 | 342 | 343 | 344 | 345 | 346 | 347 | 348 | 349 | 350 | 351 | 352 | 353 | 354 | 355 | 356 | 357 | 358 | 359 | 360 | 361 | 362 | 363 | 364 | 365 | 366 | 367 | 368 | 369 | 370 | 371 | 372 | 373 | 374 | 375 | 376 | 377 | 378 | 379 | 380 | 381 | 382 | 383 | 384 | 385 | 386 | 387 | 388 | 389 | 390 | 391 | 392 | 393 | 394 | 395 | 396 | 397 | 398 | 399 | 400 | 401 | 402 | 403 | 404 | 405 | 406 | 407 | 408 | 409 | 410 | 411 | 412 | 413 | 414 | 415 | 416 | 417 | 418 | 419 | 420 | 421 | 422 | 423 | 424 | 425 | 426 | 427 | 428 | 429 | 430 | 431 | 432 | 433 | 434 | 435 | 436 | 437 | 438 | 439 | 440 | 441 | 442 | 443 | 444 | 445 | 446 | 447 | 448 | 449 | 450 | 451 | 452 | 453 | 454 | 455 | 456 | 457 | 458 | 459 | 460 | 461 | 462 | 463 | 464 | 465 | 466 | 467 | 468 | 469 | 470 | 471 | 472 | 473 | 474 | 475 | 476 | 477 | 478 | 479 | 480 | 481 | 482 | 483 | 484 | 485 | 486 | 487 | 488 | 489 | 490 | 491 | 492 | 493 | 494 | 495 | 496 | 497 | 498 | 499 | 500 | 501 | 502 | 503 | 504 | 505 | 506 | 507 | 508 | 509 | 510 | 511 | 512 | 513 | 514 | 515 | 516 | 517 | 518 | 519 | 520 | 521 | 522 | 523 | 524 | 525 | 526 | 527 | 528 | 529 | 530 | 531 | 532 | 533 | 534 | 535 | 536 | 537 | 538 | 539 | 540 | 541 | 542 | 543 | 544 | 545 | 546 | 547 | 548 | 549 | 550 | 551 | 552 | 553 | 554 | 555 | 556 | 557 | 558 | 559 | 560 | 561 | 562 | 563 | 564 | 565 | 566 | 567 | 568 | 569 | 570 | 571 | 572 | 573 | 574 | 575 | 576 | 577 | 578 | 579 | 580 | 581 | 582 | 583 | 584 | 585 | 586 | 587 | 588 | 589 | 590 | 591 | 592 | 593 | 594 | 595 | 596 | 597 | 598 | 599 | 600 | 601 | 602 | 603 | 604 | 605 | 606 | 607 | 608 | 609 | 610 | 611 | 612 | 613 | 614 | 615 | 616 | 617 | 618 | 619 | 620 | 621 | 622 | 623 | 624 | 625 | 626 | 627 | 628 | 629 | 630 | 631 | 632 | 633 | 634 | 635 | 636 | 637 | 638 | 639 | 640 | 641 | 642 | 643 | 644 | 645 | 646 | 647 | 648 | 649 | 650 | 651 | 652 | 653 | 654 | 655 | 656 | 657 | 658 | 659 | 660 | 661 | 662 | 663 | 664 | 665 | 666 | 667 | 668 | 669 | 670 | 671 | 672 | 673 | 674 | 675 | 676 | 677 | 678 | 679 | 680 | 681 | 682 | 683 | 684 | 685 | 686 | 687 | 688 | 689 | 690 | 691 | 692 | 693 | 694 | 695 | 696 | 697 | 698 | 699 | 700 | 701 | 702 | 703 | 704 | 705 | 706 | 707 | 708 | 709 | 710 | 711 | 712 | 713 | 714 | 715 | 716 | 717 | 718 | 719 | 720 | 721 | 722 | 723 | 724 | 725 | 726 | 727 | 728 | 729 | 730 | 731 | 732 | 733 | 734 | 735 | 736 | 737 | 738 | 739 | 740 | 741 | 742 | 743 | 744 | 745 | 746 | 747 | 748 | 749 | 750 | 751 | 752 | 753 | 754 | 755 | 756 | 757 | 758 | 759 | 760 | 761 | 762 | 763 | 764 | 765 | 766 | 767 | 768 | 769 | 770 | 771 | 772 | 773 | 774 | 775 | 776 | 777 | 778 | 779 | 780 | 781 | 782 | 783 | 784 | 785 | 786 | 787 | 788 | 789 | 790 | 791 | 792 | 793 | 794 | 795 | 796 | 797 | 798 | 799 | 800 | 801 | 802 | 803 | 804 | 805 | 806 | 807 | 808 | 809 | 810 | 811 | 812 | 813 | 814 | 815 | 816 | 817 | 818 | 819 | 820 | 821 | 822 | 823 | 824 | 825 | 826 | 827 | 828 | 829 | 830 | 831 | 832 | 833 | 834 | 835 | 836 | 837 | 838 | 839 | 840 | 841 | 842 | 843 | 844 | 845 | 846 | 847 | 848 | 849 | 850 | 851 | 852 | 853 | 854 | 855 | 856 | 857 | 858 | 859 | 860 | 861 | 862 | 863 | 864 | 865 | 866 | 867 | 868 | 869 | 870 | 871 | 872 | 873 | 874 | 875 | 876 | 877 | 878 | 879 | 880 | 881 | 882 | 883 | 884 | 885 | 886 | 887 | 888 | 889 | 890 | 891 | 892 | 893 | 894 | 895 | 896 | 897 | 898 | 899 | 900 | 901 | 902 | 903 | 904 | 905 | 906 | 907 | 908 | 909 | 910 | 911 | 912 | 913 | 914 | 915 | 916 | 917 | 918 | 919 | 920 | 921 | 922 | 923 | 924 | 925 | 926 | 927 | 928 | 929 | 930 | 931 | 932 | 933 | 934 | 935 | 936 | 937 | 938 | 939 | 940 | 941 | 942 | 943 | 944 | 945 | 946 | 947 | 948 | 949 | 950 | 951 | 952 | 953 | 954 | 955 | 956 | 957 | 958 | 959 | 960 | 961 | 962 | 963 | 964 | 965 | 966 | 967 | 968 | 969 | 970 | 971 | 972 | 973 | 974 | 975 | 976 | 977 | 978 | 979 | 980 | 981 | 982 | 983 | 984 | 985 | 986 | 987 | 988 | 989 | 990 | 991 | 992 | 993 | 994 | 995 | 996 | 997 | 998 | 999 | 1000 | 1001 | 1002 | 1003 | 1004 | 1005 | 1006 | 1007 | 1008 | 1009 | 1010 | 1011 | 1012 | 1013 | 1014 | 1015 | 1016 | 1017 | 1018 | 1019 | 1020 | 1021 | 1022 | 1023 | 1024 | 1025 | 1026 | 1027 | 1028 | 1029 | 1030 | 1031 | 1032 | 1033 | 1034 | 1035 | 1036 | 1037 | 1038 | 1039 | 1040 | 1041 | 1042 | 1043 | 1044 | 1045 | 1046 | 1047 | 1048 | 1049 | 1050 | 1051 | 1052 | 1053 | 1054 | 1055 | 1056 | 1057 | 1058 | 1059 | 1060 | 1061 | 1062 | 1063 | 1064 | 1065 | 1066 | 1067 | 1068 | 1069 | 1070 | 1071 | 1072 | 1073 | 1074 | 1075 | 1076 | 1077 | 1078 | 1079 | 1080 | 1081 | 1082 | 1083 | 1084 | 1085 | 1086 | 1087 | 1088 | 1089 | 1090 | 1091 | 1092 | 1093 | 1094 | 1095 | 1096 | 1097 | 1098 | 1099 | 1100 | 1101 | 1102 | 1103 | 1104 | 1105 | 1106 | 1107 | 1108 | 1109 | 1110 | 1111 | 1112 | 1113 | 1114 | 1115 | 1116 | 1117 | 1118 | 1119 | 1120 | 1121 | 1122 | 1123 | 1124 | 1125 | 1126 | 1127 | 1128 | 1129 | 1130 | 1131 | 1132 | 1133 | 1134 | 1135 | 1136 | 1137 | 1138 | 1139 | 1140 | 1141 | 1142 | 1143 | 1144 | 1145 | 1146 | 1147 | 1148 | 1149 | 1150 | 1151 | 1152 | 1153 | 1154 | 1155 | 1156 | 1157 | 1158 | 1159 | 1160 | 1161 | 1162 | 1163 | 1164 | 1165 | 1166 | 1167 | 1168 | 1169 | 1170 | 1171 | 1172 | 1173 | 1174 | 1175 | 1176 | 1177 | 1178 | 1179 | 1180 | 1181 | 1182 | 1183 | 1184 | 1185 | 1186 | 1187 | 1188 | 1189 | 1190 | 1191 | 1192 | 1193 | 1194 | 1195 | 1196 | 1197 | 1198 | 1199 | 1200 | 1201 | 1202 | 1203 | 1204 | 1205 | 1206 | 1207 | 1208 | 1209 | 1210 | 1211 | 1212 | 1213 | 1214 | 1215 | 1216 | 1217 | 1218 | 1219 | 1220 | 1221 | 1222 | 1223 | 1224 | 1225 | 1226 | 1227 | 1228 | 1229 | 1230 | 1231 | 1232 | 1233 | 1234 | 1235 | 1236 | 1237 | 1238 | 1239 | 1240 | 1241 | 1242 | 1243 | 1244 | 1245 | 1246 | 1247 | 1248 | 1249 | 1250 | 1251 | 1252 | 1253 | 1254 | 1255 | 1256 | 1257 | 1258 | 1259 | 1260 | 1261 | 1262 | 1263 | 1264 | 1265 | 1266 | 1267 | 1268 | 1269 | 1270 | 1271 | 1272 | 1273 | 1274 | 1275 | 1276 | 1277 | 1278 | 1279 | 1280 | 1281 | 1282 | 1283 | 1284 | 1285 | 1286 | 1287 | 1288 | 1289 | 1290 | 1291 | 1292 | 1293 | 1294 | 1295 | 1296 | 1297 | 1298 | 1299 | 1300 | 1301 | 1302 | 1303 | 1304 | 1305 | 1306 | 1307 | 1308 | 1309 | 1310 | 1311 | 1312 | 1313 | 1314 | 1315 | 1316 | 1317 | 1318 | 1319 | 1320 | 1321 | 1322 | 1323 | 1324 | 1325 | 1326 | 1327 | 1328 | 1329 | 1330 | 1331 | 1332 | 1333 | 1334 | 1335 | 1336 | 1337 | 1338 | 1339 | 1340 | 1341 | 1342 | 1343 | 1344 | 1345 | 1346 | 1347 | 1348 | 1349 | 1350 | 1351 | 1352 | 1353 | 1354 | 1355 | 1356 | 1357 | 1358 | 1359 | 1360 | 1361 | 1362 | 1363 | 1364 | 1365 | 1366 | 1367 | 1368 | 1369 | 1370 | 1371 | 1372 | 1373 | 1374 | 1375 | 1376 | 1377 | 1378 | 1379 | 1380 | 1381 | 1382 | 1383 | 1384 | 1385 | 1386 | 1387 | 1388 | 1389 | 1390 | 1391 | 1392 | 1393 | 1394 | 1395 | 1396 | 1397 | 1398 | 1399 | 1400 | 1401 | 1402 | 1403 | 1404 | 1405 | 1406 | 1407 | 1408 | 1409 | 1410 | 1411 | 1412 | 1413 | 1414 | 1415 | 1416 | 1417 | 1418 | 1419 | 1420 | 1421 | 1422 | 1423 | 1424 | 1425 | 1426 | 1427 | 1428 | 1429 | 1430 | 1431 | 1432 | 1433 | 1434 | 1435 | 1436 | 1437 | 1438 | 1439 | 1440 | 1441 | 1442 | 1443 | 1444 | 1445 | 1446 | 1447 | 1448 | 1449 | 1450 | 1451 | 1452 | 1453 | 1454 | 1455 | 1456 | 1457 | 1458 | 1459 | 1460 | 1461 | 1462 | 1463 | 1464 | 1465 | 1466 | 1467 | 1468 | 1469 | 1470 | 1471 | 1472 | 1473 | 1474 | 1475 | 1476 | 1477 | 1478 | 1479 | 1480 | 1481 | 1482 | 1483 | 1484 | 1485 | 1486 | 1487 | 1488 | 1489 | 1490 | 1491 | 1492 | 1493 | 1494 | 1495 | 1496 | 1497 | 1498 | 1499 | 1500 | 1501 | 1502 | 1503 | 1504 | 1505 | 1506 | 1507 | 1508 | 1509 | 1510 | 1511 | 1512 | 1513 | 1514 | 1515 | 1516 | 1517 | 1518 | 1519 | 1520 | 1521 | 1522 | 1523 | 1524 | 1525 | 1526 | 1527 | 1528 | 1529 | 1530 | 1531 | 1532 | 1533 | 1534 | 1535 | 1536 | 1537 | 1538 | 1539 | 1540 | 1541 | 1542 | 1543 | 1544 | 1545 | 1546 | 1547 | 1548 | 1549 | 1550 | 1551 | 1552 | 1553 | 1554 | 1555 | 1556 | 1557 | 1558 | 1559 | 1560 | 1561 | 1562 | 1563 | 1564 | 1565 | 1566 | 1567 | 1568 | 1569 | 1570 | 1571 | 1572 | 1573 | 1574 | 1575 | 1576 | 1577 | 1578 | 1579 | 1580 | 1581 | 1582 | 1583 | 1584 | 1585 | 1586 | 1587 | 1588 | 1589 | 1590 | 1591 | 1592 | 1593 | 1594 | 1595 | 1596 | 1597 | 1598 | 1599 | 1600 | 1601 | 1602 | 1603 | 1604 | 1605 | 1606 | 1607 | 1608 | 1609 | 1610 | 1611 | 1612 | 1613 | 1614 | 1615 | 1616 | 1617 | 1618 | 1619 | 1620 | 1621 | 1622 | 1623 | 1624 | 1625 | 1626 | 1627 | 1628 | 1629 | 1630 | 1631 | 1632 | 1633 | 1634 | 1635 | 1636 | 1637 | 1638 | 1639 | 1640 | 1641 | 1642 | 1643 | 1644 | 1645 | 1646 | 1647 | 1648 | 1649 | 1650 | 1651 | 1652 | 1653 | 1654 | 1655 | 1656 | 1657 | 1658 | 1659 | 1660 | 1661 | 1662 | 1663 | 1664 | 1665 | 1666 | 1667 | 1668 | 1669 | 1670 | 1671 | 1672 | 1673 | 1674 | 1675 | 1676 | 1677 | 1678 | 1679 | 1680 | 1681 | 1682 | 1683 | 1684 | 1685 | 1686 | 1687 | 1688 | 1689 | 1690 | 1691 | 1692 | 1693 | 1694 | 1695 | 1696 | 1697 | 1698 | 1699 | 1700 | 1701 | 1702 | 1703 | 1704 | 1705 | 1706 | 1707 | 1708 | 1709 | 1710 | 1711 | 1712 | 1713 | 1714 | 1715 | 1716 | 1717 | 1718 | 1719 | 1720 | 1721 | 1722 | 1723 | 1724 | 1725 | 1726 | 1727 | 1728 | 1729 | 1730 | 1731 | 1732 | 1733 | 1734 | 1735 | 1736 | 1737 | 1738 | 1739 | 1740 | 1741 | 1742 | 1743 | 1744 | 1745 | 1746 | 1747 | 1748 | 1749 | 1750 | 1751 | 1752 | 1753 | 1754 | 1755 | 1756 | 1757 | 1758 | 1759 | 1760 | 1761 | 1762 | 1763 | 1764 | 1765 | 1766 | 1767 | 1768 | 1769 | 1770 | 1771 | 1772 | 1773 | 1774 | 1775 | 1776 | 1777 | 1778 | 1779 | 1780 | 1781 | 1782 | 1783 | 1784 | 1785 | 1786 | 1787 | 1788 | 1789 | 1790 | 1791 | 1792 | 1793 | 1794 | 1795 | 1796 | 1797 | 1798 | 1799 | 1800 | 1801 | 1802 | 1803 | 1804 | 1805 | 1806 | 1807 | 1808 | 1809 | 1810 | 1811 | 1812 | 1813 | 1814 | 1815 | 1816 | 1817 | 1818 | 1819 | 1820 | 1821 | 1822 | 1823 | 1824 | 1825 | 1826 | 1827 | 1828 | 1829 | 1830 | 1831 | 1832 | 1833 | 1834 | 1835 | 1836 | 1837 | 1838 | 1839 | 1840 | 1841 | 1842 | 1843 | 1844 | 1845 | 1846 | 1847 | 1848 | 1849 | 1850 | 1851 | 1852 | 1853 | 1854 | 1855 | 1856 | 1857 | 1858 | 1859 | 1860 | 1861 | 1862 | 1863 | 1864 | 1865 | 1866 | 1867 | 1868 | 1869 | 1870 | 1871 | 1872 | 1873 | 1874 | 1875 | 1876 | 1877 | 1878 | 1879 | 1880 | 1881 | 1882 | 1883 | 1884 | 1885 | 1886 | 1887 | 1888 | 1889 | 1890 | 1891 | 1892 | 1893 | 1894 | 1895 | 1896 | 1897 | 1898 | 1899 | 1900 | 1901 | 1902 | 1903 | 1904 | 1905 | 1906 | 1907 | 1908 | 1909 | 1910 | 1911 | 1912 | 1913 | 1914 | 1915 | 1916 | 1917 | 1918 | 1919 | 1920 | 1921 | 1922 | 1923 | 1924 | 1925 | 1926 | 1927 | 1928 | 1929 | 1930 | 1931 | 1932 | 1933 | 1934 | 1935 | 1936 | 1937 | 1938 | 1939 | 1940 | 1941 | 1942 | 1943 | 1944 | 1945 | 1946 | 1947 | 1948 | 1949 | 1950 | 1951 | 1952 | 1953 | 1954 | 1955 | 1956 | 1957 | 1958 | 1959 | 1960 | 1961 | 1962 | 1963 | 1964 | 1965 | 1966 | 1967 | 1968 | 1969 | 1970 | 1971 | 1972 | 1973 | 1974 | 1975 | 1976 | 1977 | 1978 | 1979 | 1980 | 1981 | 1982 | 1983 | 1984 | 1985 | 1986 | 1987 | 1988 | 1989 | 1990 | 1991 | 1992 | 1993 | 1994 | 1995 | 1996 | 1997 | 1998 | 1999 | 2000 | 2001 | 2002 | 2003 | 2004 | 2005 | 2006 | 2007 | 2008 | 2009 | 2010 | 2011 | 2012 | 2013 | 2014 | 2015 | 2016 | 2017 | 2018 | 2019 | 2020 | 2021 | 2022 | 2023 | 2024 | 2025 | 2026 | 2027 | 2028 | 2029 | 2030 | 2031 | 2032 | 2033 | 2034 | 2035 | 2036 | 2037 | 2038 | 2039 | 2040 | 2041 | 2042 | 2043 | 2044 | 2045 | 2046 | 2047
1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 2 | 1 | 2 | 2 | 0 | 0 | 2 | 2 | 1 | 2 | 4 | 3 | 5 | 4 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 3 | 1 | 0 | 2 | 2 | 0 | 1 | 0 | 2 | 0 | 1 | 2 | 1 | 0 | 1 | 1 | 0 | 2 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 2 | 1 | 1 | 0 | 0 | 2 | 2 | 0 | 0 | 1 | 2 | 0 | 2 | 2 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 2 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 2 | 1 | 2 | 1 | 1 | 1 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 4 | 0 | 2 | 3 | 2 | 0 | 0 | 1 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 2 | 5 | 3 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 2 | 1 | 1 | 2 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 2 | 0 | 1 | 1 | 2 | 1 | 0 | 1 | 0 | 1 | 1 | 2 | 2 | 0 | 0 | 2 | 3 | 2 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 2 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 3 | 2 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 2 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 2 | 1 | 1 | 0 | 0 | 4 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 2 | 2365 | 189 | 192 | 220 | 214 | 224 | 254 | 287 | 266 | 335 | 381 | 426 | 433 | 478 | 499 | 533 | 604 | 670 | 739 | 832 | 903 | 946 | 1041 | 1122 | 1251 | 1381 | 1484 | 1576 | 1708 | 1901 | 2100 | 2222 | 2347 | 2502 | 2657 | 2802 | 2936 | 3059 | 3155 | 3211 | 3391 | 3492 | 3469 | 3499 | 3527 | 3532 | 3549 | 3600 | 3646 | 3682 | 3711 | 3937 | 3710 | 3641 | 3646 | 3633 | 3648 | 3717 | 3663 | 3689 | 3669 | 3631 | 3981 | 3644 | 3693 | 3638 | 3698 | 3634 | 3626 | 3649 | 3720 | 3693 | 3680 | 3892 | 3738 | 3689 | 3664 | 3640 | 3733 | 3826 | 3806 | 3740 | 3792 | 3817 | 3830 | 3729 | 3808 | 3807 | 3703 | 3704 | 3702 | 3674 | 3713 | 3817 | 3835 | 3771 | 3793 | 3656 | 3596 | 3620 | 3613 | 3555 | 3527 | 3552 | 3661 | 3662 | 3565 | 3559 | 3555 | 3525 | 3524 | 3522 | 3558 | 3585 | 3566 | 3552 | 3206 | 3130 | 3198 | 3159 | 3152 | 3157 | 3140 | 3209 | 3239 | 3143 | 3211 | 3217 | 3238 | 3250 | 3277 | 3197 | 3205 | 3267 | 3327 | 3269 | 3270 | 3354 | 3380 | 3320 | 3244 | 3248 | 3275 | 3279 | 3297 | 3295 | 3324 | 3395 | 3387 | 3380 | 3346 | 3362 | 3393 | 3453 | 3503 | 3476 | 3548 | 3541 | 3506 | 3557 | 3665 | 3680 | 3667 | 3678 | 3783 | 3799 | 3867 | 3930 | 3938 | 3996 | 3999 | 4033 | 4057 | 4137 | 4270 | 4378 | 4371 | 4332 | 4404 | 4487 | 4531 | 4613 | 4707 | 4722 | 4736 | 4748 | 4758 | 4781 | 4817 | 4850 | 4826 | 4816 | 4850 | 4811 | 4762 | 4805 | 4813 | 4737 | 4684 | 4697 | 4639 | 4583 | 4560 | 4440 | 4407 | 4487 | 4325 | 4271 | 4364 | 4302 | 4302 | 4127 | 4148 | 4071 | 4046 | 4087 | 3945 | 3931 | 3971 | 3849 | 3809 | 3750 | 3757 | 3706 | 3680 | 3637 | 3691 | 3649 | 3562 | 3529 | 3512 | 3503 | 3446 | 3376 | 3435 | 3368 | 3308 | 3326 | 3315 | 3281 | 3249 | 3216 | 3173 | 3181 | 3161 | 3125 | 3121 | 3071 | 3005 | 3008 | 3013 | 3023 | 3074 | 3097 | 3017 | 2941 | 2855 | 2841 | 2865 | 2844 | 2766 | 2757 | 2805 | 2806 | 2739 | 2720 | 2750 | 2782 | 2764 | 2719 | 2643 | 2685 | 2706 | 2652 | 2651 | 2609 | 2528 | 2569 | 2517 | 2514 | 2583 | 2497 | 2459 | 2444 | 2469 | 2437 | 2448 | 2477 | 2415 | 2368 | 2361 | 2378 | 2369 | 2380 | 2356 | 2316 | 2310 | 2271 | 2253 | 2212 | 2180 | 2220 | 2204 | 2207 | 2227 | 2237 | 2220 | 2203 | 2239 | 2211 | 2186 | 2193 | 2184 | 2140 | 2090 | 2060 | 2101 | 2114 | 2043 | 1992 | 2082 | 2084 | 2051 | 2019 | 1950 | 1954 | 1969 | 1966 | 1929 | 1905 | 1896 | 1842 | 1827 | 1798 | 1792 | 1819 | 1763 | 1741 | 1726 | 1724 | 1716 | 1677 | 1649 | 1650 | 1654 | 1662 | 1647 | 1613 | 1626 | 1638 | 1631 | 1631 | 1606 | 1566 | 1518 | 1530 | 1511 | 1509 | 1466 | 1432 | 1457 | 1416 | 1433 | 1417 | 1466 | 1418 | 1422 | 1432 | 1392 | 1378 | 1356 | 1313 | 1374 | 1397 | 1376 | 1430 | 1364 | 1310 | 1327 | 1352 | 1343 | 1365 | 1371 | 1321 | 1374 | 1359 | 1331 | 1321 | 1320 | 1334 | 1331 | 1303 | 1269 | 1245 | 1256 | 1284 | 1301 | 1297 | 1269 | 1279 | 1308 | 1291 | 1283 | 1288 | 1252 | 1221 | 1228 | 1245 | 1280 | 1276 | 1249 | 1228 | 1236 | 1238 | 1192 | 1230 | 1245 | 1237 | 1237 | 1214 | 1273 | 1266 | 1259 | 1230 | 1267 | 1231 | 1272 | 1247 | 1322 | 1280 | 1282 | 1276 | 1244 | 1323 | 1379 | 1376 | 1380 | 1438 | 1458 | 1470 | 1519 | 1535 | 1564 | 1619 | 1684 | 1763 | 1843 | 1875 | 1951 | 2049 | 2146 | 2261 | 2391 | 2556 | 2685 | 2786 | 2941 | 3115 | 3295 | 3501 | 3746 | 4029 | 4270 | 4533 | 4873 | 5171 | 5552 | 5950 | 6360 | 6688 | 7037 | 7568 | 8249 | 8860 | 9393 | 10055 | 10656 | 11401 | 11937 | 12665 | 13234 | 13921 | 14586 | 15366 | 16162 | 16716 | 17510 | 18327 | 18796 | 19440 | 20318 | 20950 | 21614 | 22286 | 22734 | 23132 | 23610 | 24062 | 24431 | 24662 | 24687 | 24721 | 25156 | 24696 | 24613 | 24639 | 24352 | 23967 | 23652 | 23488 | 22874 | 21965 | 21374 | 21033 | 20301 | 19571 | 18858 | 17694 | 16907 | 16130 | 15345 | 14513 | 13690 | 12786 | 11893 | 11195 | 10432 | 9793 | 9073 | 8303 | 7587 | 7055 | 6498 | 5955 | 5474 | 4998 | 4564 | 4232 | 3916 | 3708 | 3491 | 3156 | 2918 | 2769 | 2622 | 2452 | 2268 | 2129 | 2084 | 2071 | 2020 | 1950 | 1875 | 1810 | 1788 | 1765 | 1708 | 1650 | 1624 | 1671 | 1687 | 1687 | 1674 | 1650 | 1588 | 1588 | 1609 | 1527 | 1572 | 1561 | 1569 | 1546 | 1577 | 1559 | 1529 | 1604 | 1537 | 1499 | 1509 | 1623 | 1554 | 1583 | 1564 | 1537 | 1520 | 1562 | 1571 | 1517 | 1533 | 1543 | 1563 | 1540 | 1513 | 1553 | 1567 | 1501 | 1474 | 1511 | 1539 | 1541 | 1507 | 1505 | 1512 | 1508 | 1485 | 1451 | 1438 | 1439 | 1431 | 1415 | 1415 | 1451 | 1454 | 1421 | 1441 | 1458 | 1427 | 1416 | 1428 | 1397 | 1392 | 1408 | 1390 | 1405 | 1406 | 1432 | 1434 | 1466 | 1378 | 1413 | 1428 | 1359 | 1369 | 1414 | 1403 | 1404 | 1380 | 1363 | 1365 | 1370 | 1369 | 1392 | 1415 | 1389 | 1388 | 1410 | 1421 | 1440 | 1440 | 1411 | 1554 | 1412 | 1448 | 1453 | 1440 | 1444 | 1460 | 1458 | 1455 | 1502 | 1522 | 1534 | 1568 | 1570 | 1567 | 1605 | 1636 | 1600 | 1622 | 1674 | 1730 | 1698 | 1681 | 1598 | 1661 | 1705 | 1706 | 1699 | 1722 | 1731 | 1760 | 1779 | 1775 | 1761 | 1715 | 1719 | 1735 | 1751 | 1779 | 1889 | 1850 | 1796 | 1822 | 1794 | 1781 | 1802 | 1788 | 1780 | 1766 | 1755 | 1779 | 1773 | 1732 | 1708 | 1679 | 1717 | 1653 | 1591 | 1695 | 1668 | 1616 | 1587 | 1614 | 1547 | 1562 | 1536 | 1475 | 1482 | 1506 | 1477 | 1451 | 1445 | 1449 | 1389 | 1383 | 1383 | 1374 | 1348 | 1335 | 1373 | 1319 | 1272 | 1275 | 1244 | 1252 | 1238 | 1209 | 1179 | 1162 | 1164 | 1161 | 1172 | 1184 | 1144 | 1113 | 1115 | 1114 | 1110 | 1145 | 1114 | 1053 | 1054 | 1039 | 999 | 1012 | 1039 | 1037 | 1005 | 971 | 988 | 1007 | 988 | 962 | 970 | 991 | 960 | 934 | 941 | 928 | 908 | 913 | 887 | 916 | 897 | 887 | 878 | 858 | 841 | 855 | 861 | 845 | 863 | 855 | 836 | 781 | 765 | 804 | 812 | 815 | 769 | 792 | 793 | 784 | 794 | 762 | 750 | 749 | 757 | 747 | 734 | 742 | 748 | 744 | 735 | 700 | 707 | 731 | 730 | 696 | 663 | 660 | 687 | 707 | 682 | 668 | 676 | 664 | 652 | 639 | 620 | 608 | 593 | 594 | 602 | 599 | 602 | 594 | 611 | 582 | 557 | 566 | 569 | 566 | 553 | 545 | 535 | 534 | 528 | 502 | 508 | 522 | 508 | 503 | 523 | 517 | 521 | 507 | 490 | 464 | 510 | 527 | 506 | 505 | 502 | 477 | 480 | 473 | 462 | 468 | 457 | 458 | 449 | 442 | 432 | 417 | 404 | 430 | 410 | 433 | 418 | 408 | 400 | 394 | 414 | 412 | 416 | 409 | 407 | 420 | 408 | 382 | 374 | 383 | 381 | 382 | 390 | 397 | 407 | 391 | 385 | 374 | 387 | 374 | 361 | 370 | 368 | 369 | 362 | 363 | 395 | 416 | 397 | 376 | 378 | 358 | 358 | 385 | 378 | 361 | 352 | 347 | 344 | 349 | 357 | 366 | 369 | 365 | 364 | 365 | 354 | 349 | 372 | 377 | 365 | 385 | 384 | 350 | 360 | 378 | 366 | 366 | 360 | 376 | 357 | 359 | 386 | 405 | 411 | 415 | 430 | 432 | 420 | 414 | 461 | 458 | 461 | 468 | 439 | 485 | 459 | 481 | 482 | 524 | 533 | 546 | 562 | 570 | 596 | 627 | 632 | 660 | 709 | 695 | 727 | 789 | 829 | 836 | 833 | 878 | 904 | 942 | 1008 | 1079 | 1122 | 1155 | 1192 | 1208 | 1234 | 1312 | 1381 | 1465 | 1525 | 1548 | 1596 | 1749 | 1720 | 1744 | 1827 | 1900 | 1929 | 1965 | 2021 | 2108 | 2156 | 2218 | 2319 | 2374 | 2391 | 2419 | 2488 | 2499 | 2597 | 2599 | 2544 | 2869 | 2777 | 2815 | 2719 | 2798 | 2790 | 2785 | 2802 | 2757 | 2748 | 2779 | 2791 | 2757 | 2751 | 2723 | 2655 | 2668 | 2692 | 2632 | 2579 | 2589 | 2408 | 2363 | 2281 | 2252 | 2238 | 2177 | 2055 | 1966 | 1897 | 1845 | 1789 | 1719 | 1640 | 1557 | 1476 | 1394 | 1354 | 1323 | 1322 | 1261 | 1091 | 1039 | 1013 | 931 | 869 | 824 | 791 | 751 | 685 | 643 | 617 | 575 | 537 | 503 | 470 | 419 | 396 | 352 | 389 | 353 | 312 | 300 | 277 | 249 | 225 | 216 | 219 | 200 | 177 | 179 | 184 | 180 | 177 | 171 | 145 | 140 | 139 | 150 | 138 | 137 | 119 | 126 | 137 | 124 | 124 | 123 | 109 | 108 | 115 | 113 | 114 | 117 | 109 | 109 | 105 | 101 | 103 | 93 | 101 | 120 | 111 | 117 | 106 | 96 | 94 | 103 | 107 | 107 | 108 | 110 | 108 | 104 | 95 | 93 | 107 | 106 | 98 | 104 | 101 | 91 | 102 | 100 | 94 | 95 | 92 | 89 | 93 | 102 | 106 | 103 | 91 | 93 | 100 | 98 | 95 | 94 | 104 | 94 | 85 | 80 | 93 | 89 | 97 | 91 | 83 | 85 | 91 | 97 | 92 | 83 | 84 | 97 | 93 | 86 | 80 | 83 | 96 | 100 | 90 | 94 | 104 | 97 | 81 | 85 | 88 | 93 | 99 | 91 | 85 | 86 | 81 | 80 | 79 | 81 | 92 | 93 | 90 | 91 | 88 | 84 | 90 | 87 | 87 | 93 | 92 | 91 | 86 | 83 | 84 | 85 | 92 | 96 | 94 | 95 | 90 | 82 | 78 | 88 | 81 | 73 | 71 | 69 | 73 | 76 | 66 | 66 | 76 | 78 | 76 | 77 | 77 | 81 | 72 | 71 | 76 | 69 | 72 | 78 | 77 | 78 | 69 | 74 | 84 | 86 | 79 | 63 | 60 | 68 | 67 | 69 | 72 | 72 | 68 | 72 | 71 | 73 | 67 | 72 | 64 | 58 | 65 | 62 | 54 | 60 | 66 | 60 | 58 | 59 | 56 | 56 | 59 | 61 | 60 | 65 | 65 | 67 | 61 | 54 | 53 | 54 | 51 | 47 | 52 | 60 | 62 | 55 | 48 | 51 | 52 | 52 | 50 | 46 | 54 | 48 | 56 | 58 | 52 | 48 | 42 | 44 | 46 | 50 | 50 | 48 | 49 | 54 | 55 | 47 | 44 | 44 | 44 | 50 | 44 | 45 | 47 | 57 | 53 | 48 | 48 | 37 | 41 | 52 | 52 | 43 | 45 | 47 | 42 | 47 | 51 | 40 | 42 | 42 | 41 | 43 | 47 | 41 | 42 | 37 | 40 | 39 | 35 | 32 | 35 | 38 | 36 | 33 | 31 | 34 | 40 | 37 | 33 | 39 | 41 | 35 | 44 | 48 | 42 | 40 | 43 | 42 | 35 | 31 | 36 | 40 | 37 | 38 | 44 | 41 | 37 | 40 | 40 | 34 | 40 | 39 | 30 | 30 | 32 | 37 | 44 | 40 | 33 | 31 | 36 | 37 | 34 | 36 | 43 | 38 | 41 | 48 | 39 | 30 | 32 | 49 | 40 | 39 | 34 | 28 | 31 | 33 | 28 | 26 | 36 | 34 | 30 | 37 | 40 | 33 | 28 | 26 | 22 | 25 | 27 | 32 | 37 | 27 | 32 | 33 | 34 | 44 | 36 | 29 | 31 | 35 | 34 | 33 | 35 | 37 | 35 | 30 | 27 | 30 | 39 | 34 | 35 | 38 | 33 | 33 | 37 | 41 | 44 | 38 | 33 | 32 | 38 | 30 | 39 | 43 | 41 | 48 | 50 | 33 | 40 | 37 | 32 | 48 | 39 | 44 | 45 | 43 | 48 | 44 | 42 | 41 | 37 | 49 | 37 | 42 | 42 | 46 | 49 | 44 | 41 | 45 | 37 | 41 | 41 | 32 | 38 | 43 | 48 | 42 | 42 | 46 | 48 | 46 | 44 | 43 | 43 | 52 | 51 | 48 | 48 | 53 | 50 | 44 | 41 | 44 | 50 | 48 | 50 | 53 | 55 | 49 | 45 | 52 | 55 | 54 | 49 | 53 | 53 | 48 | 46 | 46 | 48 | 45 | 51 | 54 | 55 | 50 | 49 | 58 | 47 | 52 | 63 | 50 | 56 | 59 | 50 | 49 | 51 | 49 | 47 | 46 | 45 | 39 | 34 | 33 | 44 | 37 | 35 | 42 | 38 | 34 | 28 | 36 | 45 | 27 | 40 | 33 | 32 | 22 | 25 | 23 | 29 | 31 | 25 | 28 | 24 | 23 | 26 | 24 | 22 | 27 | 22 | 23 | 30 | 30 | 21 | 17 | 22 | 20 | 16 | 19 | 18 | 17 | 21 | 18 | 15 | 18 | 18 | 18 | 22 | 16 | 12 | 15 | 12 | 12 | 17 | 19 | 15 | 10 | 9 | 4 | 9 | 15 | 9 | 16 | 9 | 11 | 15 | 18 | 8 | 14 | 8 | 16 | 12 | 11 | 8 | 6 | 9 | 11 | 9 | 12 | 8 | 5 | 9 | 13 | 9 | 11 | 11 | 9 | 9 | 6 | 6 | 8 | 10 | 9 | 7 | 9 | 11 | 7 | 6 | 7 | 10 | 11 | 9 | 10 | 9 | 9 | 10 | 10 | 8 | 8 | 8 | 9 | 10 | 9 | 9 | 10 | 7 | 9 | 11 | 8 | 7 | 10 | 11 | 13 | 10 | 7 | 7 | 6 | 7 | 9 | 6 | 8 | 8 | 5 | 10 | 9 | 11 | 13 | 13 | 8 | 11 | 6 | 8 | 8 | 5 | 9 | 7 | 8 | 6 | 8 | 13 | 6 | 6 | 9 | 6 | 11 | 5 | 6 | 6 | 9 | 11 | 9 | 7 | 8 | 8 | 6 | 6 | 7 | 8 | 10 | 7 | 11 | 10 | 7 | 8 | 7 | 8 | 8 | 8 | 7 | 7 | 7 | 5 | 4 | 5 | 5 | 9 | 7 | 4 | 5 | 6 | 7 | 4 | 5 | 6 | 6 | 5 | 8 | 8 | 8 | 2 | 3 | 4 | 5 | 7 | 4 | 6 | 4 | 1 | 5 | 4 | 2 | 7 | 9 | 4 | 3 | 6 | 12 | 7 | 3 | 4 | 4 | 7 | 5 | 7 | 8 | 6 | 3 | 7 | 9 | 5 | 7 | 5 | 6 | 4 | 6 | 5 | 4 | 2 | 2 | 4 | 3 | 5 | 6 | 4 | 2 | 2 | 3 | 6 | 7 | 3 | 2 | 3 | 4 | 4 | 5 | 5 | 4 | 4 | 5 | 6 | 4 | 3 | 4 | 3 | 3 | 4 | 6 | 3 | 2 | 3 | 4 | 3 | 2 | 4 | 4 | 3 | 3 | 5 | 8 | 7 | 5 | 1 | 5 | 3 | 4 | 4 | 3 | 3 | 5 | 3 | 7 | 1 | 3 | 4 | 3 | 2 | 2 | 2 | 4 | 1 | 1 | 4 | 6 | 3 | 4 | 4 | 2 | 3 | 2 | 3 | 3 | 4 | 4 | 2 | 5 | 1 | 3 | 5 | 2 | 2 | 2 | 3 | 3 | 4 | 4 | 4 | 2 | 4 | 4 | 1 | 3 | 2 | 1 | 2 | 4 | 3 | 4 | 3 | 2 | 2 | 2 | 3 | 3 | 4 | 3 | 1 | 3 | 3 | 2 | 1 | 3 | 1 | 4 | 2 | 5 | 3 | 4 | 1 | 0 | 7 | 1 | 2 | 2 | 5 | 4 | 2 | 3 | 1 | 2 | 4 | 0 | 0 | 3 | 0 | 3 | 0 | 2 | 2 | 1 | 3 | 6 | 4 | 3 | 4 | 3 | 1 | 3 | 2 | 4 | 5 | 4 | 3 | 2 | 1 | 1 | 1 | 2 | 3 | 3 | 4 | 3 | 3 | 4 | 2 | 1 | 2 | 2 | 1 | 2 | 1 | 0 | 5 | 4 | 5 | 4 | 4 | 4 | 3 | 5 | 1 | 2 | 0 | 2 | 3 | 4 | 2 | 2 | 2 | 1 | 3 | 3 | 3 | 2 | 1 | 3 | 5 | 2 | 1 | 2 | 4 | 3 | 3 | 4 | 4 | 8 | 3 | 4 | 0 | 3 | 1 | 2 | 2 | 3 | 2 | 2 | 5 | 3 | 3 | 2 | 3 | 1 | 3 | 4 | 4 | 4 | 4 | 5 | 4 | 2 | 0 | 1 | 3 | 3 | 3 | 3 | 3 | 3 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 3 | 2276 | 195 | 198 | 204 | 233 | 238 | 236 | 253 | 287 | 316 | 354 | 393 | 429 | 473 | 517 | 551 | 582 | 629 | 727 | 823 | 906 | 984 | 1014 | 1123 | 1253 | 1392 | 1521 | 1624 | 1736 | 1887 | 2081 | 2208 | 2282 | 2422 | 2586 | 2713 | 2841 | 2984 | 3149 | 3253 | 3384 | 3499 | 3475 | 3478 | 3492 | 3479 | 3477 | 3583 | 3691 | 3689 | 3672 | 3801 | 3654 | 3596 | 3565 | 3581 | 3564 | 3572 | 3605 | 3551 | 3570 | 3591 | 3928 | 3749 | 3700 | 3661 | 3682 | 3684 | 3662 | 3670 | 3677 | 3636 | 3781 | 3826 | 3586 | 3672 | 3717 | 3655 | 3599 | 3662 | 3720 | 3709 | 3734 | 3773 | 3792 | 3732 | 3799 | 3729 | 3689 | 3730 | 3727 | 3712 | 3697 | 3746 | 3835 | 3722 | 3665 | 3612 | 3609 | 3621 | 3542 | 3523 | 3541 | 3489 | 3596 | 3668 | 3584 | 3560 | 3530 | 3539 | 3545 | 3515 | 3457 | 3451 | 3463 | 3449 | 3240 | 3123 | 3193 | 3160 | 3123 | 3128 | 3217 | 3191 | 3169 | 3156 | 3217 | 3156 | 3228 | 3253 | 3244 | 3181 | 3271 | 3253 | 3215 | 3165 | 3271 | 3310 | 3283 | 3262 | 3222 | 3183 | 3262 | 3282 | 3258 | 3295 | 3288 | 3288 | 3337 | 3360 | 3346 | 3318 | 3396 | 3448 | 3413 | 3387 | 3448 | 3491 | 3575 | 3580 | 3565 | 3525 | 3578 | 3668 | 3719 | 3764 | 3780 | 3921 | 3912 | 3969 | 4047 | 4101 | 4159 | 4207 | 4217 | 4206 | 4273 | 4365 | 4475 | 4485 | 4533 | 4604 | 4631 | 4695 | 4722 | 4714 | 4763 | 4798 | 4806 | 4754 | 4780 | 4849 | 4808 | 4773 | 4784 | 4787 | 4738 | 4709 | 4700 | 4629 | 4582 | 4539 | 4397 | 4409 | 4435 | 4401 | 4397 | 4347 | 4320 | 4259 | 4209 | 4136 | 4115 | 4089 | 4068 | 4002 | 3846 | 3974 | 3900 | 3853 | 3819 | 3823 | 3699 | 3682 | 3722 | 3625 | 3611 | 3570 | 3611 | 3566 | 3442 | 3397 | 3411 | 3387 | 3391 | 3333 | 3241 | 3283 | 3304 | 3276 | 3202 | 3178 | 3183 | 3137 | 3126 | 3119 | 3093 | 3083 | 3035 | 3016 | 3037 | 2991 | 2930 | 2923 | 2920 | 2890 | 2889 | 2855 | 2859 | 2859 | 2819 | 2832 | 2785 | 2747 | 2813 | 2817 | 2758 | 2703 | 2667 | 2794 | 2720 | 2699 | 2651 | 2622 | 2570 | 2574 | 2548 | 2650 | 2587 | 2563 | 2518 | 2547 | 2463 | 2442 | 2417 | 2446 | 2407 | 2362 | 2371 | 2427 | 2360 | 2324 | 2386 | 2400 | 2329 | 2300 | 2288 | 2271 | 2308 | 2307 | 2258 | 2237 | 2211 | 2194 | 2195 | 2224 | 2198 | 2148 | 2125 | 2136 | 2165 | 2123 | 2092 | 2063 | 2087 | 2141 | 2121 | 2066 | 2068 | 2067 | 2068 | 2070 | 2019 | 1988 | 1974 | 2016 | 2011 | 1974 | 1942 | 1886 | 1890 | 1866 | 1835 | 1828 | 1851 | 1814 | 1770 | 1764 | 1700 | 1673 | 1710 | 1714 | 1686 | 1662 | 1615 | 1599 | 1634 | 1639 | 1617 | 1635 | 1644 | 1606 | 1565 | 1522 | 1528 | 1536 | 1476 | 1528 | 1493 | 1496 | 1490 | 1439 | 1467 | 1426 | 1411 | 1481 | 1410 | 1387 | 1394 | 1426 | 1392 | 1359 | 1413 | 1415 | 1373 | 1364 | 1380 | 1378 | 1360 | 1326 | 1337 | 1348 | 1321 | 1308 | 1324 | 1312 | 1310 | 1298 | 1289 | 1292 | 1315 | 1303 | 1308 | 1307 | 1302 | 1295 | 1282 | 1269 | 1237 | 1244 | 1297 | 1279 | 1253 | 1255 | 1250 | 1251 | 1256 | 1306 | 1293 | 1241 | 1252 | 1231 | 1228 | 1247 | 1245 | 1255 | 1251 | 1241 | 1230 | 1229 | 1236 | 1251 | 1240 | 1229 | 1281 | 1264 | 1287 | 1274 | 1284 | 1216 | 1273 | 1297 | 1313 | 1343 | 1367 | 1363 | 1410 | 1392 | 1433 | 1456 | 1519 | 1561 | 1527 | 1585 | 1685 | 1726 | 1774 | 1833 | 1872 | 2034 | 2198 | 2265 | 2319 | 2398 | 2558 | 2745 | 2913 | 3093 | 3283 | 3536 | 3723 | 3874 | 4164 | 4493 | 4889 | 5260 | 5534 | 5832 | 6309 | 6811 | 7197 | 7759 | 8421 | 8844 | 9400 | 10131 | 10731 | 11212 | 11870 | 12532 | 13071 | 13803 | 14587 | 15374 | 16078 | 16722 | 17365 | 18113 | 18949 | 19528 | 20236 | 20956 | 21461 | 21942 | 22374 | 22981 | 23531 | 23865 | 24216 | 24542 | 24581 | 24585 | 24695 | 24703 | 24605 | 24336 | 24161 | 24046 | 23825 | 23683 | 22730 | 21899 | 21387 | 20733 | 19974 | 19413 | 18641 | 17595 | 16959 | 16066 | 15116 | 14326 | 13383 | 12620 | 11910 | 11126 | 10292 | 9634 | 8890 | 8276 | 7669 | 7018 | 6500 | 6031 | 5508 | 4899 | 4545 | 4244 | 3863 | 3602 | 3336 | 3080 | 2838 | 2670 | 2570 | 2457 | 2358 | 2232 | 2068 | 1972 | 1916 | 1868 | 1839 | 1812 | 1784 | 1762 | 1713 | 1683 | 1698 | 1675 | 1633 | 1621 | 1617 | 1636 | 1629 | 1668 | 1664 | 1613 | 1600 | 1580 | 1594 | 1597 | 1546 | 1566 | 1553 | 1576 | 1561 | 1558 | 1584 | 1585 | 1528 | 1568 | 1566 | 1544 | 1542 | 1567 | 1522 | 1549 | 1506 | 1608 | 1561 | 1540 | 1497 | 1506 | 1538 | 1510 | 1460 | 1496 | 1477 | 1459 | 1451 | 1454 | 1458 | 1444 | 1436 | 1446 | 1448 | 1443 | 1444 | 1433 | 1424 | 1423 | 1421 | 1423 | 1431 | 1423 | 1418 | 1423 | 1391 | 1402 | 1433 | 1389 | 1369 | 1396 | 1402 | 1364 | 1335 | 1366 | 1395 | 1423 | 1427 | 1357 | 1342 | 1364 | 1344 | 1371 | 1417 | 1380 | 1351 | 1337 | 1342 | 1361 | 1359 | 1372 | 1402 | 1408 | 1409 | 1382 | 1383 | 1410 | 1446 | 1426 | 1462 | 1455 | 1489 | 1467 | 1465 | 1525 | 1496 | 1489 | 1518 | 1554 | 1587 | 1589 | 1603 | 1604 | 1579 | 1608 | 1586 | 1666 | 1664 | 1627 | 1663 | 1642 | 1621 | 1662 | 1717 | 1724 | 1728 | 1713 | 1697 | 1722 | 1751 | 1752 | 1755 | 1765 | 1767 | 1764 | 1742 | 1784 | 1817 | 1756 | 1731 | 1744 | 1780 | 1772 | 1737 | 1741 | 1729 | 1691 | 1680 | 1701 | 1730 | 1680 | 1663 | 1662 | 1657 | 1656 | 1706 | 1656 | 1630 | 1600 | 1524 | 1551 | 1558 | 1560 | 1581 | 1570 | 1505 | 1461 | 1460 | 1453 | 1431 | 1400 | 1353 | 1352 | 1310 | 1323 | 1295 | 1391 | 1338 | 1294 | 1304 | 1296 | 1240 | 1248 | 1250 | 1277 | 1242 | 1201 | 1171 | 1147 | 1163 | 1145 | 1143 | 1149 | 1157 | 1117 | 1093 | 1085 | 1035 | 1034 | 1031 | 992 | 996 | 969 | 945 | 962 | 974 | 956 | 961 | 979 | 980 | 964 | 937 | 923 | 932 | 915 | 912 | 942 | 925 | 899 | 919 | 895 | 859 | 817 | 797 | 831 | 860 | 832 | 793 | 814 | 834 | 834 | 824 | 792 | 802 | 820 | 832 | 770 | 766 | 755 | 750 | 756 | 770 | 733 | 753 | 752 | 715 | 718 | 737 | 735 | 705 | 678 | 668 | 685 | 699 | 667 | 686 | 684 | 660 | 677 | 668 | 666 | 659 | 625 | 610 | 620 | 634 | 626 | 620 | 615 | 590 | 604 | 619 | 625 | 592 | 623 | 599 | 576 | 566 | 561 | 564 | 559 | 547 | 541 | 549 | 548 | 539 | 521 | 512 | 513 | 520 | 517 | 513 | 515 | 526 | 499 | 468 | 486 | 500 | 473 | 457 | 461 | 478 | 476 | 457 | 446 | 450 | 464 | 462 | 452 | 437 | 440 | 465 | 429 | 436 | 423 | 438 | 426 | 421 | 416 | 419 | 412 | 396 | 411 | 415 | 414 | 413 | 392 | 387 | 388 | 377 | 393 | 388 | 383 | 403 | 368 | 377 | 342 | 358 | 389 | 396 | 400 | 385 | 378 | 385 | 379 | 368 | 368 | 363 | 343 | 336 | 350 | 357 | 389 | 398 | 364 | 347 | 339 | 330 | 338 | 337 | 330 | 344 | 358 | 354 | 347 | 350 | 354 | 359 | 372 | 360 | 356 | 364 | 355 | 373 | 372 | 359 | 354 | 369 | 392 | 374 | 365 | 365 | 379 | 398 | 391 | 376 | 385 | 406 | 409 | 416 | 429 | 414 | 486 | 463 | 441 | 466 | 492 | 479 | 504 | 502 | 503 | 539 | 565 | 577 | 587 | 580 | 597 | 641 | 658 | 706 | 752 | 766 | 806 | 805 | 832 | 909 | 954 | 963 | 984 | 1020 | 1058 | 1099 | 1155 | 1217 | 1259 | 1300 | 1372 | 1407 | 1415 | 1483 | 1550 | 1654 | 1751 | 1745 | 1790 | 1871 | 1887 | 1958 | 2054 | 2106 | 2104 | 2134 | 2250 | 2343 | 2366 | 2391 | 2469 | 2478 | 2574 | 2610 | 2618 | 2816 | 2742 | 2725 | 2693 | 2768 | 2833 | 2786 | 2737 | 2729 | 2740 | 2789 | 2791 | 2728 | 2722 | 2723 | 2660 | 2615 | 2513 | 2576 | 2597 | 2620 | 2359 | 2336 | 2240 | 2216 | 2116 | 2077 | 1996 | 1930 | 1901 | 1848 | 1771 | 1699 | 1628 | 1567 | 1517 | 1468 | 1418 | 1327 | 1282 | 1213 | 1120 | 1032 | 938 | 902 | 862 | 815 | 745 | 699 | 681 | 655 | 619 | 573 | 524 | 483 | 438 | 413 | 403 | 372 | 340 | 321 | 315 | 288 | 271 | 247 | 228 | 220 | 208 | 187 | 176 | 174 | 162 | 160 | 157 | 149 | 144 | 137 | 148 | 146 | 139 | 128 | 127 | 114 | 116 | 108 | 105 | 108 | 107 | 111 | 124 | 127 | 123 | 117 | 108 | 111 | 109 | 97 | 95 | 126 | 112 | 115 | 93 | 95 | 118 | 106 | 98 | 93 | 95 | 100 | 101 | 94 | 97 | 109 | 107 | 100 | 94 | 93 | 88 | 98 | 95 | 84 | 95 | 98 | 95 | 93 | 88 | 86 | 88 | 93 | 102 | 106 | 98 | 92 | 91 | 93 | 95 | 96 | 93 | 89 | 95 | 104 | 100 | 88 | 82 | 85 | 92 | 92 | 89 | 87 | 90 | 91 | 91 | 101 | 88 | 101 | 94 | 91 | 93 | 102 | 99 | 99 | 96 | 90 | 94 | 89 | 93 | 87 | 87 | 86 | 81 | 86 | 87 | 90 | 86 | 91 | 83 | 81 | 89 | 78 | 81 | 74 | 87 | 88 | 88 | 74 | 71 | 79 | 88 | 96 | 88 | 74 | 79 | 89 | 82 | 80 | 80 | 85 | 92 | 87 | 82 | 74 | 79 | 86 | 80 | 79 | 83 | 83 | 75 | 68 | 76 | 79 | 78 | 82 | 76 | 65 | 64 | 74 | 61 | 67 | 78 | 66 | 62 | 69 | 67 | 67 | 74 | 70 | 60 | 64 | 68 | 64 | 66 | 65 | 72 | 60 | 70 | 55 | 70 | 68 | 59 | 75 | 63 | 64 | 58 | 53 | 54 | 66 | 65 | 54 | 51 | 53 | 57 | 55 | 57 | 55 | 49 | 63 | 62 | 60 | 66 | 64 | 63 | 51 | 47 | 53 | 57 | 51 | 51 | 52 | 48 | 49 | 52 | 53 | 52 | 46 | 47 | 51 | 51 | 51 | 50 | 47 | 45 | 45 | 46 | 44 | 47 | 53 | 52 | 51 | 51 | 51 | 49 | 46 | 51 | 50 | 40 | 41 | 39 | 43 | 47 | 34 | 36 | 34 | 34 | 41 | 38 | 35 | 40 | 41 | 43 | 43 | 35 | 39 | 52 | 48 | 42 | 45 | 43 | 40 | 34 | 42 | 38 | 35 | 37 | 40 | 35 | 38 | 49 | 47 | 38 | 42 | 41 | 39 | 38 | 36 | 35 | 38 | 42 | 45 | 43 | 41 | 39 | 41 | 39 | 38 | 41 | 38 | 33 | 33 | 34 | 38 | 38 | 34 | 32 | 35 | 41 | 38 | 33 | 35 | 38 | 31 | 32 | 31 | 31 | 34 | 40 | 42 | 37 | 41 | 39 | 32 | 31 | 24 | 26 | 43 | 39 | 34 | 31 | 32 | 29 | 29 | 36 | 37 | 32 | 31 | 36 | 38 | 37 | 29 | 32 | 36 | 34 | 36 | 33 | 35 | 36 | 40 | 32 | 31 | 32 | 35 | 37 | 33 | 31 | 36 | 37 | 31 | 32 | 33 | 33 | 39 | 46 | 37 | 34 | 41 | 39 | 39 | 34 | 39 | 44 | 44 | 43 | 35 | 33 | 36 | 32 | 29 | 32 | 33 | 38 | 36 | 39 | 39 | 38 | 39 | 36 | 38 | 40 | 39 | 38 | 37 | 32 | 35 | 42 | 46 | 40 | 38 | 47 | 41 | 37 | 37 | 37 | 36 | 55 | 34 | 36 | 48 | 56 | 43 | 47 | 34 | 41 | 49 | 45 | 43 | 44 | 49 | 42 | 41 | 39 | 54 | 47 | 51 | 48 | 51 | 45 | 36 | 45 | 51 | 51 | 58 | 40 | 43 | 45 | 47 | 43 | 41 | 47 | 56 | 51 | 52 | 58 | 54 | 50 | 56 | 57 | 49 | 47 | 49 | 48 | 53 | 49 | 47 | 48 | 47 | 48 | 47 | 39 | 43 | 42 | 44 | 45 | 42 | 40 | 44 | 48 | 45 | 48 | 45 | 42 | 37 | 36 | 42 | 37 | 51 | 35 | 29 | 29 | 35 | 30 | 44 | 38 | 27 | 28 | 37 | 32 | 24 | 29 | 19 | 29 | 23 | 29 | 26 | 29 | 27 | 23 | 24 | 29 | 18 | 22 | 17 | 12 | 19 | 26 | 21 | 14 | 13 | 17 | 18 | 19 | 15 | 12 | 14 | 17 | 13 | 12 | 13 | 11 | 12 | 14 | 12 | 18 | 18 | 11 | 8 | 11 | 14 | 11 | 12 | 11 | 11 | 10 | 11 | 11 | 10 | 9 | 7 | 9 | 8 | 10 | 7 | 9 | 9 | 9 | 6 | 12 | 5 | 8 | 8 | 9 | 17 | 9 | 10 | 8 | 12 | 6 | 8 | 10 | 11 | 10 | 6 | 10 | 8 | 13 | 10 | 11 | 9 | 9 | 10 | 16 | 13 | 10 | 10 | 10 | 9 | 7 | 11 | 8 | 7 | 10 | 8 | 9 | 8 | 8 | 6 | 8 | 6 | 7 | 8 | 9 | 11 | 8 | 7 | 10 | 11 | 7 | 8 | 9 | 7 | 9 | 6 | 7 | 12 | 9 | 8 | 8 | 8 | 6 | 9 | 4 | 4 | 11 | 8 | 8 | 4 | 6 | 8 | 6 | 5 | 6 | 7 | 6 | 10 | 7 | 4 | 9 | 5 | 5 | 7 | 5 | 5 | 6 | 9 | 7 | 3 | 8 | 9 | 8 | 4 | 7 | 6 | 4 | 6 | 6 | 7 | 10 | 5 | 7 | 8 | 11 | 6 | 4 | 6 | 7 | 8 | 10 | 7 | 3 | 4 | 6 | 5 | 8 | 10 | 7 | 6 | 6 | 6 | 6 | 4 | 8 | 9 | 8 | 5 | 3 | 4 | 7 | 11 | 9 | 4 | 4 | 4 | 5 | 7 | 6 | 2 | 4 | 7 | 7 | 4 | 5 | 6 | 4 | 7 | 3 | 7 | 4 | 4 | 8 | 6 | 3 | 7 | 2 | 5 | 4 | 2 | 7 | 3 | 4 | 5 | 5 | 3 | 2 | 1 | 4 | 2 | 2 | 3 | 4 | 2 | 0 | 1 | 3 | 4 | 4 | 4 | 4 | 3 | 4 | 4 | 7 | 7 | 4 | 4 | 6 | 5 | 4 | 3 | 3 | 6 | 5 | 2 | 3 | 4 | 4 | 3 | 3 | 4 | 4 | 5 | 5 | 5 | 4 | 4 | 1 | 2 | 4 | 2 | 3 | 3 | 3 | 2 | 1 | 3 | 2 | 2 | 2 | 3 | 2 | 3 | 3 | 3 | 4 | 7 | 3 | 3 | 6 | 3 | 3 | 1 | 2 | 4 | 1 | 2 | 4 | 1 | 5 | 3 | 4 | 0 | 0 | 1 | 2 | 4 | 3 | 5 | 1 | 1 | 0 | 1 | 3 | 3 | 2 | 1 | 2 | 3 | 1 | 2 | 3 | 3 | 3 | 2 | 2 | 3 | 3 | 3 | 2 | 3 | 2 | 1 | 1 | 2 | 3 | 4 | 2 | 2 | 3 | 3 | 3 | 3 | 5 | 2 | 2 | 2 | 1 | 2 | 1 | 1 | 3 | 2 | 0 | 3 | 2 | 2 | 1 | 1 | 1 | 2 | 4 | 5 | 2 | 4 | 3 | 3 | 1 | 3 | 2 | 3 | 3 | 3 | 5 | 2 | 3 | 5 | 4 | 3 | 1 | 3 | 3 | 2 | 4 | 1 | 1 | 2 | 1 | 2 | 3 | 1 | 1 | 2 | 0 | 2 | 3 | 3 | 2 | 2 | 2 | 2 | 3 | 2 | 2 | 4 | 2 | 3 | 2 | 2 | 2 | 2 | 2 | 3 | 2 | 2 | 4 | 4 | 4 | 2 | 3 | 5 | 2 | 5 | 2 | 2 | 3 | 2 | 3 | 3 | 2 | 5 | 3 | 1 | 4 | 3 | 3 | 2 | 6 | 2 | 4 | 3 | 5 | 1 | 4 | 2 | 4 | 2 | 4 | 2 | 1 | 4 | 5 | 5 | 6 | 5 | 3 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 4 | 2130 | 173 | 191 | 219 | 208 | 232 | 242 | 268 | 285 | 322 | 360 | 396 | 429 | 442 | 486 | 549 | 604 | 656 | 728 | 795 | 912 | 955 | 1017 | 1096 | 1204 | 1336 | 1457 | 1616 | 1765 | 1869 | 2048 | 2194 | 2296 | 2462 | 2599 | 2681 | 2811 | 2941 | 3075 | 3205 | 3337 | 3435 | 3455 | 3477 | 3466 | 3445 | 3513 | 3585 | 3583 | 3547 | 3548 | 3720 | 3710 | 3638 | 3632 | 3513 | 3584 | 3666 | 3587 | 3629 | 3672 | 3598 | 3855 | 3612 | 3598 | 3642 | 3622 | 3513 | 3583 | 3568 | 3663 | 3628 | 3657 | 3662 | 3652 | 3710 | 3702 | 3648 | 3720 | 3701 | 3637 | 3695 | 3782 | 3848 | 3746 | 3654 | 3678 | 3623 | 3621 | 3676 | 3680 | 3682 | 3687 | 3784 | 3822 | 3709 | 3653 | 3604 | 3594 | 3559 | 3558 | 3533 | 3498 | 3481 | 3555 | 3603 | 3534 | 3456 | 3398 | 3421 | 3439 | 3452 | 3480 | 3457 | 3380 | 3379 | 3223 | 3102 | 3107 | 3118 | 3118 | 3137 | 3176 | 3211 | 3119 | 3069 | 3175 | 3188 | 3158 | 3184 | 3181 | 3169 | 3197 | 3173 | 3208 | 3225 | 3231 | 3198 | 3177 | 3206 | 3212 | 3237 | 3231 | 3167 | 3195 | 3309 | 3339 | 3280 | 3311 | 3338 | 3331 | 3366 | 3241 | 3327 | 3351 | 3310 | 3442 | 3354 | 3508 | 3529 | 3527 | 3594 | 3595 | 3523 | 3648 | 3710 | 3829 | 3843 | 3848 | 3952 | 3990 | 4003 | 4024 | 4126 | 4230 | 4258 | 4297 | 4350 | 4433 | 4425 | 4484 | 4514 | 4549 | 4652 | 4703 | 4747 | 4743 | 4742 | 4795 | 4734 | 4670 | 4686 | 4716 | 4725 | 4667 | 4614 | 4644 | 4618 | 4542 | 4568 | 4597 | 4566 | 4418 | 4336 | 4342 | 4306 | 4283 | 4241 | 4163 | 4151 | 4129 | 4034 | 3997 | 4040 | 4049 | 3948 | 3879 | 3776 | 3853 | 3842 | 3722 | 3645 | 3562 | 3576 | 3626 | 3546 | 3516 | 3515 | 3468 | 3477 | 3388 | 3370 | 3332 | 3292 | 3312 | 3325 | 3244 | 3182 | 3198 | 3187 | 3160 | 3145 | 3153 | 3174 | 3136 | 3095 | 3059 | 2970 | 2921 | 2946 | 2960 | 2951 | 2916 | 2889 | 2896 | 2880 | 2893 | 2877 | 2798 | 2745 | 2783 | 2867 | 2816 | 2747 | 2730 | 2656 | 2657 | 2658 | 2594 | 2651 | 2677 | 2592 | 2574 | 2580 | 2572 | 2549 | 2519 | 2482 | 2508 | 2484 | 2469 | 2482 | 2381 | 2356 | 2361 | 2394 | 2369 | 2397 | 2328 | 2293 | 2334 | 2340 | 2360 | 2316 | 2252 | 2283 | 2335 | 2295 | 2269 | 2252 | 2242 | 2223 | 2183 | 2220 | 2167 | 2147 | 2150 | 2141 | 2161 | 2077 | 2082 | 2056 | 2088 | 2076 | 2059 | 2039 | 2015 | 2048 | 2086 | 2029 | 1967 | 1980 | 1970 | 1917 | 1930 | 1990 | 1999 | 1907 | 1887 | 1865 | 1831 | 1827 | 1848 | 1848 | 1814 | 1791 | 1761 | 1751 | 1733 | 1701 | 1702 | 1694 | 1647 | 1615 | 1624 | 1618 | 1581 | 1564 | 1589 | 1589 | 1530 | 1491 | 1497 | 1495 | 1492 | 1514 | 1487 | 1468 | 1454 | 1414 | 1411 | 1412 | 1418 | 1436 | 1390 | 1439 | 1440 | 1422 | 1401 | 1386 | 1416 | 1439 | 1367 | 1406 | 1366 | 1396 | 1326 | 1304 | 1329 | 1357 | 1327 | 1341 | 1338 | 1314 | 1326 | 1319 | 1259 | 1261 | 1302 | 1289 | 1292 | 1306 | 1300 | 1274 | 1250 | 1246 | 1259 | 1258 | 1235 | 1228 | 1230 | 1214 | 1229 | 1247 | 1238 | 1242 | 1245 | 1249 | 1248 | 1218 | 1217 | 1238 | 1227 | 1208 | 1221 | 1228 | 1248 | 1207 | 1171 | 1227 | 1252 | 1196 | 1171 | 1195 | 1266 | 1242 | 1210 | 1245 | 1278 | 1234 | 1288 | 1283 | 1286 | 1290 | 1282 | 1350 | 1378 | 1364 | 1414 | 1409 | 1437 | 1491 | 1600 | 1627 | 1663 | 1742 | 1779 | 1872 | 1958 | 1935 | 2036 | 2202 | 2343 | 2426 | 2540 | 2746 | 2887 | 3051 | 3271 | 3514 | 3758 | 3922 | 4182 | 4454 | 4824 | 5257 | 5607 | 5991 | 6367 | 6866 | 7371 | 7708 | 8234 | 8863 | 9427 | 9996 | 10642 | 11221 | 11914 | 12682 | 13265 | 13961 | 14684 | 15377 | 16100 | 16803 | 17460 | 17957 | 18625 | 19415 | 20162 | 20700 | 21426 | 22067 | 22445 | 22866 | 23230 | 23501 | 23879 | 24200 | 24357 | 24380 | 24400 | 24403 | 24284 | 24170 | 23918 | 23544 | 23213 | 22843 | 22151 | 21348 | 20845 | 20306 | 19721 | 18936 | 18117 | 17107 | 16375 | 15415 | 14655 | 13861 | 13104 | 12255 | 11328 | 10636 | 9824 | 9228 | 8515 | 7885 | 7336 | 6723 | 6127 | 5521 | 5168 | 4789 | 4353 | 4092 | 3698 | 3423 | 3185 | 2982 | 2756 | 2577 | 2425 | 2340 | 2206 | 2124 | 2054 | 1976 | 1902 | 1824 | 1778 | 1741 | 1710 | 1708 | 1686 | 1683 | 1688 | 1655 | 1643 | 1648 | 1618 | 1616 | 1612 | 1592 | 1605 | 1603 | 1603 | 1540 | 1556 | 1530 | 1511 | 1522 | 1590 | 1591 | 1578 | 1584 | 1558 | 1536 | 1560 | 1533 | 1532 | 1510 | 1520 | 1522 | 1549 | 1507 | 1492 | 1500 | 1526 | 1533 | 1506 | 1544 | 1516 | 1498 | 1496 | 1511 | 1474 | 1439 | 1437 | 1424 | 1411 | 1412 | 1435 | 1443 | 1423 | 1411 | 1393 | 1377 | 1377 | 1417 | 1453 | 1417 | 1394 | 1409 | 1412 | 1387 | 1388 | 1380 | 1366 | 1341 | 1369 | 1391 | 1368 | 1364 | 1384 | 1347 | 1345 | 1406 | 1419 | 1392 | 1359 | 1336 | 1343 | 1329 | 1310 | 1325 | 1360 | 1377 | 1392 | 1401 | 1397 | 1412 | 1413 | 1405 | 1383 | 1374 | 1359 | 1370 | 1460 | 1434 | 1446 | 1442 | 1443 | 1513 | 1507 | 1463 | 1472 | 1503 | 1507 | 1501 | 1509 | 1534 | 1564 | 1557 | 1520 | 1585 | 1586 | 1680 | 1747 | 1640 | 1714 | 1714 | 1654 | 1689 | 1692 | 1695 | 1684 | 1690 | 1696 | 1728 | 1753 | 1734 | 1755 | 1751 | 1745 | 1728 | 1734 | 1783 | 1768 | 1739 | 1759 | 1737 | 1723 | 1753 | 1767 | 1731 | 1681 | 1665 | 1674 | 1685 | 1686 | 1678 | 1686 | 1643 | 1591 | 1625 | 1660 | 1673 | 1629 | 1579 | 1578 | 1547 | 1484 | 1458 | 1473 | 1477 | 1481 | 1469 | 1423 | 1401 | 1379 | 1361 | 1376 | 1362 | 1349 | 1308 | 1297 | 1281 | 1284 | 1263 | 1249 | 1205 | 1198 | 1226 | 1226 | 1215 | 1179 | 1164 | 1174 | 1183 | 1190 | 1125 | 1117 | 1105 | 1137 | 1094 | 1124 | 1105 | 1081 | 1012 | 1068 | 1011 | 1040 | 1004 | 985 | 983 | 986 | 973 | 944 | 936 | 941 | 931 | 899 | 900 | 930 | 876 | 909 | 903 | 919 | 876 | 861 | 858 | 872 | 871 | 872 | 863 | 829 | 795 | 795 | 809 | 816 | 820 | 798 | 796 | 797 | 787 | 786 | 784 | 764 | 769 | 762 | 747 | 736 | 721 | 726 | 729 | 726 | 713 | 700 | 709 | 698 | 705 | 692 | 675 | 653 | 678 | 698 | 649 | 674 | 649 | 656 | 657 | 661 | 647 | 633 | 635 | 630 | 619 | 622 | 631 | 604 | 554 | 578 | 585 | 579 | 601 | 569 | 571 | 551 | 513 | 546 | 573 | 530 | 533 | 528 | 517 | 521 | 525 | 534 | 525 | 498 | 496 | 486 | 502 | 506 | 510 | 500 | 469 | 475 | 496 | 489 | 473 | 441 | 434 | 462 | 469 | 458 | 445 | 437 | 438 | 445 | 441 | 433 | 419 | 424 | 415 | 432 | 434 | 411 | 424 | 437 | 421 | 401 | 402 | 426 | 443 | 423 | 378 | 363 | 401 | 411 | 365 | 391 | 389 | 413 | 375 | 375 | 370 | 370 | 394 | 376 | 348 | 376 | 375 | 347 | 353 | 360 | 332 | 334 | 366 | 361 | 349 | 350 | 357 | 370 | 352 | 340 | 369 | 320 | 338 | 349 | 343 | 348 | 347 | 334 | 340 | 349 | 337 | 340 | 346 | 352 | 374 | 370 | 364 | 350 | 355 | 366 | 367 | 376 | 379 | 371 | 375 | 379 | 378 | 375 | 366 | 370 | 382 | 394 | 405 | 411 | 421 | 433 | 465 | 456 | 449 | 449 | 473 | 483 | 495 | 524 | 538 | 552 | 552 | 547 | 590 | 626 | 612 | 628 | 670 | 681 | 700 | 759 | 840 | 845 | 890 | 874 | 903 | 985 | 1017 | 1069 | 1102 | 1128 | 1159 | 1216 | 1287 | 1336 | 1376 | 1437 | 1556 | 1550 | 1580 | 1675 | 1806 | 1814 | 1916 | 1937 | 1948 | 2029 | 2106 | 2103 | 2112 | 2137 | 2188 | 2295 | 2367 | 2391 | 2469 | 2543 | 2581 | 2643 | 2713 | 2810 | 2764 | 2698 | 2759 | 2785 | 2800 | 2826 | 2825 | 2815 | 2829 | 2819 | 2765 | 2728 | 2736 | 2697 | 2608 | 2608 | 2594 | 2536 | 2557 | 2526 | 2392 | 2303 | 2254 | 2215 | 2136 | 2093 | 2030 | 1990 | 1894 | 1820 | 1797 | 1714 | 1625 | 1558 | 1481 | 1333 | 1340 | 1246 | 1250 | 1210 | 1061 | 1013 | 950 | 909 | 813 | 805 | 727 | 688 | 659 | 596 | 537 | 491 | 477 | 486 | 428 | 387 | 359 | 337 | 325 | 299 | 262 | 248 | 229 | 231 | 229 | 221 | 202 | 186 | 183 | 174 | 161 | 156 | 148 | 139 | 137 | 142 | 142 | 127 | 130 | 136 | 125 | 123 | 131 | 127 | 118 | 124 | 124 | 108 | 100 | 111 | 117 | 112 | 108 | 116 | 116 | 109 | 109 | 101 | 112 | 125 | 100 | 101 | 103 | 98 | 108 | 103 | 100 | 92 | 91 | 103 | 104 | 98 | 92 | 95 | 101 | 113 | 104 | 105 | 85 | 90 | 89 | 95 | 100 | 103 | 88 | 86 | 90 | 96 | 89 | 88 | 97 | 94 | 93 | 100 | 106 | 93 | 87 | 101 | 92 | 88 | 87 | 81 | 88 | 93 | 97 | 94 | 84 | 89 | 96 | 94 | 89 | 83 | 80 | 87 | 90 | 86 | 80 | 87 | 82 | 77 | 89 | 92 | 90 | 89 | 94 | 87 | 84 | 91 | 91 | 84 | 73 | 77 | 92 | 92 | 78 | 82 | 99 | 97 | 77 | 93 | 89 | 79 | 85 | 81 | 81 | 72 | 69 | 77 | 83 | 89 | 85 | 76 | 78 | 87 | 88 | 81 | 84 | 92 | 72 | 80 | 75 | 83 | 76 | 73 | 75 | 88 | 87 | 81 | 79 | 73 | 72 | 72 | 71 | 74 | 78 | 77 | 70 | 73 | 79 | 71 | 67 | 69 | 67 | 69 | 67 | 70 | 71 | 67 | 65 | 69 | 68 | 68 | 74 | 70 | 56 | 61 | 66 | 65 | 62 | 65 | 65 | 55 | 57 | 63 | 63 | 63 | 62 | 59 | 55 | 52 | 56 | 55 | 57 | 58 | 65 | 61 | 56 | 44 | 50 | 49 | 39 | 46 | 45 | 49 | 52 | 53 | 50 | 50 | 55 | 53 | 55 | 53 | 54 | 53 | 48 | 51 | 47 | 40 | 44 | 47 | 47 | 45 | 53 | 47 | 40 | 40 | 40 | 39 | 42 | 43 | 43 | 46 | 43 | 45 | 49 | 49 | 37 | 28 | 34 | 39 | 37 | 37 | 42 | 41 | 40 | 41 | 45 | 46 | 45 | 43 | 47 | 42 | 38 | 40 | 40 | 30 | 41 | 45 | 43 | 43 | 44 | 45 | 41 | 43 | 42 | 35 | 32 | 33 | 36 | 30 | 36 | 37 | 37 | 36 | 34 | 40 | 32 | 35 | 43 | 35 | 35 | 32 | 34 | 37 | 38 | 36 | 31 | 32 | 42 | 44 | 36 | 32 | 29 | 41 | 45 | 31 | 41 | 38 | 38 | 38 | 37 | 29 | 30 | 34 | 34 | 39 | 42 | 38 | 34 | 33 | 30 | 30 | 34 | 37 | 36 | 36 | 38 | 37 | 34 | 38 | 37 | 33 | 32 | 34 | 35 | 35 | 35 | 31 | 26 | 29 | 31 | 36 | 30 | 28 | 29 | 34 | 33 | 34 | 34 | 33 | 36 | 33 | 28 | 28 | 33 | 35 | 32 | 32 | 31 | 40 | 32 | 32 | 39 | 38 | 41 | 43 | 40 | 42 | 38 | 40 | 30 | 36 | 42 | 31 | 38 | 44 | 44 | 40 | 36 | 39 | 35 | 36 | 30 | 35 | 32 | 30 | 32 | 41 | 50 | 38 | 35 | 33 | 32 | 36 | 36 | 39 | 43 | 44 | 41 | 33 | 28 | 35 | 45 | 43 | 39 | 42 | 34 | 27 | 35 | 48 | 45 | 41 | 42 | 45 | 45 | 42 | 49 | 43 | 45 | 42 | 42 | 40 | 51 | 50 | 53 | 44 | 46 | 42 | 63 | 50 | 38 | 41 | 45 | 53 | 45 | 54 | 58 | 46 | 46 | 56 | 50 | 55 | 47 | 40 | 54 | 62 | 40 | 61 | 45 | 46 | 46 | 47 | 48 | 50 | 38 | 47 | 45 | 45 | 55 | 51 | 40 | 41 | 43 | 37 | 42 | 43 | 38 | 40 | 44 | 46 | 40 | 35 | 40 | 36 | 34 | 30 | 35 | 36 | 33 | 28 | 29 | 37 | 35 | 27 | 26 | 25 | 25 | 23 | 20 | 26 | 29 | 24 | 23 | 23 | 22 | 24 | 23 | 22 | 24 | 17 | 14 | 13 | 18 | 24 | 19 | 20 | 18 | 16 | 3 | 10 | 14 | 10 | 19 | 16 | 12 | 13 | 18 | 16 | 12 | 8 | 12 | 11 | 18 | 16 | 8 | 11 | 10 | 8 | 9 | 7 | 12 | 13 | 15 | 7 | 6 | 12 | 12 | 9 | 9 | 11 | 5 | 10 | 12 | 10 | 9 | 11 | 12 | 12 | 11 | 10 | 11 | 11 | 10 | 5 | 5 | 7 | 7 | 8 | 7 | 6 | 9 | 13 | 12 | 7 | 6 | 12 | 11 | 8 | 14 | 13 | 6 | 7 | 10 | 9 | 7 | 11 | 8 | 10 | 3 | 5 | 11 | 8 | 9 | 5 | 4 | 8 | 4 | 10 | 9 | 10 | 7 | 13 | 11 | 7 | 4 | 9 | 8 | 5 | 4 | 6 | 7 | 10 | 10 | 10 | 9 | 9 | 6 | 7 | 4 | 8 | 10 | 10 | 8 | 8 | 6 | 4 | 7 | 10 | 8 | 5 | 6 | 6 | 8 | 9 | 9 | 7 | 8 | 10 | 9 | 7 | 4 | 6 | 7 | 5 | 9 | 12 | 8 | 7 | 6 | 7 | 7 | 5 | 12 | 11 | 9 | 11 | 7 | 6 | 6 | 8 | 6 | 4 | 4 | 6 | 10 | 10 | 7 | 3 | 7 | 1 | 4 | 9 | 8 | 3 | 8 | 8 | 7 | 6 | 8 | 4 | 6 | 8 | 8 | 9 | 6 | 7 | 4 | 6 | 7 | 1 | 3 | 4 | 8 | 7 | 9 | 4 | 5 | 3 | 4 | 4 | 3 | 4 | 3 | 4 | 4 | 5 | 5 | 3 | 3 | 5 | 5 | 3 | 4 | 5 | 4 | 2 | 3 | 4 | 4 | 3 | 1 | 3 | 4 | 2 | 4 | 3 | 4 | 6 | 5 | 4 | 4 | 6 | 5 | 4 | 3 | 5 | 4 | 8 | 9 | 2 | 5 | 3 | 2 | 3 | 6 | 5 | 8 | 5 | 5 | 4 | 4 | 5 | 3 | 2 | 3 | 5 | 3 | 1 | 2 | 1 | 5 | 3 | 3 | 2 | 1 | 4 | 5 | 4 | 3 | 4 | 4 | 3 | 2 | 4 | 5 | 1 | 2 | 5 | 3 | 2 | 3 | 2 | 3 | 3 | 1 | 4 | 4 | 3 | 2 | 3 | 6 | 5 | 3 | 1 | 2 | 2 | 1 | 2 | 1 | 1 | 2 | 2 | 5 | 4 | 2 | 3 | 2 | 2 | 1 | 3 | 3 | 1 | 4 | 1 | 6 | 1 | 0 | 3 | 2 | 3 | 2 | 3 | 2 | 3 | 1 | 1 | 0 | 2 | 1 | 2 | 0 | 1 | 1 | 1 | 0 | 1 | 2 | 0 | 1 | 1 | 3 | 5 | 3 | 3 | 2 | 3 | 5 | 3 | 3 | 4 | 4 | 4 | 2 | 2 | 1 | 0 | 1 | 2 | 4 | 2 | 1 | 3 | 2 | 2 | 3 | 4 | 2 | 1 | 1 | 2 | 2 | 1 | 1 | 3 | 2 | 2 | 3 | 1 | 1 | 3 | 3 | 1 | 2 | 1 | 3 | 4 | 3 | 2 | 2 | 2 | 3 | 1 | 1 | 2 | 1 | 4 | 0 | 4 | 2 | 7 | 2 | 2 | 2 | 4 | 1 | 1 | 3 | 2 | 5 | 2 | 2 | 3 | 1 | 3 | 6 | 2 | 6 | 4 | 3 | 4 | 2 | 3 | 2 | 2 | 2 | 4 | 3 | 3 | 3 | 1 | 1 | 3 | 3 | 2 | 3 | 4 | 3 | 3 | 3 | 3 | 2 | 3 | 2 | 4 | 3 | 4 | 3 | 3 | 2 | 2 | 3 | 3 | 3 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 5 | 2201 | 192 | 187 | 211 | 230 | 220 | 233 | 250 | 269 | 357 | 370 | 418 | 432 | 466 | 513 | 581 | 619 | 656 | 719 | 808 | 912 | 972 | 1081 | 1161 | 1266 | 1401 | 1513 | 1605 | 1744 | 1951 | 2163 | 2278 | 2329 | 2497 | 2629 | 2759 | 2948 | 3019 | 3044 | 3164 | 3336 | 3463 | 3472 | 3416 | 3396 | 3482 | 3473 | 3440 | 3530 | 3540 | 3505 | 3721 | 3696 | 3578 | 3539 | 3550 | 3621 | 3639 | 3613 | 3567 | 3597 | 3582 | 3773 | 3624 | 3534 | 3556 | 3594 | 3578 | 3581 | 3617 | 3541 | 3621 | 3586 | 3771 | 3706 | 3677 | 3661 | 3678 | 3588 | 3557 | 3657 | 3648 | 3616 | 3735 | 3854 | 3641 | 3710 | 3702 | 3700 | 3684 | 3718 | 3676 | 3627 | 3672 | 3732 | 3689 | 3651 | 3564 | 3534 | 3538 | 3517 | 3504 | 3504 | 3485 | 3541 | 3617 | 3445 | 3388 | 3458 | 3462 | 3433 | 3396 | 3406 | 3431 | 3434 | 3473 | 3256 | 3096 | 3084 | 3106 | 3152 | 3156 | 3143 | 3149 | 3159 | 3119 | 3097 | 3091 | 3171 | 3169 | 3167 | 3177 | 3224 | 3167 | 3135 | 3121 | 3129 | 3124 | 3135 | 3203 | 3304 | 3348 | 3278 | 3212 | 3252 | 3301 | 3275 | 3277 | 3313 | 3320 | 3271 | 3406 | 3447 | 3380 | 3316 | 3359 | 3489 | 3538 | 3476 | 3538 | 3509 | 3546 | 3593 | 3654 | 3652 | 3703 | 3784 | 3840 | 3965 | 3929 | 3970 | 4094 | 4117 | 4037 | 4149 | 4265 | 4254 | 4358 | 4419 | 4486 | 4531 | 4550 | 4612 | 4670 | 4693 | 4709 | 4746 | 4789 | 4771 | 4777 | 4780 | 4762 | 4745 | 4729 | 4682 | 4644 | 4659 | 4621 | 4566 | 4554 | 4564 | 4527 | 4457 | 4433 | 4394 | 4333 | 4215 | 4173 | 4131 | 4180 | 4261 | 4036 | 4090 | 4020 | 3953 | 3953 | 3902 | 3770 | 3829 | 3892 | 3730 | 3657 | 3616 | 3769 | 3664 | 3542 | 3555 | 3460 | 3462 | 3434 | 3419 | 3421 | 3402 | 3354 | 3257 | 3290 | 3301 | 3248 | 3206 | 3182 | 3188 | 3172 | 3145 | 3120 | 3055 | 2996 | 2980 | 2998 | 3015 | 2987 | 2935 | 2939 | 2964 | 2917 | 2890 | 2871 | 2856 | 2891 | 2854 | 2780 | 2788 | 2753 | 2706 | 2694 | 2727 | 2695 | 2655 | 2641 | 2604 | 2646 | 2668 | 2612 | 2562 | 2632 | 2616 | 2599 | 2547 | 2538 | 2489 | 2478 | 2489 | 2462 | 2416 | 2470 | 2510 | 2483 | 2421 | 2396 | 2433 | 2439 | 2365 | 2357 | 2307 | 2290 | 2347 | 2343 | 2305 | 2272 | 2256 | 2249 | 2266 | 2270 | 2165 | 2205 | 2264 | 2192 | 2147 | 2170 | 2131 | 2126 | 2124 | 2106 | 2175 | 2120 | 2077 | 2066 | 2068 | 2092 | 2093 | 2005 | 1974 | 2009 | 1987 | 1908 | 1889 | 1878 | 1908 | 1881 | 1834 | 1805 | 1823 | 1838 | 1829 | 1818 | 1826 | 1808 | 1812 | 1787 | 1730 | 1690 | 1655 | 1640 | 1646 | 1636 | 1606 | 1591 | 1592 | 1586 | 1584 | 1583 | 1565 | 1551 | 1537 | 1501 | 1480 | 1496 | 1485 | 1475 | 1413 | 1393 | 1408 | 1390 | 1433 | 1446 | 1433 | 1468 | 1448 | 1414 | 1446 | 1362 | 1387 | 1422 | 1416 | 1408 | 1369 | 1331 | 1379 | 1404 | 1339 | 1350 | 1376 | 1315 | 1338 | 1291 | 1294 | 1349 | 1335 | 1317 | 1290 | 1299 | 1276 | 1271 | 1277 | 1302 | 1269 | 1260 | 1275 | 1262 | 1256 | 1265 | 1255 | 1241 | 1271 | 1287 | 1267 | 1250 | 1253 | 1261 | 1223 | 1186 | 1177 | 1189 | 1225 | 1242 | 1228 | 1213 | 1191 | 1211 | 1229 | 1215 | 1197 | 1223 | 1218 | 1188 | 1228 | 1265 | 1306 | 1248 | 1198 | 1274 | 1266 | 1238 | 1284 | 1323 | 1373 | 1342 | 1371 | 1400 | 1433 | 1459 | 1504 | 1523 | 1554 | 1583 | 1651 | 1741 | 1842 | 1929 | 1960 | 2027 | 2100 | 2244 | 2419 | 2469 | 2599 | 2809 | 2974 | 3089 | 3325 | 3562 | 3792 | 4054 | 4397 | 4693 | 4981 | 5277 | 5666 | 6265 | 6680 | 7049 | 7495 | 7921 | 8489 | 8990 | 9525 | 10271 | 11089 | 11611 | 12282 | 12686 | 13332 | 14172 | 14910 | 15646 | 16407 | 16984 | 17730 | 18491 | 19149 | 19887 | 20602 | 21167 | 21681 | 22327 | 22756 | 23059 | 23424 | 23774 | 24002 | 24209 | 24421 | 24483 | 24481 | 24472 | 24316 | 24031 | 23614 | 23179 | 22968 | 22661 | 22008 | 21072 | 20432 | 19822 | 18983 | 18465 | 17701 | 16682 | 16048 | 15200 | 14129 | 13640 | 12458 | 11711 | 11119 | 10090 | 9584 | 9060 | 8062 | 7411 | 6861 | 6414 | 5882 | 5334 | 4898 | 4527 | 4229 | 3887 | 3548 | 3303 | 3086 | 2882 | 2690 | 2490 | 2367 | 2291 | 2158 | 2079 | 1994 | 1895 | 1828 | 1794 | 1794 | 1766 | 1716 | 1692 | 1676 | 1652 | 1641 | 1618 | 1609 | 1641 | 1640 | 1599 | 1568 | 1580 | 1583 | 1605 | 1598 | 1546 | 1535 | 1566 | 1567 | 1541 | 1534 | 1556 | 1556 | 1564 | 1513 | 1539 | 1565 | 1555 | 1554 | 1498 | 1510 | 1497 | 1518 | 1449 | 1429 | 1484 | 1517 | 1481 | 1456 | 1486 | 1473 | 1442 | 1398 | 1420 | 1442 | 1442 | 1425 | 1415 | 1460 | 1458 | 1413 | 1409 | 1424 | 1435 | 1433 | 1400 | 1404 | 1376 | 1359 | 1378 | 1362 | 1383 | 1421 | 1389 | 1344 | 1408 | 1437 | 1409 | 1373 | 1363 | 1329 | 1322 | 1405 | 1367 | 1378 | 1384 | 1378 | 1379 | 1355 | 1371 | 1379 | 1372 | 1383 | 1396 | 1378 | 1367 | 1373 | 1378 | 1374 | 1368 | 1352 | 1363 | 1386 | 1413 | 1409 | 1423 | 1509 | 1470 | 1443 | 1462 | 1451 | 1464 | 1495 | 1477 | 1458 | 1503 | 1546 | 1523 | 1519 | 1553 | 1515 | 1504 | 1543 | 1532 | 1554 | 1610 | 1673 | 1598 | 1604 | 1679 | 1706 | 1667 | 1651 | 1659 | 1674 | 1702 | 1697 | 1721 | 1758 | 1726 | 1715 | 1748 | 1749 | 1765 | 1745 | 1729 | 1814 | 1790 | 1734 | 1731 | 1752 | 1756 | 1732 | 1684 | 1673 | 1717 | 1721 | 1685 | 1661 | 1656 | 1661 | 1647 | 1636 | 1632 | 1614 | 1612 | 1594 | 1543 | 1580 | 1574 | 1506 | 1475 | 1482 | 1480 | 1447 | 1386 | 1372 | 1430 | 1417 | 1354 | 1333 | 1337 | 1363 | 1345 | 1276 | 1320 | 1316 | 1216 | 1211 | 1206 | 1208 | 1194 | 1151 | 1160 | 1166 | 1131 | 1109 | 1117 | 1123 | 1110 | 1122 | 1117 | 1099 | 1092 | 1047 | 1111 | 1056 | 1057 | 1033 | 1002 | 1004 | 1003 | 1001 | 972 | 949 | 939 | 953 | 978 | 956 | 922 | 938 | 913 | 858 | 877 | 916 | 904 | 883 | 889 | 878 | 859 | 831 | 839 | 851 | 863 | 854 | 825 | 829 | 846 | 823 | 783 | 757 | 758 | 810 | 835 | 772 | 740 | 757 | 765 | 769 | 763 | 728 | 720 | 720 | 713 | 724 | 716 | 691 | 697 | 685 | 662 | 681 | 660 | 648 | 683 | 660 | 662 | 644 | 650 | 612 | 627 | 650 | 636 | 613 | 577 | 570 | 597 | 617 | 624 | 604 | 599 | 584 | 611 | 564 | 583 | 591 | 612 | 563 | 576 | 563 | 557 | 509 | 485 | 502 | 508 | 506 | 493 | 483 | 515 | 518 | 486 | 484 | 465 | 481 | 495 | 496 | 480 | 474 | 464 | 458 | 468 | 471 | 466 | 445 | 441 | 436 | 442 | 462 | 440 | 421 | 425 | 420 | 421 | 411 | 425 | 429 | 409 | 396 | 403 | 403 | 406 | 409 | 403 | 401 | 391 | 385 | 374 | 346 | 363 | 370 | 372 | 381 | 389 | 374 | 376 | 397 | 369 | 377 | 358 | 365 | 374 | 362 | 353 | 350 | 345 | 336 | 332 | 337 | 334 | 340 | 333 | 330 | 357 | 380 | 348 | 343 | 376 | 367 | 356 | 358 | 345 | 346 | 356 | 342 | 336 | 346 | 342 | 347 | 348 | 335 | 328 | 317 | 333 | 331 | 347 | 364 | 353 | 358 | 361 | 352 | 349 | 374 | 387 | 385 | 394 | 398 | 390 | 403 | 416 | 402 | 408 | 428 | 460 | 481 | 458 | 455 | 470 | 511 | 519 | 524 | 550 | 552 | 558 | 581 | 616 | 654 | 654 | 673 | 674 | 725 | 756 | 801 | 873 | 806 | 840 | 940 | 993 | 1006 | 1004 | 1046 | 1082 | 1111 | 1161 | 1228 | 1326 | 1420 | 1397 | 1470 | 1532 | 1554 | 1580 | 1687 | 1794 | 1813 | 1822 | 1880 | 1920 | 2042 | 2089 | 2089 | 2149 | 2270 | 2376 | 2420 | 2437 | 2470 | 2528 | 2597 | 2622 | 2657 | 2607 | 2668 | 2764 | 2698 | 2773 | 2753 | 2737 | 2752 | 2720 | 2695 | 2717 | 2740 | 2713 | 2684 | 2654 | 2636 | 2602 | 2572 | 2498 | 2456 | 2392 | 2395 | 2366 | 2288 | 2191 | 2089 | 2024 | 1984 | 1960 | 1903 | 1812 | 1741 | 1702 | 1638 | 1542 | 1472 | 1396 | 1345 | 1253 | 1157 | 1161 | 1147 | 1024 | 990 | 922 | 813 | 780 | 728 | 700 | 675 | 643 | 595 | 558 | 512 | 464 | 443 | 411 | 375 | 357 | 369 | 321 | 288 | 286 | 248 | 250 | 229 | 222 | 203 | 183 | 176 | 171 | 162 | 157 | 151 | 148 | 153 | 146 | 126 | 111 | 112 | 118 | 113 | 111 | 110 | 99 | 105 | 116 | 125 | 122 | 111 | 105 | 106 | 103 | 96 | 94 | 97 | 97 | 104 | 97 | 105 | 116 | 111 | 99 | 101 | 96 | 92 | 84 | 84 | 91 | 93 | 94 | 96 | 95 | 93 | 89 | 88 | 99 | 104 | 86 | 100 | 104 | 101 | 98 | 89 | 90 | 104 | 87 | 92 | 92 | 94 | 98 | 96 | 92 | 90 | 93 | 88 | 87 | 99 | 88 | 79 | 81 | 85 | 87 | 89 | 84 | 87 | 91 | 93 | 93 | 87 | 85 | 87 | 85 | 86 | 91 | 92 | 97 | 91 | 83 | 90 | 91 | 90 | 89 | 92 | 94 | 94 | 90 | 81 | 78 | 81 | 82 | 77 | 74 | 85 | 90 | 94 | 90 | 94 | 85 | 78 | 80 | 75 | 79 | 74 | 72 | 80 | 80 | 80 | 80 | 83 | 83 | 80 | 81 | 75 | 83 | 85 | 94 | 72 | 69 | 72 | 81 | 75 | 67 | 74 | 67 | 75 | 71 | 76 | 78 | 77 | 80 | 72 | 66 | 68 | 64 | 66 | 65 | 68 | 66 | 64 | 71 | 56 | 60 | 62 | 57 | 61 | 67 | 66 | 66 | 65 | 60 | 58 | 61 | 62 | 60 | 59 | 60 | 62 | 60 | 61 | 60 | 58 | 56 | 58 | 64 | 61 | 58 | 63 | 65 | 59 | 53 | 55 | 61 | 57 | 65 | 61 | 58 | 62 | 64 | 51 | 56 | 69 | 61 | 55 | 53 | 46 | 47 | 54 | 52 | 51 | 49 | 46 | 47 | 40 | 37 | 51 | 63 | 43 | 49 | 38 | 42 | 38 | 50 | 46 | 53 | 49 | 42 | 47 | 52 | 49 | 45 | 42 | 48 | 62 | 52 | 43 | 51 | 41 | 40 | 46 | 46 | 51 | 46 | 38 | 41 | 46 | 46 | 46 | 47 | 46 | 49 | 49 | 44 | 41 | 39 | 42 | 45 | 49 | 40 | 34 | 37 | 39 | 41 | 40 | 37 | 35 | 34 | 34 | 34 | 33 | 33 | 36 | 39 | 42 | 46 | 36 | 28 | 39 | 30 | 35 | 39 | 40 | 42 | 32 | 32 | 41 | 40 | 33 | 33 | 32 | 39 | 43 | 31 | 26 | 40 | 39 | 40 | 38 | 32 | 44 | 33 | 27 | 34 | 35 | 33 | 32 | 31 | 31 | 31 | 34 | 35 | 36 | 37 | 34 | 37 | 41 | 31 | 32 | 37 | 34 | 33 | 36 | 35 | 31 | 30 | 32 | 34 | 34 | 31 | 28 | 28 | 26 | 32 | 37 | 38 | 33 | 33 | 37 | 32 | 30 | 36 | 39 | 35 | 37 | 36 | 35 | 34 | 31 | 33 | 27 | 27 | 28 | 38 | 41 | 37 | 37 | 39 | 40 | 36 | 33 | 37 | 36 | 36 | 51 | 36 | 33 | 27 | 31 | 41 | 32 | 33 | 38 | 32 | 39 | 34 | 41 | 41 | 48 | 40 | 34 | 46 | 34 | 38 | 42 | 42 | 45 | 42 | 42 | 36 | 38 | 38 | 37 | 37 | 41 | 42 | 44 | 43 | 39 | 35 | 41 | 45 | 45 | 44 | 42 | 44 | 49 | 52 | 49 | 43 | 41 | 39 | 50 | 52 | 40 | 40 | 48 | 50 | 53 | 48 | 45 | 48 | 45 | 47 | 44 | 49 | 44 | 43 | 42 | 46 | 53 | 56 | 52 | 49 | 39 | 50 | 54 | 56 | 49 | 56 | 50 | 43 | 47 | 51 | 43 | 53 | 44 | 43 | 48 | 50 | 43 | 44 | 48 | 41 | 43 | 42 | 39 | 41 | 40 | 44 | 39 | 39 | 37 | 36 | 29 | 30 | 36 | 35 | 40 | 35 | 33 | 33 | 30 | 30 | 28 | 23 | 25 | 35 | 30 | 23 | 29 | 28 | 27 | 27 | 21 | 21 | 19 | 14 | 16 | 18 | 19 | 21 | 19 | 17 | 18 | 19 | 24 | 23 | 16 | 10 | 10 | 14 | 18 | 17 | 16 | 17 | 10 | 12 | 13 | 11 | 8 | 10 | 17 | 7 | 13 | 10 | 7 | 15 | 11 | 6 | 9 | 9 | 6 | 7 | 10 | 6 | 10 | 6 | 11 | 14 | 12 | 8 | 9 | 13 | 9 | 10 | 9 | 8 | 7 | 11 | 12 | 12 | 8 | 8 | 8 | 9 | 7 | 10 | 12 | 8 | 8 | 10 | 13 | 14 | 11 | 9 | 9 | 9 | 10 | 13 | 11 | 10 | 7 | 9 | 11 | 8 | 9 | 12 | 8 | 10 | 6 | 4 | 6 | 9 | 9 | 8 | 6 | 8 | 8 | 8 | 5 | 4 | 5 | 5 | 7 | 9 | 6 | 8 | 6 | 6 | 12 | 7 | 4 | 6 | 5 | 10 | 11 | 8 | 6 | 7 | 7 | 8 | 5 | 12 | 5 | 11 | 11 | 9 | 6 | 8 | 6 | 8 | 9 | 6 | 5 | 8 | 12 | 9 | 10 | 10 | 7 | 7 | 7 | 12 | 12 | 12 | 11 | 7 | 6 | 5 | 9 | 7 | 4 | 5 | 7 | 6 | 6 | 7 | 7 | 5 | 6 | 6 | 6 | 7 | 5 | 6 | 6 | 6 | 2 | 4 | 5 | 7 | 5 | 3 | 7 | 9 | 6 | 5 | 4 | 7 | 4 | 4 | 3 | 5 | 3 | 4 | 7 | 7 | 5 | 3 | 4 | 4 | 4 | 4 | 7 | 4 | 5 | 5 | 5 | 4 | 6 | 6 | 2 | 7 | 7 | 3 | 4 | 2 | 2 | 3 | 3 | 4 | 6 | 5 | 3 | 3 | 3 | 5 | 3 | 2 | 3 | 3 | 3 | 5 | 6 | 5 | 4 | 3 | 5 | 6 | 5 | 3 | 3 | 3 | 2 | 2 | 3 | 3 | 5 | 5 | 3 | 4 | 3 | 3 | 2 | 2 | 4 | 3 | 7 | 2 | 4 | 5 | 3 | 4 | 1 | 2 | 7 | 5 | 3 | 2 | 0 | 4 | 5 | 2 | 4 | 3 | 1 | 4 | 3 | 3 | 2 | 2 | 3 | 2 | 3 | 2 | 1 | 3 | 4 | 2 | 2 | 2 | 2 | 3 | 4 | 4 | 3 | 2 | 2 | 2 | 3 | 3 | 4 | 3 | 1 | 3 | 4 | 2 | 1 | 3 | 3 | 1 | 2 | 3 | 2 | 1 | 3 | 6 | 5 | 2 | 3 | 3 | 2 | 2 | 3 | 3 | 3 | 3 | 2 | 3 | 2 | 1 | 3 | 3 | 2 | 4 | 2 | 0 | 4 | 2 | 2 | 5 | 3 | 2 | 6 | 3 | 4 | 2 | 1 | 2 | 1 | 2 | 2 | 2 | 4 | 1 | 3 | 3 | 1 | 3 | 2 | 3 | 2 | 3 | 2 | 4 | 3 | 4 | 2 | 2 | 3 | 2 | 3 | 3 | 4 | 1 | 2 | 1 | 0 | 1 | 3 | 4 | 3 | 2 | 2 | 1 | 2 | 3 | 3 | 2 | 1 | 1 | 2 | 2 | 2 | 4 | 3 | 4 | 3 | 2 | 3 | 3 | 2 | 0 | 1 | 1 | 3 | 3 | 2 | 2 | 4 | 2 | 3 | 3 | 3 | 3 | 2 | 2 | 6 | 2 | 2 | 1 | 2 | 3 | 1 | 3 | 5 | 6 | 4 | 6 | 2 | 2 | 6 | 1 | 1 | 6 | 0 | 2 | 3 | 0 | 2 | 2 | 3 | 2 | 2 | 2 | 3 | 3 | 4 | 5 | 4 | 2 | 2 | 3 | 2 | 2 | 2 | 1 | 3 | 3 | 3 | 2 | 2 | 3 | 2 | 4 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 6 | 379 | 30 | 41 | 22 | 36 | 39 | 46 | 56 | 42 | 39 | 75 | 71 | 89 | 113 | 126 | 143 | 175 | 204 | 255 | 289 | 334 | 380 | 474 | 568 | 617 | 686 | 787 | 897 | 1002 | 1098 | 1207 | 1295 | 1420 | 1382 | 1393 | 1391 | 1409 | 1437 | 1495 | 1491 | 1453 | 1509 | 1535 | 1560 | 1564 | 1557 | 1549 | 1548 | 1506 | 1511 | 1506 | 1515 | 1528 | 1543 | 1578 | 1520 | 1505 | 1459 | 1464 | 1491 | 1495 | 1508 | 1511 | 1506 | 1557 | 1596 | 1499 | 1501 | 1502 | 1480 | 1487 | 1520 | 1524 | 1499 | 1489 | 1568 | 1632 | 1562 | 1563 | 1536 | 1547 | 1605 | 1618 | 1579 | 1577 | 1618 | 1617 | 1596 | 1567 | 1553 | 1518 | 1513 | 1550 | 1568 | 1549 | 1486 | 1487 | 1519 | 1500 | 1467 | 1424 | 1417 | 1450 | 1460 | 1439 | 1424 | 1422 | 1454 | 1445 | 1420 | 1434 | 1441 | 1422 | 1410 | 1415 | 1427 | 1412 | 1387 | 1428 | 1362 | 1329 | 1299 | 1280 | 1266 | 1269 | 1341 | 1331 | 1317 | 1284 | 1327 | 1326 | 1290 | 1289 | 1287 | 1328 | 1336 | 1320 | 1315 | 1311 | 1326 | 1349 | 1367 | 1357 | 1335 | 1319 | 1320 | 1343 | 1338 | 1314 | 1304 | 1319 | 1375 | 1406 | 1375 | 1358 | 1386 | 1410 | 1412 | 1397 | 1415 | 1482 | 1504 | 1525 | 1557 | 1604 | 1622 | 1657 | 1631 | 1729 | 1677 | 1795 | 1811 | 1821 | 1841 | 1924 | 1880 | 1996 | 1936 | 1971 | 2038 | 2062 | 2057 | 2065 | 2094 | 2010 | 2004 | 2002 | 1997 | 1992 | 1975 | 1995 | 1918 | 1916 | 1913 | 1833 | 1813 | 1833 | 1872 | 1840 | 1764 | 1746 | 1707 | 1700 | 1704 | 1687 | 1658 | 1674 | 1674 | 1649 | 1601 | 1597 | 1588 | 1573 | 1547 | 1531 | 1480 | 1450 | 1440 | 1447 | 1452 | 1485 | 1397 | 1445 | 1371 | 1334 | 1270 | 1300 | 1347 | 1352 | 1363 | 1326 | 1270 | 1273 | 1271 | 1283 | 1350 | 1318 | 1236 | 1220 | 1218 | 1241 | 1300 | 1259 | 1223 | 1184 | 1161 | 1152 | 1147 | 1134 | 1108 | 1109 | 1123 | 1133 | 1119 | 1110 | 1127 | 1098 | 1045 | 1031 | 1049 | 1078 | 1048 | 1016 | 1038 | 1039 | 997 | 991 | 1015 | 992 | 978 | 995 | 995 | 1006 | 998 | 980 | 920 | 937 | 967 | 979 | 931 | 894 | 930 | 938 | 926 | 927 | 905 | 909 | 931 | 914 | 880 | 862 | 871 | 872 | 855 | 856 | 856 | 855 | 836 | 809 | 811 | 816 | 814 | 800 | 783 | 791 | 785 | 772 | 781 | 785 | 794 | 792 | 793 | 791 | 756 | 753 | 754 | 745 | 755 | 752 | 742 | 753 | 719 | 696 | 694 | 698 | 712 | 676 | 680 | 662 | 669 | 659 | 624 | 652 | 656 | 655 | 626 | 612 | 642 | 647 | 636 | 641 | 626 | 613 | 591 | 592 | 606 | 605 | 575 | 549 | 549 | 567 | 572 | 565 | 558 | 553 | 565 | 577 | 576 | 550 | 547 | 574 | 548 | 532 | 543 | 559 | 572 | 561 | 547 | 531 | 533 | 518 | 523 | 536 | 530 | 506 | 506 | 543 | 557 | 542 | 504 | 525 | 482 | 473 | 514 | 514 | 489 | 501 | 506 | 503 | 498 | 471 | 489 | 471 | 462 | 450 | 467 | 476 | 483 | 472 | 496 | 491 | 460 | 454 | 467 | 474 | 467 | 464 | 465 | 461 | 450 | 455 | 459 | 445 | 463 | 458 | 476 | 440 | 437 | 447 | 441 | 453 | 457 | 476 | 447 | 457 | 430 | 436 | 450 | 477 | 463 | 454 | 445 | 465 | 462 | 440 | 468 | 490 | 426 | 471 | 469 | 472 | 485 | 505 | 538 | 560 | 598 | 561 | 602 | 654 | 689 | 717 | 769 | 805 | 868 | 953 | 1026 | 1124 | 1210 | 1283 | 1388 | 1521 | 1648 | 1800 | 2024 | 2211 | 2399 | 2621 | 2850 | 3104 | 3343 | 3617 | 3930 | 4216 | 4589 | 4989 | 5299 | 5700 | 6157 | 6518 | 6901 | 7210 | 7715 | 8132 | 8515 | 9003 | 9292 | 9661 | 10079 | 10270 | 10755 | 10969 | 11255 | 11462 | 11652 | 11782 | 11897 | 11928 | 11893 | 11898 | 11856 | 11671 | 11520 | 11313 | 11039 | 10807 | 10568 | 10183 | 9708 | 9252 | 8843 | 8405 | 8037 | 7631 | 7165 | 6772 | 6302 | 5856 | 5429 | 4996 | 4659 | 4302 | 3873 | 3543 | 3290 | 2994 | 2732 | 2517 | 2348 | 2135 | 1941 | 1782 | 1623 | 1528 | 1444 | 1315 | 1177 | 1080 | 1039 | 1027 | 967 | 924 | 899 | 872 | 858 | 842 | 821 | 812 | 796 | 806 | 797 | 795 | 788 | 778 | 763 | 771 | 774 | 764 | 762 | 785 | 802 | 773 | 766 | 787 | 792 | 786 | 761 | 747 | 762 | 760 | 754 | 764 | 752 | 740 | 730 | 716 | 740 | 726 | 727 | 778 | 742 | 701 | 707 | 745 | 736 | 698 | 721 | 710 | 750 | 718 | 721 | 685 | 668 | 698 | 688 | 697 | 674 | 688 | 673 | 710 | 702 | 677 | 699 | 720 | 686 | 661 | 664 | 672 | 670 | 646 | 655 | 681 | 676 | 662 | 661 | 664 | 670 | 677 | 667 | 674 | 653 | 616 | 630 | 645 | 646 | 667 | 671 | 660 | 656 | 673 | 652 | 632 | 665 | 707 | 700 | 671 | 626 | 637 | 634 | 683 | 662 | 670 | 661 | 647 | 641 | 678 | 661 | 666 | 649 | 636 | 649 | 670 | 691 | 680 | 663 | 663 | 669 | 674 | 717 | 753 | 725 | 723 | 736 | 741 | 757 | 767 | 789 | 809 | 803 | 795 | 789 | 793 | 811 | 820 | 850 | 845 | 820 | 788 | 789 | 848 | 844 | 830 | 780 | 810 | 826 | 841 | 835 | 847 | 806 | 796 | 840 | 854 | 826 | 809 | 827 | 810 | 815 | 798 | 798 | 814 | 793 | 822 | 806 | 771 | 746 | 712 | 727 | 714 | 691 | 683 | 695 | 718 | 706 | 669 | 655 | 652 | 667 | 680 | 674 | 653 | 624 | 631 | 623 | 635 | 608 | 593 | 587 | 585 | 575 | 567 | 558 | 548 | 543 | 540 | 540 | 534 | 522 | 518 | 529 | 536 | 507 | 479 | 486 | 483 | 474 | 490 | 500 | 480 | 453 | 439 | 436 | 445 | 449 | 456 | 465 | 441 | 411 | 409 | 440 | 426 | 411 | 416 | 420 | 421 | 413 | 422 | 425 | 408 | 389 | 373 | 369 | 383 | 384 | 368 | 360 | 362 | 355 | 371 | 395 | 380 | 365 | 337 | 316 | 321 | 347 | 345 | 330 | 326 | 332 | 336 | 329 | 324 | 313 | 304 | 308 | 298 | 338 | 301 | 284 | 314 | 306 | 301 | 319 | 322 | 297 | 282 | 284 | 281 | 263 | 265 | 287 | 281 | 260 | 264 | 250 | 255 | 271 | 262 | 240 | 259 | 255 | 253 | 238 | 237 | 219 | 219 | 229 | 226 | 222 | 222 | 233 | 243 | 205 | 206 | 225 | 216 | 220 | 228 | 193 | 199 | 215 | 207 | 194 | 185 | 182 | 193 | 205 | 204 | 180 | 160 | 163 | 167 | 186 | 183 | 173 | 161 | 150 | 165 | 178 | 184 | 156 | 164 | 169 | 163 | 160 | 164 | 167 | 159 | 157 | 161 | 161 | 172 | 167 | 155 | 164 | 161 | 150 | 166 | 148 | 146 | 169 | 153 | 156 | 146 | 161 | 158 | 145 | 130 | 126 | 130 | 122 | 119 | 127 | 123 | 141 | 142 | 142 | 128 | 112 | 123 | 133 | 140 | 147 | 147 | 131 | 123 | 128 | 126 | 127 | 125 | 130 | 144 | 142 | 141 | 135 | 134 | 123 | 130 | 126 | 122 | 115 | 122 | 122 | 128 | 133 | 136 | 136 | 132 | 136 | 139 | 139 | 139 | 130 | 126 | 134 | 137 | 124 | 129 | 146 | 153 | 143 | 138 | 153 | 162 | 154 | 155 | 159 | 166 | 183 | 191 | 191 | 201 | 208 | 209 | 224 | 255 | 273 | 277 | 282 | 283 | 301 | 323 | 327 | 364 | 391 | 390 | 414 | 441 | 466 | 507 | 539 | 571 | 613 | 630 | 632 | 655 | 727 | 746 | 794 | 860 | 900 | 916 | 945 | 1008 | 1086 | 1122 | 1136 | 1170 | 1194 | 1221 | 1296 | 1354 | 1381 | 1376 | 1418 | 1456 | 1563 | 1572 | 1595 | 1616 | 1626 | 1642 | 1666 | 1678 | 1692 | 1711 | 1719 | 1716 | 1730 | 1731 | 1687 | 1686 | 1657 | 1667 | 1621 | 1611 | 1599 | 1532 | 1537 | 1506 | 1509 | 1463 | 1416 | 1313 | 1268 | 1213 | 1175 | 1179 | 1137 | 1080 | 1039 | 982 | 927 | 888 | 880 | 771 | 764 | 665 | 639 | 610 | 569 | 525 | 475 | 471 | 443 | 406 | 367 | 336 | 324 | 299 | 269 | 258 | 244 | 216 | 174 | 174 | 167 | 131 | 136 | 119 | 113 | 90 | 94 | 92 | 82 | 71 | 71 | 73 | 63 | 57 | 53 | 49 | 37 | 38 | 27 | 30 | 39 | 28 | 27 | 26 | 28 | 26 | 28 | 32 | 30 | 30 | 27 | 22 | 20 | 21 | 18 | 21 | 21 | 21 | 23 | 20 | 21 | 21 | 19 | 21 | 24 | 24 | 29 | 26 | 23 | 26 | 28 | 29 | 27 | 25 | 23 | 22 | 24 | 22 | 21 | 15 | 21 | 24 | 26 | 26 | 25 | 26 | 21 | 23 | 25 | 24 | 23 | 20 | 17 | 18 | 20 | 21 | 21 | 14 | 19 | 23 | 21 | 15 | 23 | 23 | 19 | 20 | 19 | 19 | 20 | 15 | 14 | 21 | 26 | 19 | 14 | 20 | 24 | 20 | 16 | 14 | 14 | 18 | 17 | 18 | 20 | 17 | 18 | 21 | 19 | 18 | 17 | 17 | 17 | 14 | 12 | 14 | 15 | 16 | 15 | 16 | 18 | 18 | 18 | 18 | 17 | 17 | 16 | 14 | 17 | 17 | 16 | 17 | 19 | 16 | 18 | 14 | 12 | 15 | 14 | 16 | 18 | 18 | 18 | 20 | 19 | 19 | 19 | 17 | 14 | 16 | 19 | 16 | 15 | 17 | 16 | 15 | 11 | 9 | 14 | 16 | 12 | 10 | 13 | 14 | 13 | 15 | 18 | 15 | 15 | 15 | 14 | 16 | 17 | 18 | 15 | 13 | 11 | 10 | 8 | 13 | 15 | 16 | 10 | 11 | 12 | 12 | 12 | 13 | 11 | 13 | 15 | 15 | 16 | 13 | 16 | 13 | 9 | 6 | 15 | 11 | 11 | 15 | 10 | 10 | 12 | 11 | 11 | 10 | 8 | 12 | 13 | 9 | 11 | 11 | 9 | 7 | 15 | 11 | 10 | 10 | 12 | 8 | 11 | 9 | 11 | 7 | 5 | 5 | 7 | 9 | 8 | 7 | 11 | 10 | 7 | 11 | 8 | 8 | 9 | 7 | 13 | 8 | 6 | 7 | 10 | 7 | 6 | 7 | 7 | 9 | 10 | 7 | 10 | 9 | 8 | 5 | 8 | 9 | 7 | 9 | 8 | 6 | 6 | 8 | 6 | 7 | 7 | 7 | 5 | 7 | 8 | 8 | 8 | 7 | 11 | 12 | 9 | 9 | 11 | 5 | 7 | 7 | 4 | 6 | 6 | 6 | 7 | 6 | 7 | 7 | 8 | 7 | 6 | 7 | 7 | 9 | 9 | 9 | 9 | 7 | 7 | 6 | 6 | 6 | 7 | 7 | 7 | 8 | 7 | 6 | 6 | 7 | 6 | 5 | 7 | 7 | 4 | 7 | 9 | 8 | 7 | 10 | 9 | 7 | 6 | 4 | 5 | 6 | 9 | 8 | 8 | 7 | 6 | 6 | 7 | 7 | 7 | 6 | 7 | 9 | 11 | 8 | 4 | 6 | 5 | 5 | 7 | 9 | 9 | 8 | 8 | 7 | 7 | 6 | 4 | 4 | 7 | 9 | 8 | 6 | 8 | 8 | 7 | 6 | 5 | 5 | 6 | 9 | 10 | 7 | 8 | 10 | 9 | 8 | 6 | 5 | 7 | 12 | 9 | 5 | 7 | 6 | 9 | 6 | 6 | 6 | 6 | 8 | 10 | 11 | 9 | 12 | 4 | 8 | 14 | 10 | 5 | 5 | 12 | 10 | 7 | 4 | 6 | 10 | 8 | 5 | 6 | 13 | 6 | 9 | 6 | 9 | 5 | 8 | 4 | 9 | 6 | 10 | 8 | 10 | 10 | 6 | 10 | 12 | 9 | 7 | 10 | 8 | 14 | 13 | 6 | 14 | 10 | 9 | 10 | 10 | 6 | 6 | 13 | 9 | 13 | 12 | 4 | 14 | 14 | 10 | 11 | 10 | 8 | 13 | 11 | 15 | 7 | 11 | 7 | 11 | 12 | 9 | 8 | 6 | 8 | 11 | 5 | 9 | 12 | 6 | 10 | 7 | 7 | 5 | 6 | 8 | 7 | 7 | 4 | 4 | 6 | 7 | 4 | 4 | 8 | 7 | 8 | 4 | 3 | 4 | 6 | 7 | 6 | 5 | 6 | 4 | 3 | 4 | 2 | 5 | 2 | 3 | 2 | 2 | 2 | 2 | 2 | 1 | 2 | 4 | 7 | 5 | 2 | 3 | 3 | 3 | 5 | 6 | 2 | 1 | 2 | 3 | 5 | 2 | 0 | 1 | 3 | 4 | 1 | 2 | 4 | 5 | 3 | 2 | 1 | 1 | 2 | 1 | 0 | 1 | 1 | 2 | 1 | 2 | 2 | 1 | 2 | 4 | 3 | 3 | 1 | 3 | 3 | 3 | 2 | 2 | 1 | 1 | 3 | 2 | 2 | 3 | 2 | 2 | 1 | 2 | 2 | 1 | 2 | 2 | 1 | 1 | 3 | 2 | 2 | 1 | 1 | 4 | 4 | 1 | 1 | 2 | 1 | 1 | 4 | 2 | 2 | 1 | 1 | 1 | 4 | 2 | 3 | 3 | 2 | 2 | 1 | 1 | 1 | 2 | 1 | 0 | 2 | 0 | 0 | 1 | 3 | 1 | 0 | 0 | 2 | 2 | 1 | 1 | 0 | 1 | 3 | 3 | 1 | 3 | 2 | 1 | 0 | 2 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 0 | 4 | 2 | 4 | 1 | 2 | 2 | 0 | 2 | 3 | 2 | 0 | 3 | 0 | 2 | 1 | 1 | 2 | 1 | 2 | 5 | 3 | 1 | 2 | 2 | 4 | 0 | 0 | 2 | 2 | 2 | 2 | 3 | 1 | 3 | 3 | 1 | 0 | 2 | 2 | 0 | 2 | 0 | 2 | 1 | 1 | 1 | 1 | 0 | 2 | 1 | 1 | 0 | 0 | 1 | 0 | 2 | 1 | 3 | 0 | 0 | 2 | 3 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 2 | 1 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 2 | 2 | 1 | 2 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 2 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 2 | 2 | 0 | 0 | 0 | 2 | 1 | 1 | 1 | 2 | 0 | 0 | 1 | 3 | 1 | 1 | 0 | 0 | 1 | 2 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 2 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 2 | 1 | 2 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 2 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 2 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 2 | 2 | 2 | 0 | 1 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 1 | 0 | 2 | 3 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 7 | 368 | 29 | 29 | 34 | 27 | 28 | 31 | 33 | 36 | 40 | 62 | 84 | 89 | 102 | 106 | 129 | 145 | 167 | 212 | 262 | 302 | 355 | 445 | 494 | 575 | 700 | 800 | 869 | 934 | 1046 | 1181 | 1238 | 1326 | 1387 | 1332 | 1343 | 1353 | 1407 | 1394 | 1399 | 1418 | 1424 | 1439 | 1457 | 1459 | 1444 | 1442 | 1428 | 1421 | 1453 | 1474 | 1473 | 1458 | 1478 | 1553 | 1426 | 1410 | 1424 | 1418 | 1427 | 1425 | 1468 | 1468 | 1434 | 1428 | 1484 | 1434 | 1464 | 1455 | 1420 | 1389 | 1403 | 1451 | 1458 | 1446 | 1514 | 1517 | 1440 | 1455 | 1485 | 1465 | 1491 | 1522 | 1501 | 1491 | 1499 | 1527 | 1496 | 1467 | 1468 | 1465 | 1490 | 1488 | 1501 | 1452 | 1386 | 1441 | 1495 | 1443 | 1409 | 1407 | 1408 | 1426 | 1405 | 1335 | 1343 | 1399 | 1467 | 1444 | 1356 | 1391 | 1409 | 1381 | 1368 | 1357 | 1365 | 1381 | 1387 | 1362 | 1302 | 1253 | 1232 | 1259 | 1284 | 1283 | 1266 | 1224 | 1243 | 1243 | 1230 | 1273 | 1270 | 1255 | 1262 | 1269 | 1277 | 1258 | 1269 | 1290 | 1284 | 1283 | 1284 | 1266 | 1254 | 1262 | 1280 | 1299 | 1272 | 1242 | 1283 | 1323 | 1339 | 1345 | 1349 | 1341 | 1314 | 1360 | 1377 | 1368 | 1409 | 1413 | 1432 | 1421 | 1461 | 1476 | 1548 | 1583 | 1526 | 1583 | 1569 | 1692 | 1693 | 1714 | 1803 | 1772 | 1822 | 1801 | 1868 | 1923 | 1923 | 1956 | 1990 | 2008 | 1936 | 1894 | 1924 | 1939 | 1909 | 1924 | 1901 | 1890 | 1912 | 1864 | 1827 | 1799 | 1779 | 1776 | 1811 | 1786 | 1698 | 1699 | 1709 | 1692 | 1640 | 1615 | 1623 | 1586 | 1559 | 1594 | 1572 | 1527 | 1522 | 1510 | 1507 | 1513 | 1450 | 1343 | 1363 | 1431 | 1322 | 1359 | 1361 | 1321 | 1346 | 1297 | 1359 | 1330 | 1240 | 1263 | 1294 | 1281 | 1264 | 1244 | 1241 | 1247 | 1200 | 1262 | 1215 | 1213 | 1212 | 1158 | 1115 | 1147 | 1178 | 1154 | 1108 | 1164 | 1182 | 1125 | 1124 | 1108 | 1075 | 1085 | 1077 | 1067 | 1075 | 1057 | 1054 | 1040 | 1026 | 1037 | 1002 | 989 | 1008 | 1015 | 990 | 986 | 1029 | 1000 | 932 | 960 | 944 | 965 | 924 | 954 | 933 | 903 | 897 | 898 | 920 | 885 | 886 | 862 | 903 | 905 | 866 | 878 | 876 | 864 | 847 | 843 | 856 | 857 | 843 | 821 | 803 | 788 | 793 | 801 | 815 | 819 | 798 | 776 | 771 | 769 | 766 | 784 | 784 | 767 | 756 | 741 | 744 | 754 | 770 | 745 | 732 | 730 | 730 | 708 | 689 | 712 | 703 | 683 | 702 | 695 | 675 | 703 | 696 | 671 | 657 | 644 | 629 | 648 | 650 | 635 | 617 | 594 | 605 | 606 | 595 | 613 | 616 | 612 | 609 | 589 | 573 | 561 | 550 | 545 | 547 | 547 | 544 | 556 | 558 | 543 | 548 | 551 | 542 | 534 | 517 | 512 | 521 | 533 | 541 | 539 | 530 | 517 | 510 | 520 | 507 | 495 | 502 | 521 | 503 | 468 | 450 | 510 | 510 | 488 | 474 | 508 | 506 | 491 | 471 | 460 | 457 | 463 | 481 | 456 | 453 | 479 | 463 | 455 | 468 | 495 | 471 | 464 | 463 | 453 | 459 | 461 | 449 | 438 | 438 | 450 | 455 | 438 | 433 | 442 | 436 | 448 | 450 | 420 | 421 | 446 | 422 | 409 | 417 | 418 | 428 | 437 | 415 | 419 | 405 | 405 | 405 | 444 | 442 | 437 | 443 | 477 | 425 | 420 | 440 | 453 | 422 | 422 | 438 | 457 | 471 | 467 | 496 | 479 | 519 | 529 | 535 | 573 | 595 | 610 | 622 | 660 | 705 | 769 | 805 | 849 | 926 | 993 | 1094 | 1157 | 1229 | 1408 | 1550 | 1669 | 1817 | 1982 | 2185 | 2409 | 2607 | 2856 | 3148 | 3341 | 3564 | 3902 | 4320 | 4614 | 4896 | 5246 | 5726 | 6049 | 6563 | 6774 | 7250 | 7739 | 7989 | 8575 | 8910 | 9145 | 9492 | 9955 | 10147 | 10439 | 10713 | 10901 | 11187 | 11339 | 11428 | 11448 | 11618 | 11597 | 11521 | 11418 | 11219 | 11086 | 10941 | 10675 | 10354 | 10011 | 9614 | 9226 | 8874 | 8445 | 7950 | 7507 | 7144 | 6740 | 6301 | 5896 | 5406 | 5009 | 4666 | 4284 | 3901 | 3593 | 3301 | 3025 | 2793 | 2483 | 2271 | 2046 | 1880 | 1746 | 1634 | 1519 | 1411 | 1307 | 1237 | 1149 | 1058 | 1035 | 978 | 929 | 878 | 853 | 842 | 834 | 819 | 808 | 803 | 787 | 764 | 749 | 765 | 755 | 732 | 726 | 735 | 754 | 737 | 707 | 707 | 714 | 723 | 725 | 714 | 715 | 716 | 721 | 724 | 719 | 713 | 720 | 713 | 706 | 733 | 737 | 734 | 724 | 706 | 710 | 672 | 664 | 698 | 712 | 719 | 707 | 670 | 690 | 682 | 690 | 690 | 700 | 689 | 681 | 666 | 655 | 664 | 672 | 674 | 654 | 649 | 670 | 668 | 657 | 658 | 662 | 651 | 658 | 657 | 652 | 646 | 638 | 631 | 631 | 650 | 666 | 656 | 641 | 632 | 620 | 619 | 632 | 644 | 620 | 598 | 596 | 626 | 635 | 646 | 660 | 626 | 617 | 651 | 625 | 621 | 634 | 633 | 627 | 671 | 647 | 632 | 657 | 659 | 638 | 642 | 635 | 619 | 636 | 648 | 672 | 677 | 672 | 703 | 667 | 653 | 656 | 674 | 717 | 655 | 728 | 697 | 716 | 706 | 694 | 737 | 732 | 769 | 754 | 740 | 770 | 778 | 755 | 741 | 753 | 799 | 771 | 776 | 801 | 825 | 803 | 809 | 788 | 793 | 790 | 855 | 830 | 812 | 821 | 824 | 821 | 811 | 793 | 809 | 836 | 802 | 798 | 815 | 783 | 780 | 816 | 752 | 761 | 768 | 761 | 748 | 730 | 719 | 726 | 695 | 703 | 707 | 678 | 674 | 651 | 669 | 662 | 631 | 612 | 666 | 632 | 619 | 615 | 632 | 622 | 596 | 560 | 543 | 545 | 551 | 565 | 565 | 547 | 540 | 535 | 517 | 504 | 482 | 477 | 503 | 527 | 480 | 476 | 498 | 488 | 486 | 467 | 456 | 448 | 459 | 473 | 459 | 445 | 445 | 432 | 430 | 438 | 413 | 404 | 406 | 405 | 428 | 412 | 389 | 397 | 399 | 391 | 402 | 411 | 383 | 355 | 352 | 351 | 356 | 363 | 358 | 351 | 354 | 365 | 358 | 346 | 329 | 327 | 327 | 314 | 312 | 313 | 318 | 320 | 322 | 327 | 320 | 302 | 293 | 303 | 303 | 286 | 295 | 301 | 305 | 285 | 257 | 277 | 279 | 268 | 266 | 269 | 274 | 278 | 281 | 282 | 280 | 272 | 260 | 251 | 240 | 238 | 238 | 254 | 249 | 253 | 247 | 250 | 240 | 221 | 228 | 228 | 226 | 227 | 227 | 228 | 216 | 209 | 216 | 213 | 204 | 199 | 202 | 217 | 223 | 214 | 203 | 187 | 179 | 175 | 168 | 180 | 184 | 183 | 179 | 173 | 177 | 171 | 171 | 172 | 184 | 171 | 178 | 156 | 180 | 182 | 166 | 153 | 158 | 162 | 148 | 145 | 157 | 153 | 143 | 145 | 153 | 155 | 152 | 153 | 151 | 142 | 134 | 135 | 146 | 139 | 140 | 149 | 142 | 153 | 143 | 145 | 150 | 151 | 139 | 121 | 123 | 140 | 134 | 124 | 137 | 120 | 133 | 143 | 114 | 101 | 119 | 113 | 136 | 132 | 125 | 131 | 120 | 130 | 129 | 116 | 123 | 125 | 134 | 119 | 120 | 118 | 108 | 119 | 121 | 119 | 119 | 111 | 118 | 119 | 122 | 121 | 121 | 123 | 116 | 112 | 122 | 129 | 129 | 137 | 135 | 124 | 133 | 133 | 140 | 141 | 152 | 143 | 152 | 157 | 162 | 160 | 161 | 159 | 171 | 188 | 189 | 190 | 200 | 214 | 226 | 219 | 221 | 251 | 256 | 269 | 299 | 336 | 329 | 319 | 339 | 371 | 411 | 421 | 418 | 433 | 467 | 517 | 557 | 601 | 647 | 653 | 699 | 778 | 788 | 780 | 821 | 916 | 979 | 974 | 985 | 1014 | 1045 | 1111 | 1168 | 1187 | 1206 | 1262 | 1313 | 1325 | 1347 | 1392 | 1478 | 1489 | 1488 | 1545 | 1586 | 1620 | 1633 | 1647 | 1636 | 1644 | 1668 | 1666 | 1673 | 1694 | 1665 | 1618 | 1611 | 1631 | 1612 | 1552 | 1562 | 1567 | 1512 | 1476 | 1427 | 1382 | 1365 | 1325 | 1248 | 1205 | 1190 | 1134 | 1069 | 1038 | 1019 | 1012 | 960 | 902 | 872 | 777 | 773 | 720 | 681 | 623 | 558 | 515 | 506 | 472 | 446 | 410 | 372 | 359 | 334 | 300 | 277 | 261 | 239 | 203 | 182 | 189 | 169 | 150 | 139 | 126 | 102 | 101 | 94 | 81 | 74 | 70 | 68 | 60 | 53 | 50 | 40 | 38 | 41 | 39 | 38 | 30 | 35 | 28 | 32 | 28 | 22 | 23 | 26 | 28 | 26 | 22 | 21 | 22 | 22 | 25 | 23 | 20 | 17 | 20 | 20 | 18 | 26 | 23 | 17 | 20 | 25 | 26 | 19 | 17 | 16 | 17 | 17 | 13 | 14 | 16 | 19 | 24 | 21 | 14 | 17 | 19 | 22 | 13 | 14 | 23 | 20 | 18 | 16 | 18 | 15 | 12 | 18 | 21 | 17 | 18 | 15 | 17 | 16 | 15 | 24 | 29 | 20 | 23 | 25 | 24 | 23 | 20 | 15 | 17 | 18 | 18 | 19 | 19 | 19 | 14 | 14 | 17 | 15 | 20 | 23 | 11 | 17 | 19 | 16 | 18 | 24 | 21 | 22 | 21 | 18 | 14 | 12 | 16 | 19 | 17 | 17 | 13 | 12 | 14 | 15 | 15 | 19 | 14 | 13 | 16 | 14 | 12 | 11 | 13 | 17 | 17 | 15 | 14 | 14 | 16 | 21 | 20 | 18 | 14 | 10 | 13 | 15 | 16 | 20 | 14 | 12 | 14 | 16 | 15 | 15 | 17 | 17 | 15 | 15 | 14 | 11 | 11 | 12 | 13 | 13 | 15 | 9 | 9 | 11 | 12 | 13 | 11 | 9 | 10 | 10 | 12 | 15 | 14 | 13 | 15 | 14 | 11 | 12 | 10 | 11 | 15 | 10 | 10 | 11 | 11 | 13 | 12 | 11 | 10 | 10 | 11 | 12 | 14 | 14 | 10 | 10 | 10 | 8 | 12 | 15 | 14 | 9 | 9 | 10 | 8 | 9 | 14 | 13 | 10 | 8 | 8 | 9 | 11 | 10 | 11 | 10 | 6 | 9 | 8 | 8 | 7 | 6 | 7 | 9 | 9 | 13 | 12 | 9 | 11 | 13 | 12 | 11 | 11 | 12 | 11 | 10 | 9 | 10 | 11 | 10 | 12 | 6 | 7 | 9 | 7 | 8 | 8 | 6 | 4 | 5 | 8 | 10 | 10 | 10 | 8 | 8 | 7 | 5 | 11 | 10 | 7 | 7 | 9 | 6 | 7 | 5 | 7 | 6 | 5 | 7 | 8 | 9 | 8 | 5 | 9 | 8 | 8 | 11 | 7 | 7 | 8 | 9 | 8 | 8 | 5 | 4 | 7 | 7 | 7 | 5 | 6 | 7 | 5 | 5 | 6 | 6 | 7 | 6 | 7 | 3 | 10 | 11 | 4 | 9 | 4 | 9 | 6 | 5 | 6 | 6 | 6 | 4 | 2 | 4 | 7 | 4 | 6 | 6 | 7 | 4 | 6 | 9 | 4 | 6 | 8 | 6 | 6 | 6 | 6 | 6 | 7 | 9 | 8 | 6 | 7 | 8 | 8 | 4 | 2 | 5 | 5 | 5 | 10 | 8 | 6 | 5 | 5 | 4 | 6 | 7 | 6 | 5 | 6 | 7 | 6 | 5 | 7 | 4 | 6 | 7 | 5 | 9 | 12 | 7 | 8 | 7 | 6 | 7 | 9 | 9 | 7 | 6 | 8 | 9 | 7 | 6 | 8 | 11 | 7 | 7 | 6 | 7 | 9 | 8 | 6 | 8 | 8 | 7 | 6 | 6 | 6 | 9 | 8 | 8 | 12 | 10 | 10 | 10 | 7 | 6 | 7 | 8 | 7 | 5 | 7 | 11 | 11 | 11 | 7 | 8 | 9 | 6 | 7 | 8 | 7 | 7 | 10 | 12 | 10 | 15 | 13 | 7 | 7 | 10 | 10 | 7 | 6 | 6 | 8 | 8 | 7 | 12 | 10 | 9 | 9 | 6 | 6 | 5 | 6 | 10 | 11 | 11 | 10 | 11 | 16 | 14 | 12 | 9 | 7 | 6 | 9 | 7 | 10 | 10 | 7 | 9 | 8 | 8 | 10 | 7 | 8 | 3 | 5 | 9 | 8 | 9 | 7 | 11 | 6 | 5 | 8 | 6 | 8 | 11 | 6 | 5 | 5 | 8 | 10 | 6 | 2 | 6 | 5 | 5 | 5 | 8 | 6 | 2 | 1 | 3 | 6 | 3 | 1 | 5 | 9 | 2 | 4 | 4 | 2 | 5 | 1 | 1 | 5 | 1 | 2 | 2 | 1 | 2 | 0 | 5 | 3 | 4 | 3 | 2 | 3 | 1 | 2 | 3 | 1 | 3 | 0 | 1 | 2 | 1 | 2 | 0 | 2 | 1 | 2 | 3 | 3 | 4 | 3 | 3 | 3 | 2 | 5 | 2 | 0 | 0 | 0 | 1 | 2 | 3 | 2 | 0 | 1 | 1 | 0 | 1 | 2 | 0 | 0 | 1 | 2 | 3 | 1 | 0 | 1 | 2 | 1 | 1 | 3 | 1 | 1 | 2 | 1 | 1 | 1 | 2 | 1 | 0 | 0 | 2 | 5 | 3 | 1 | 2 | 1 | 1 | 1 | 2 | 2 | 1 | 1 | 2 | 2 | 4 | 4 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 3 | 4 | 2 | 1 | 0 | 3 | 3 | 1 | 1 | 0 | 2 | 2 | 0 | 0 | 2 | 4 | 3 | 2 | 1 | 0 | 1 | 0 | 0 | 2 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 2 | 2 | 1 | 1 | 2 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 3 | 3 | 2 | 1 | 0 | 2 | 2 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 2 | 2 | 2 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 2 | 2 | 1 | 1 | 2 | 1 | 2 | 1 | 2 | 1 | 0 | 1 | 1 | 0 | 2 | 1 | 0 | 2 | 2 | 0 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 2 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 2 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 3 | 1 | 1 | 2 | 1 | 1 | 0 | 1 | 0 | 2 | 0 | 1 | 3 | 0 | 0 | 1 | 1 | 1 | 2 | 0 | 0 | 1 | 2 | 1 | 0 | 2 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 0 | 2 | 2 | 1 | 1 | 0 | 0 | 1 | 0 | 2 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 2 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 3 | 2 | 1 | 0 | 1 | 3 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 1 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 2 | 0 | 2 | 1 | 2 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 4 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 3 | 1 | 2 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 8 | 342 | 30 | 40 | 27 | 33 | 50 | 35 | 40 | 44 | 62 | 36 | 67 | 87 | 102 | 118 | 151 | 151 | 183 | 216 | 244 | 284 | 359 | 441 | 497 | 558 | 671 | 742 | 856 | 967 | 1040 | 1101 | 1126 | 1239 | 1307 | 1319 | 1306 | 1339 | 1321 | 1381 | 1390 | 1344 | 1315 | 1359 | 1465 | 1450 | 1386 | 1410 | 1428 | 1466 | 1450 | 1417 | 1419 | 1441 | 1493 | 1473 | 1425 | 1445 | 1418 | 1414 | 1431 | 1465 | 1438 | 1384 | 1396 | 1440 | 1486 | 1425 | 1418 | 1395 | 1404 | 1404 | 1410 | 1434 | 1429 | 1413 | 1470 | 1507 | 1453 | 1416 | 1401 | 1448 | 1475 | 1469 | 1462 | 1449 | 1505 | 1525 | 1489 | 1488 | 1474 | 1436 | 1408 | 1425 | 1439 | 1434 | 1415 | 1417 | 1452 | 1424 | 1403 | 1406 | 1364 | 1367 | 1369 | 1363 | 1342 | 1326 | 1370 | 1403 | 1376 | 1374 | 1357 | 1361 | 1367 | 1342 | 1334 | 1351 | 1369 | 1357 | 1301 | 1247 | 1221 | 1201 | 1241 | 1280 | 1263 | 1271 | 1279 | 1245 | 1227 | 1234 | 1230 | 1211 | 1213 | 1228 | 1242 | 1279 | 1288 | 1251 | 1228 | 1243 | 1270 | 1259 | 1240 | 1233 | 1248 | 1283 | 1291 | 1269 | 1231 | 1229 | 1243 | 1304 | 1345 | 1320 | 1312 | 1333 | 1325 | 1318 | 1342 | 1380 | 1398 | 1392 | 1404 | 1448 | 1486 | 1529 | 1499 | 1555 | 1545 | 1608 | 1552 | 1694 | 1794 | 1780 | 1834 | 1800 | 1841 | 1837 | 1832 | 1874 | 1904 | 1907 | 1895 | 1911 | 1958 | 1934 | 1860 | 1824 | 1838 | 1868 | 1865 | 1821 | 1740 | 1755 | 1803 | 1768 | 1721 | 1696 | 1683 | 1677 | 1660 | 1634 | 1618 | 1607 | 1625 | 1612 | 1561 | 1511 | 1476 | 1470 | 1528 | 1492 | 1430 | 1438 | 1410 | 1421 | 1364 | 1367 | 1383 | 1305 | 1313 | 1361 | 1337 | 1297 | 1283 | 1304 | 1213 | 1202 | 1292 | 1257 | 1217 | 1292 | 1221 | 1148 | 1216 | 1181 | 1156 | 1171 | 1159 | 1145 | 1175 | 1132 | 1120 | 1123 | 1108 | 1098 | 1079 | 1070 | 1069 | 1073 | 1083 | 1092 | 1075 | 1023 | 1032 | 1039 | 1013 | 1024 | 1002 | 994 | 974 | 1011 | 988 | 983 | 975 | 982 | 953 | 951 | 933 | 953 | 952 | 920 | 916 | 908 | 881 | 917 | 879 | 871 | 859 | 867 | 857 | 875 | 843 | 862 | 875 | 874 | 852 | 854 | 849 | 854 | 817 | 809 | 825 | 823 | 797 | 793 | 815 | 816 | 802 | 799 | 796 | 786 | 789 | 786 | 764 | 755 | 760 | 762 | 747 | 736 | 734 | 733 | 727 | 708 | 710 | 706 | 705 | 717 | 704 | 716 | 710 | 708 | 702 | 651 | 671 | 689 | 666 | 639 | 623 | 627 | 616 | 591 | 628 | 632 | 596 | 590 | 603 | 621 | 612 | 605 | 609 | 588 | 572 | 565 | 547 | 534 | 550 | 566 | 566 | 558 | 570 | 576 | 551 | 528 | 521 | 535 | 546 | 546 | 563 | 564 | 535 | 520 | 523 | 519 | 512 | 525 | 521 | 489 | 489 | 511 | 521 | 498 | 476 | 513 | 501 | 481 | 493 | 474 | 448 | 462 | 470 | 455 | 442 | 477 | 495 | 487 | 478 | 458 | 466 | 452 | 450 | 448 | 459 | 459 | 465 | 476 | 454 | 436 | 446 | 475 | 463 | 454 | 467 | 444 | 430 | 443 | 442 | 439 | 437 | 438 | 430 | 402 | 420 | 437 | 437 | 423 | 426 | 437 | 425 | 413 | 428 | 449 | 438 | 432 | 406 | 412 | 434 | 417 | 434 | 440 | 443 | 415 | 437 | 424 | 425 | 461 | 439 | 452 | 465 | 452 | 483 | 518 | 521 | 534 | 525 | 541 | 568 | 626 | 673 | 679 | 680 | 755 | 786 | 829 | 915 | 1006 | 1099 | 1151 | 1252 | 1420 | 1571 | 1698 | 1828 | 1992 | 2229 | 2388 | 2575 | 2832 | 3056 | 3331 | 3664 | 3833 | 4197 | 4475 | 4880 | 5200 | 5555 | 5988 | 6447 | 6671 | 7150 | 7686 | 7957 | 8280 | 8731 | 9091 | 9509 | 9718 | 10038 | 10335 | 10614 | 10817 | 10944 | 11085 | 11318 | 11288 | 11362 | 11406 | 11435 | 11386 | 11062 | 10835 | 10711 | 10407 | 10032 | 9767 | 9472 | 9093 | 8680 | 8204 | 7773 | 7411 | 7008 | 6542 | 6081 | 5724 | 5321 | 4891 | 4491 | 4121 | 3795 | 3538 | 3272 | 2937 | 2732 | 2477 | 2257 | 2024 | 1862 | 1711 | 1535 | 1426 | 1321 | 1247 | 1158 | 1110 | 1035 | 971 | 947 | 913 | 881 | 851 | 801 | 784 | 790 | 791 | 772 | 758 | 748 | 736 | 711 | 723 | 727 | 713 | 696 | 719 | 775 | 776 | 738 | 725 | 733 | 730 | 718 | 712 | 717 | 730 | 730 | 717 | 701 | 715 | 725 | 701 | 689 | 690 | 704 | 704 | 706 | 704 | 703 | 698 | 705 | 698 | 673 | 672 | 680 | 699 | 691 | 702 | 717 | 684 | 686 | 681 | 659 | 659 | 653 | 642 | 660 | 665 | 641 | 631 | 632 | 654 | 659 | 649 | 653 | 644 | 631 | 621 | 612 | 623 | 651 | 643 | 622 | 644 | 645 | 624 | 627 | 632 | 638 | 605 | 598 | 587 | 584 | 597 | 601 | 630 | 639 | 621 | 622 | 658 | 640 | 642 | 614 | 608 | 609 | 590 | 602 | 654 | 612 | 648 | 633 | 624 | 620 | 636 | 621 | 607 | 625 | 626 | 627 | 651 | 657 | 678 | 669 | 652 | 672 | 658 | 671 | 721 | 709 | 707 | 699 | 710 | 716 | 719 | 765 | 751 | 748 | 752 | 731 | 754 | 795 | 785 | 792 | 764 | 760 | 798 | 788 | 826 | 784 | 788 | 767 | 805 | 826 | 799 | 791 | 773 | 789 | 785 | 774 | 801 | 805 | 792 | 774 | 791 | 767 | 761 | 804 | 787 | 777 | 732 | 705 | 710 | 716 | 736 | 688 | 687 | 696 | 679 | 665 | 658 | 658 | 640 | 590 | 611 | 666 | 644 | 630 | 612 | 598 | 598 | 598 | 568 | 556 | 567 | 553 | 538 | 552 | 560 | 555 | 557 | 556 | 542 | 512 | 515 | 533 | 522 | 518 | 549 | 476 | 442 | 478 | 480 | 483 | 473 | 450 | 441 | 448 | 449 | 437 | 432 | 437 | 443 | 444 | 427 | 442 | 442 | 415 | 413 | 373 | 395 | 406 | 375 | 371 | 371 | 373 | 355 | 349 | 355 | 361 | 369 | 370 | 369 | 362 | 355 | 367 | 355 | 355 | 352 | 328 | 343 | 338 | 321 | 315 | 328 | 325 | 325 | 321 | 313 | 320 | 315 | 299 | 303 | 307 | 301 | 293 | 284 | 272 | 279 | 290 | 287 | 275 | 258 | 257 | 268 | 274 | 275 | 275 | 266 | 254 | 247 | 250 | 250 | 234 | 230 | 225 | 240 | 242 | 238 | 250 | 254 | 243 | 234 | 224 | 215 | 217 | 221 | 211 | 197 | 200 | 215 | 213 | 202 | 193 | 190 | 184 | 196 | 190 | 191 | 198 | 181 | 186 | 195 | 192 | 188 | 196 | 192 | 171 | 167 | 171 | 159 | 157 | 155 | 159 | 186 | 176 | 151 | 161 | 150 | 167 | 159 | 153 | 154 | 152 | 155 | 157 | 143 | 140 | 148 | 145 | 143 | 151 | 142 | 143 | 131 | 159 | 163 | 147 | 138 | 125 | 137 | 137 | 126 | 128 | 131 | 132 | 127 | 132 | 137 | 135 | 137 | 124 | 123 | 123 | 105 | 133 | 134 | 150 | 138 | 118 | 113 | 117 | 118 | 127 | 125 | 119 | 116 | 116 | 114 | 107 | 112 | 122 | 112 | 97 | 109 | 95 | 119 | 112 | 110 | 133 | 109 | 123 | 121 | 127 | 124 | 120 | 119 | 124 | 137 | 137 | 125 | 135 | 133 | 141 | 126 | 127 | 131 | 141 | 114 | 139 | 140 | 145 | 132 | 149 | 155 | 157 | 165 | 171 | 177 | 186 | 185 | 187 | 202 | 204 | 218 | 242 | 254 | 252 | 282 | 302 | 320 | 336 | 363 | 361 | 374 | 403 | 439 | 466 | 483 | 505 | 526 | 550 | 581 | 615 | 656 | 704 | 747 | 730 | 769 | 794 | 850 | 910 | 932 | 980 | 1017 | 1046 | 1107 | 1171 | 1188 | 1206 | 1268 | 1330 | 1372 | 1398 | 1461 | 1514 | 1573 | 1535 | 1542 | 1635 | 1663 | 1650 | 1634 | 1622 | 1607 | 1604 | 1636 | 1668 | 1662 | 1669 | 1662 | 1638 | 1634 | 1585 | 1546 | 1549 | 1545 | 1497 | 1455 | 1422 | 1404 | 1354 | 1279 | 1246 | 1222 | 1175 | 1159 | 1127 | 1052 | 1004 | 997 | 958 | 872 | 812 | 786 | 764 | 713 | 652 | 600 | 540 | 499 | 478 | 444 | 405 | 382 | 367 | 345 | 322 | 301 | 267 | 232 | 210 | 203 | 197 | 185 | 169 | 135 | 125 | 119 | 110 | 93 | 77 | 72 | 73 | 74 | 71 | 61 | 50 | 46 | 47 | 43 | 45 | 40 | 31 | 29 | 36 | 39 | 35 | 29 | 29 | 29 | 30 | 26 | 21 | 21 | 23 | 23 | 21 | 21 | 19 | 20 | 24 | 19 | 27 | 24 | 24 | 22 | 22 | 16 | 18 | 18 | 19 | 21 | 18 | 16 | 17 | 18 | 21 | 19 | 18 | 21 | 25 | 17 | 17 | 17 | 23 | 21 | 21 | 26 | 21 | 15 | 11 | 18 | 19 | 20 | 19 | 17 | 19 | 20 | 23 | 22 | 24 | 17 | 18 | 14 | 17 | 17 | 21 | 15 | 14 | 15 | 17 | 20 | 20 | 24 | 23 | 17 | 14 | 17 | 15 | 18 | 19 | 13 | 23 | 21 | 23 | 22 | 19 | 15 | 10 | 15 | 18 | 20 | 19 | 17 | 16 | 18 | 21 | 18 | 14 | 17 | 21 | 15 | 15 | 13 | 18 | 16 | 13 | 16 | 21 | 16 | 13 | 16 | 15 | 13 | 11 | 13 | 15 | 13 | 14 | 13 | 16 | 22 | 19 | 16 | 18 | 15 | 16 | 13 | 14 | 16 | 17 | 12 | 10 | 14 | 17 | 14 | 11 | 12 | 10 | 9 | 15 | 23 | 17 | 15 | 15 | 10 | 12 | 14 | 13 | 13 | 13 | 9 | 10 | 9 | 12 | 15 | 13 | 10 | 12 | 10 | 9 | 9 | 10 | 10 | 6 | 11 | 12 | 10 | 11 | 13 | 15 | 14 | 12 | 10 | 11 | 11 | 11 | 13 | 15 | 13 | 10 | 7 | 11 | 13 | 10 | 12 | 14 | 11 | 7 | 3 | 4 | 8 | 9 | 8 | 9 | 7 | 7 | 8 | 9 | 9 | 7 | 6 | 9 | 8 | 5 | 8 | 12 | 13 | 11 | 11 | 9 | 7 | 5 | 6 | 6 | 6 | 9 | 9 | 8 | 10 | 11 | 8 | 8 | 4 | 8 | 14 | 10 | 6 | 7 | 9 | 10 | 9 | 7 | 8 | 9 | 8 | 7 | 8 | 7 | 8 | 7 | 7 | 9 | 6 | 5 | 8 | 5 | 10 | 10 | 5 | 7 | 9 | 7 | 7 | 7 | 7 | 8 | 7 | 7 | 7 | 4 | 10 | 6 | 6 | 6 | 7 | 7 | 9 | 8 | 8 | 8 | 7 | 6 | 5 | 5 | 8 | 4 | 6 | 12 | 7 | 2 | 7 | 4 | 7 | 9 | 9 | 5 | 3 | 5 | 5 | 6 | 7 | 6 | 4 | 5 | 5 | 7 | 4 | 6 | 3 | 5 | 13 | 6 | 6 | 5 | 5 | 4 | 3 | 8 | 7 | 4 | 4 | 6 | 8 | 7 | 8 | 6 | 10 | 6 | 5 | 4 | 4 | 8 | 5 | 4 | 3 | 6 | 6 | 5 | 6 | 8 | 6 | 3 | 4 | 5 | 5 | 5 | 6 | 5 | 5 | 1 | 4 | 8 | 8 | 4 | 5 | 5 | 5 | 5 | 6 | 8 | 6 | 5 | 7 | 7 | 6 | 6 | 4 | 4 | 7 | 6 | 6 | 7 | 7 | 6 | 6 | 6 | 7 | 10 | 9 | 5 | 3 | 5 | 6 | 5 | 7 | 8 | 10 | 10 | 5 | 6 | 5 | 8 | 9 | 7 | 7 | 8 | 5 | 4 | 5 | 8 | 8 | 10 | 10 | 7 | 11 | 13 | 10 | 8 | 7 | 10 | 10 | 10 | 10 | 9 | 10 | 12 | 10 | 8 | 7 | 11 | 13 | 8 | 6 | 8 | 9 | 8 | 9 | 9 | 8 | 9 | 7 | 6 | 10 | 11 | 9 | 9 | 9 | 10 | 10 | 7 | 8 | 10 | 10 | 15 | 12 | 7 | 7 | 7 | 8 | 9 | 10 | 8 | 7 | 8 | 6 | 7 | 5 | 5 | 5 | 5 | 8 | 7 | 7 | 6 | 5 | 6 | 6 | 7 | 8 | 8 | 8 | 10 | 8 | 7 | 3 | 4 | 6 | 7 | 7 | 2 | 3 | 3 | 5 | 2 | 2 | 2 | 4 | 8 | 8 | 5 | 2 | 1 | 3 | 5 | 3 | 5 | 4 | 3 | 3 | 2 | 4 | 1 | 4 | 1 | 1 | 3 | 2 | 1 | 4 | 2 | 3 | 0 | 2 | 3 | 2 | 3 | 5 | 1 | 0 | 1 | 0 | 1 | 2 | 2 | 3 | 2 | 1 | 2 | 4 | 0 | 0 | 2 | 1 | 3 | 3 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 2 | 1 | 1 | 0 | 2 | 0 | 2 | 1 | 0 | 2 | 0 | 2 | 4 | 4 | 0 | 4 | 7 | 3 | 3 | 5 | 2 | 1 | 2 | 2 | 1 | 3 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 3 | 1 | 2 | 1 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 2 | 3 | 1 | 0 | 0 | 2 | 3 | 2 | 2 | 1 | 2 | 2 | 1 | 2 | 1 | 1 | 1 | 2 | 2 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 2 | 1 | 0 | 0 | 1 | 2 | 3 | 4 | 4 | 2 | 2 | 0 | 0 | 0 | 3 | 4 | 2 | 2 | 1 | 2 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 2 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 2 | 2 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 2 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 2 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 2 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 2 | 1 | 0 | 0 | 1 | 1 | 1 | 2 | 1 | 2 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 2 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 2 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 2 | 1 | 1 | 1 | 0 | 1 | 3 | 1 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 9 | 323 | 25 | 31 | 35 | 26 | 26 | 47 | 35 | 46 | 48 | 49 | 69 | 76 | 88 | 117 | 144 | 148 | 166 | 208 | 257 | 286 | 335 | 454 | 485 | 591 | 661 | 748 | 840 | 953 | 1056 | 1114 | 1164 | 1278 | 1310 | 1331 | 1321 | 1394 | 1414 | 1373 | 1355 | 1338 | 1346 | 1392 | 1421 | 1488 | 1405 | 1396 | 1379 | 1382 | 1389 | 1432 | 1455 | 1378 | 1395 | 1492 | 1371 | 1465 | 1439 | 1418 | 1411 | 1394 | 1407 | 1381 | 1379 | 1422 | 1457 | 1395 | 1360 | 1393 | 1452 | 1444 | 1422 | 1407 | 1405 | 1428 | 1468 | 1504 | 1410 | 1391 | 1389 | 1426 | 1464 | 1459 | 1444 | 1470 | 1490 | 1490 | 1473 | 1411 | 1433 | 1446 | 1424 | 1406 | 1405 | 1411 | 1421 | 1434 | 1451 | 1422 | 1383 | 1380 | 1386 | 1337 | 1336 | 1373 | 1365 | 1340 | 1382 | 1421 | 1356 | 1356 | 1366 | 1336 | 1352 | 1374 | 1365 | 1371 | 1352 | 1345 | 1317 | 1239 | 1203 | 1228 | 1260 | 1256 | 1235 | 1262 | 1263 | 1241 | 1251 | 1285 | 1271 | 1224 | 1221 | 1236 | 1239 | 1229 | 1243 | 1252 | 1231 | 1221 | 1239 | 1283 | 1295 | 1267 | 1263 | 1267 | 1251 | 1243 | 1274 | 1301 | 1313 | 1328 | 1334 | 1322 | 1309 | 1297 | 1338 | 1379 | 1380 | 1388 | 1372 | 1379 | 1402 | 1423 | 1470 | 1460 | 1480 | 1562 | 1565 | 1593 | 1677 | 1765 | 1755 | 1774 | 1824 | 1844 | 1823 | 1860 | 1870 | 1903 | 1927 | 1957 | 1940 | 1904 | 1905 | 1891 | 1874 | 1868 | 1845 | 1833 | 1848 | 1850 | 1821 | 1820 | 1783 | 1707 | 1707 | 1706 | 1693 | 1682 | 1631 | 1633 | 1632 | 1596 | 1512 | 1493 | 1523 | 1546 | 1513 | 1467 | 1470 | 1459 | 1436 | 1461 | 1431 | 1381 | 1328 | 1340 | 1327 | 1283 | 1296 | 1366 | 1277 | 1285 | 1340 | 1236 | 1321 | 1257 | 1295 | 1256 | 1213 | 1221 | 1175 | 1134 | 1144 | 1160 | 1214 | 1160 | 1160 | 1157 | 1123 | 1121 | 1100 | 1099 | 1132 | 1158 | 1154 | 1105 | 1063 | 1047 | 1018 | 1014 | 1012 | 985 | 998 | 990 | 989 | 1022 | 1020 | 1003 | 1039 | 991 | 938 | 928 | 992 | 964 | 942 | 952 | 949 | 945 | 942 | 964 | 910 | 909 | 898 | 920 | 902 | 890 | 880 | 870 | 857 | 869 | 851 | 829 | 838 | 858 | 857 | 869 | 868 | 847 | 808 | 840 | 844 | 802 | 798 | 801 | 792 | 807 | 809 | 793 | 796 | 807 | 795 | 757 | 738 | 750 | 769 | 766 | 742 | 742 | 761 | 734 | 715 | 717 | 741 | 732 | 719 | 723 | 721 | 713 | 667 | 651 | 677 | 662 | 655 | 668 | 680 | 666 | 626 | 611 | 624 | 617 | 615 | 598 | 589 | 609 | 615 | 595 | 584 | 580 | 581 | 594 | 598 | 583 | 565 | 560 | 559 | 564 | 577 | 577 | 556 | 536 | 527 | 513 | 516 | 535 | 534 | 521 | 517 | 516 | 505 | 499 | 499 | 504 | 508 | 489 | 502 | 509 | 486 | 484 | 500 | 488 | 473 | 485 | 520 | 508 | 482 | 475 | 497 | 507 | 490 | 471 | 479 | 454 | 431 | 462 | 479 | 480 | 465 | 460 | 477 | 463 | 447 | 445 | 434 | 429 | 434 | 449 | 461 | 455 | 455 | 457 | 441 | 434 | 447 | 446 | 430 | 414 | 412 | 423 | 438 | 434 | 428 | 429 | 428 | 438 | 442 | 438 | 412 | 407 | 417 | 412 | 440 | 449 | 415 | 411 | 471 | 453 | 421 | 434 | 440 | 442 | 417 | 428 | 436 | 456 | 446 | 444 | 451 | 484 | 494 | 496 | 533 | 519 | 520 | 537 | 562 | 599 | 676 | 688 | 672 | 737 | 777 | 822 | 893 | 947 | 1042 | 1160 | 1219 | 1341 | 1530 | 1650 | 1824 | 2004 | 2161 | 2313 | 2604 | 2851 | 3041 | 3333 | 3611 | 3783 | 4187 | 4464 | 4902 | 5290 | 5591 | 5915 | 6435 | 6677 | 7210 | 7530 | 7943 | 8175 | 8690 | 8993 | 9429 | 9746 | 10092 | 10400 | 10684 | 10792 | 10946 | 10992 | 11305 | 11284 | 11259 | 11407 | 11286 | 11250 | 10987 | 10758 | 10647 | 10436 | 10165 | 9730 | 9243 | 8895 | 8687 | 8320 | 7791 | 7302 | 6850 | 6532 | 6156 | 5658 | 5178 | 4719 | 4365 | 4098 | 3788 | 3530 | 3233 | 2987 | 2724 | 2398 | 2147 | 2037 | 1889 | 1725 | 1553 | 1422 | 1327 | 1248 | 1153 | 1043 | 1007 | 942 | 920 | 916 | 859 | 800 | 783 | 773 | 765 | 784 | 804 | 778 | 731 | 728 | 720 | 731 | 755 | 734 | 709 | 719 | 718 | 702 | 709 | 722 | 741 | 755 | 744 | 711 | 689 | 704 | 725 | 719 | 703 | 696 | 685 | 677 | 691 | 683 | 680 | 711 | 710 | 709 | 685 | 671 | 694 | 686 | 671 | 676 | 697 | 714 | 719 | 687 | 688 | 711 | 681 | 660 | 662 | 644 | 629 | 653 | 676 | 645 | 645 | 658 | 648 | 642 | 634 | 633 | 621 | 606 | 604 | 640 | 668 | 637 | 601 | 600 | 626 | 655 | 657 | 628 | 603 | 607 | 612 | 610 | 600 | 585 | 602 | 614 | 617 | 603 | 617 | 620 | 603 | 616 | 635 | 584 | 590 | 599 | 620 | 597 | 614 | 646 | 631 | 609 | 616 | 641 | 626 | 636 | 640 | 633 | 643 | 648 | 646 | 631 | 607 | 629 | 645 | 676 | 627 | 649 | 665 | 663 | 695 | 721 | 702 | 695 | 672 | 689 | 729 | 726 | 722 | 743 | 750 | 746 | 744 | 752 | 796 | 719 | 784 | 782 | 841 | 874 | 819 | 798 | 799 | 801 | 804 | 808 | 811 | 785 | 783 | 801 | 803 | 794 | 786 | 798 | 825 | 761 | 759 | 759 | 718 | 811 | 755 | 728 | 702 | 695 | 707 | 708 | 688 | 687 | 683 | 691 | 696 | 666 | 654 | 661 | 645 | 594 | 618 | 625 | 619 | 627 | 636 | 614 | 590 | 609 | 563 | 539 | 568 | 573 | 545 | 518 | 523 | 536 | 520 | 511 | 516 | 510 | 505 | 515 | 526 | 522 | 487 | 475 | 461 | 458 | 471 | 452 | 461 | 447 | 432 | 441 | 447 | 441 | 428 | 424 | 441 | 429 | 412 | 411 | 399 | 386 | 376 | 411 | 399 | 400 | 390 | 361 | 354 | 366 | 377 | 365 | 365 | 382 | 366 | 338 | 327 | 325 | 326 | 331 | 345 | 343 | 342 | 325 | 327 | 354 | 356 | 346 | 311 | 291 | 293 | 303 | 300 | 296 | 302 | 300 | 288 | 293 | 310 | 296 | 291 | 296 | 282 | 274 | 272 | 266 | 282 | 289 | 287 | 278 | 265 | 258 | 261 | 267 | 254 | 242 | 233 | 240 | 249 | 251 | 258 | 254 | 250 | 232 | 213 | 216 | 229 | 230 | 222 | 216 | 212 | 216 | 218 | 214 | 220 | 222 | 205 | 204 | 200 | 208 | 210 | 201 | 190 | 184 | 190 | 188 | 180 | 188 | 187 | 174 | 174 | 175 | 163 | 156 | 159 | 162 | 164 | 167 | 168 | 164 | 180 | 168 | 171 | 151 | 156 | 160 | 158 | 159 | 163 | 153 | 139 | 146 | 157 | 144 | 142 | 130 | 130 | 142 | 138 | 135 | 139 | 158 | 150 | 138 | 120 | 129 | 132 | 139 | 143 | 130 | 125 | 130 | 124 | 116 | 118 | 129 | 128 | 138 | 119 | 132 | 132 | 126 | 118 | 112 | 111 | 125 | 128 | 127 | 117 | 113 | 116 | 129 | 131 | 105 | 113 | 121 | 115 | 103 | 118 | 126 | 137 | 114 | 110 | 106 | 119 | 129 | 121 | 118 | 122 | 126 | 127 | 131 | 130 | 122 | 124 | 129 | 129 | 134 | 123 | 133 | 138 | 129 | 150 | 137 | 150 | 146 | 153 | 161 | 156 | 162 | 175 | 184 | 188 | 170 | 179 | 207 | 214 | 209 | 223 | 236 | 265 | 275 | 268 | 318 | 325 | 338 | 373 | 397 | 406 | 425 | 437 | 450 | 480 | 506 | 534 | 582 | 605 | 638 | 637 | 716 | 751 | 786 | 827 | 865 | 907 | 950 | 959 | 991 | 1053 | 1100 | 1124 | 1162 | 1203 | 1233 | 1303 | 1369 | 1398 | 1419 | 1450 | 1520 | 1475 | 1529 | 1561 | 1566 | 1565 | 1568 | 1577 | 1582 | 1612 | 1611 | 1600 | 1610 | 1621 | 1619 | 1589 | 1562 | 1549 | 1516 | 1530 | 1541 | 1452 | 1414 | 1376 | 1341 | 1290 | 1249 | 1227 | 1207 | 1163 | 1119 | 1075 | 1037 | 993 | 946 | 893 | 841 | 784 | 717 | 716 | 706 | 631 | 582 | 552 | 518 | 493 | 461 | 436 | 407 | 378 | 352 | 324 | 300 | 269 | 243 | 219 | 201 | 182 | 161 | 146 | 143 | 134 | 128 | 117 | 100 | 99 | 97 | 90 | 79 | 68 | 59 | 51 | 47 | 43 | 41 | 39 | 35 | 34 | 34 | 38 | 29 | 28 | 24 | 20 | 23 | 24 | 23 | 22 | 21 | 21 | 23 | 24 | 20 | 18 | 21 | 24 | 20 | 15 | 22 | 24 | 27 | 21 | 16 | 20 | 18 | 19 | 18 | 20 | 20 | 21 | 19 | 15 | 17 | 25 | 25 | 22 | 19 | 16 | 15 | 24 | 21 | 15 | 18 | 15 | 20 | 20 | 17 | 19 | 20 | 17 | 15 | 14 | 15 | 18 | 18 | 24 | 14 | 19 | 21 | 15 | 11 | 18 | 15 | 17 | 18 | 16 | 16 | 12 | 12 | 12 | 17 | 21 | 17 | 17 | 15 | 16 | 23 | 15 | 14 | 11 | 17 | 15 | 13 | 11 | 20 | 20 | 16 | 19 | 17 | 15 | 16 | 16 | 16 | 15 | 16 | 18 | 12 | 15 | 16 | 21 | 17 | 11 | 13 | 17 | 21 | 17 | 14 | 14 | 16 | 17 | 13 | 12 | 20 | 19 | 17 | 21 | 14 | 16 | 14 | 15 | 8 | 11 | 17 | 15 | 10 | 12 | 17 | 14 | 11 | 10 | 11 | 13 | 15 | 16 | 14 | 11 | 13 | 17 | 15 | 13 | 16 | 18 | 15 | 16 | 12 | 12 | 11 | 10 | 10 | 11 | 9 | 9 | 12 | 12 | 13 | 11 | 8 | 8 | 13 | 15 | 10 | 11 | 9 | 8 | 10 | 9 | 9 | 11 | 11 | 12 | 14 | 11 | 9 | 12 | 14 | 10 | 9 | 10 | 12 | 9 | 7 | 10 | 10 | 8 | 10 | 11 | 8 | 6 | 7 | 10 | 11 | 7 | 8 | 13 | 11 | 10 | 8 | 9 | 12 | 9 | 10 | 12 | 9 | 8 | 6 | 5 | 8 | 10 | 13 | 14 | 10 | 5 | 6 | 8 | 8 | 7 | 9 | 9 | 14 | 11 | 9 | 8 | 8 | 8 | 8 | 9 | 8 | 6 | 5 | 6 | 7 | 8 | 8 | 7 | 4 | 9 | 11 | 11 | 5 | 6 | 6 | 5 | 8 | 6 | 6 | 6 | 5 | 6 | 7 | 8 | 8 | 6 | 8 | 7 | 2 | 8 | 8 | 4 | 8 | 7 | 7 | 5 | 5 | 8 | 7 | 5 | 4 | 6 | 8 | 6 | 6 | 5 | 8 | 5 | 9 | 6 | 7 | 8 | 5 | 10 | 5 | 6 | 2 | 3 | 3 | 5 | 6 | 6 | 6 | 5 | 7 | 7 | 6 | 7 | 5 | 10 | 6 | 6 | 4 | 5 | 3 | 5 | 5 | 5 | 5 | 3 | 4 | 7 | 7 | 6 | 4 | 3 | 5 | 5 | 4 | 4 | 5 | 6 | 6 | 4 | 7 | 7 | 5 | 2 | 7 | 6 | 5 | 4 | 5 | 7 | 5 | 7 | 5 | 4 | 6 | 4 | 4 | 5 | 5 | 8 | 6 | 8 | 8 | 5 | 6 | 8 | 6 | 4 | 4 | 6 | 4 | 6 | 9 | 6 | 6 | 8 | 4 | 6 | 5 | 7 | 6 | 10 | 8 | 9 | 10 | 7 | 9 | 4 | 5 | 6 | 8 | 11 | 7 | 2 | 6 | 8 | 7 | 7 | 7 | 5 | 7 | 8 | 10 | 7 | 8 | 10 | 7 | 6 | 7 | 3 | 8 | 9 | 7 | 8 | 8 | 7 | 7 | 10 | 10 | 8 | 9 | 9 | 9 | 10 | 9 | 8 | 7 | 5 | 4 | 10 | 11 | 6 | 10 | 10 | 5 | 6 | 9 | 8 | 6 | 4 | 8 | 9 | 10 | 12 | 11 | 9 | 10 | 12 | 15 | 12 | 5 | 5 | 5 | 5 | 7 | 8 | 5 | 4 | 7 | 8 | 9 | 9 | 8 | 8 | 7 | 5 | 7 | 8 | 8 | 8 | 7 | 5 | 8 | 7 | 6 | 6 | 5 | 8 | 6 | 4 | 6 | 5 | 4 | 5 | 3 | 3 | 3 | 5 | 5 | 5 | 4 | 5 | 4 | 2 | 2 | 4 | 3 | 4 | 4 | 4 | 2 | 2 | 3 | 5 | 3 | 1 | 5 | 3 | 2 | 1 | 3 | 3 | 1 | 0 | 2 | 1 | 3 | 4 | 4 | 3 | 3 | 1 | 3 | 4 | 1 | 1 | 4 | 2 | 3 | 2 | 0 | 3 | 4 | 3 | 2 | 2 | 1 | 4 | 1 | 0 | 0 | 2 | 2 | 0 | 4 | 2 | 1 | 1 | 1 | 0 | 2 | 3 | 1 | 2 | 1 | 1 | 1 | 2 | 2 | 1 | 3 | 2 | 3 | 0 | 3 | 0 | 2 | 0 | 1 | 3 | 1 | 1 | 4 | 0 | 0 | 5 | 0 | 1 | 1 | 2 | 1 | 2 | 3 | 3 | 0 | 2 | 3 | 3 | 0 | 1 | 2 | 2 | 2 | 0 | 1 | 2 | 1 | 3 | 3 | 0 | 1 | 1 | 2 | 0 | 2 | 2 | 1 | 2 | 2 | 0 | 0 | 2 | 2 | 0 | 0 | 1 | 2 | 2 | 0 | 2 | 3 | 1 | 3 | 1 | 1 | 2 | 3 | 2 | 1 | 2 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 2 | 2 | 1 | 2 | 3 | 1 | 0 | 0 | 1 | 2 | 2 | 1 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 0 | 1 | 2 | 4 | 3 | 1 | 1 | 1 | 3 | 2 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 2 | 3 | 2 | 2 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 2 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 2 | 3 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 2 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 2 | 1 | 0 | 2 | 0 | 1 | 0 | 0 | 2 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 2 | 1 | 1 | 0 | 1 | 0 | 0 | 2 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 2 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 2 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 1 | 1 | 1 | 2 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 2 | 1 | 0 | 0 | 1 | 0 | 2 | 2 | 1 | 2 | 1 | 1 | 3 | 5 | 5 | 4 | 5 | 3 | 2 | 1 | 4 | 6 | 5 | 4 | 2 | 3 | 2 | 4 | 3 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 3 | 3 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 3 | 0 | 1 | 1 | 0 | 3 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 2 | 2 | 2 | 2 | 1 | 1 | 0 | 1 | 1 | 1 | 2 | 2 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 2 | 2 | 3 | 1 | 0 | 0 | 1 | 1 | 1 | 2 | 2 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 2 | 0 | 1 | 0 | 0 | 1 | 1 | 2 | 1 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 2 | 2 | 2 | 0 | 1 | 4 | 2 | 0 | 2 | 0 | 1 | 2 | 1 | 3 | 0 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 3 | 2 | 1 | 2 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 2 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 2 | 2 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 2 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 2 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0