## Results
Hidex Automatic Gamma Counter
Counting 18F
Run
Template | Counting 18F
Run status | Finished
Run started | 2021-09-01 19:12
Run type | Count only, single rack
Software version | 1.8.0.0
AMG serial number | 2210227
Firmware versions | 0: 50, 1: 50, F: 100
Counting time | 60
Result normalization | 2021-09-01 11:48
Comment
Counting windows
Window start | Window end | Additional windows | Isotope | Counting efficiency | Half-life
400 | 650 |  | ¹⁸F | 1 | 1.82951 h
Racks
Rack | Time | Barcode | Slot count | Latest QC check | QC status
1 | 2021-09-01 19:12 | 006 | 10 | 2021-09-01 10:39 | Calibration suggested
Results
Rack | Vial | Time | Counted time (s) | Dead time factor | ¹⁸F (counts) | ¹⁸F (CPM) | Normalized ¹⁸F (CPM) | Normalized ¹⁸F (Bq) | Tags
1 | 1 | 2021-09-01 19:13 | 60 | 1.000 | 54 | 54 | 897 | 15
1 | 2 | 2021-09-01 19:14 | 60 | 1.311 | 1187056 | 1556230 | 26081872 | 434698
1 | 3 | 2021-09-01 19:16 | 60 | 1.310 | 1178156 | 1543011 | 26103589 | 435060
1 | 4 | 2021-09-01 19:17 | 60 | 1.303 | 1160725 | 1512523 | 25828541 | 430476
1 | 5 | 2021-09-01 19:19 | 60 | 1.303 | 1160259 | 1511574 | 26055063 | 434251
1 | 6 | 2021-09-01 19:20 | 60 | 1.105 | 484369 | 535324 | 9314764 | 155246
1 | 7 | 2021-09-01 19:22 | 60 | 1.102 | 466913 | 514633 | 9039556 | 150659
1 | 8 | 2021-09-01 19:23 | 60 | 1.100 | 458851 | 504801 | 8950826 | 149180
1 | 9 | 2021-09-01 19:25 | 60 | 1.099 | 457066 | 502481 | 8993953 | 149899
1 | 10 | 2021-09-01 19:26 | 60 | 1.000 | 53 | 53 | 958 | 16

## Spectra
Hidex Automatic Gamma Counter
Counting 18F
Run
Template | Counting 18F
Run status | Finished
Run started | 2021-09-01 19:12
Run type | Count only, single rack
Software version | 1.8.0.0
AMG serial number | 2210227
Firmware versions | 0: 50, 1: 50, F: 100
Counting time | 60
Result normalization | 2021-09-01 11:48
Comment
Spectra
Channel
Rack | Vial | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66 | 67 | 68 | 69 | 70 | 71 | 72 | 73 | 74 | 75 | 76 | 77 | 78 | 79 | 80 | 81 | 82 | 83 | 84 | 85 | 86 | 87 | 88 | 89 | 90 | 91 | 92 | 93 | 94 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 117 | 118 | 119 | 120 | 121 | 122 | 123 | 124 | 125 | 126 | 127 | 128 | 129 | 130 | 131 | 132 | 133 | 134 | 135 | 136 | 137 | 138 | 139 | 140 | 141 | 142 | 143 | 144 | 145 | 146 | 147 | 148 | 149 | 150 | 151 | 152 | 153 | 154 | 155 | 156 | 157 | 158 | 159 | 160 | 161 | 162 | 163 | 164 | 165 | 166 | 167 | 168 | 169 | 170 | 171 | 172 | 173 | 174 | 175 | 176 | 177 | 178 | 179 | 180 | 181 | 182 | 183 | 184 | 185 | 186 | 187 | 188 | 189 | 190 | 191 | 192 | 193 | 194 | 195 | 196 | 197 | 198 | 199 | 200 | 201 | 202 | 203 | 204 | 205 | 206 | 207 | 208 | 209 | 210 | 211 | 212 | 213 | 214 | 215 | 216 | 217 | 218 | 219 | 220 | 221 | 222 | 223 | 224 | 225 | 226 | 227 | 228 | 229 | 230 | 231 | 232 | 233 | 234 | 235 | 236 | 237 | 238 | 239 | 240 | 241 | 242 | 243 | 244 | 245 | 246 | 247 | 248 | 249 | 250 | 251 | 252 | 253 | 254 | 255 | 256 | 257 | 258 | 259 | 260 | 261 | 262 | 263 | 264 | 265 | 266 | 267 | 268 | 269 | 270 | 271 | 272 | 273 | 274 | 275 | 276 | 277 | 278 | 279 | 280 | 281 | 282 | 283 | 284 | 285 | 286 | 287 | 288 | 289 | 290 | 291 | 292 | 293 | 294 | 295 | 296 | 297 | 298 | 299 | 300 | 301 | 302 | 303 | 304 | 305 | 306 | 307 | 308 | 309 | 310 | 311 | 312 | 313 | 314 | 315 | 316 | 317 | 318 | 319 | 320 | 321 | 322 | 323 | 324 | 325 | 326 | 327 | 328 | 329 | 330 | 331 | 332 | 333 | 334 | 335 | 336 | 337 | 338 | 339 | 340 | 341 | 342 | 343 | 344 | 345 | 346 | 347 | 348 | 349 | 350 | 351 | 352 | 353 | 354 | 355 | 356 | 357 | 358 | 359 | 360 | 361 | 362 | 363 | 364 | 365 | 366 | 367 | 368 | 369 | 370 | 371 | 372 | 373 | 374 | 375 | 376 | 377 | 378 | 379 | 380 | 381 | 382 | 383 | 384 | 385 | 386 | 387 | 388 | 389 | 390 | 391 | 392 | 393 | 394 | 395 | 396 | 397 | 398 | 399 | 400 | 401 | 402 | 403 | 404 | 405 | 406 | 407 | 408 | 409 | 410 | 411 | 412 | 413 | 414 | 415 | 416 | 417 | 418 | 419 | 420 | 421 | 422 | 423 | 424 | 425 | 426 | 427 | 428 | 429 | 430 | 431 | 432 | 433 | 434 | 435 | 436 | 437 | 438 | 439 | 440 | 441 | 442 | 443 | 444 | 445 | 446 | 447 | 448 | 449 | 450 | 451 | 452 | 453 | 454 | 455 | 456 | 457 | 458 | 459 | 460 | 461 | 462 | 463 | 464 | 465 | 466 | 467 | 468 | 469 | 470 | 471 | 472 | 473 | 474 | 475 | 476 | 477 | 478 | 479 | 480 | 481 | 482 | 483 | 484 | 485 | 486 | 487 | 488 | 489 | 490 | 491 | 492 | 493 | 494 | 495 | 496 | 497 | 498 | 499 | 500 | 501 | 502 | 503 | 504 | 505 | 506 | 507 | 508 | 509 | 510 | 511 | 512 | 513 | 514 | 515 | 516 | 517 | 518 | 519 | 520 | 521 | 522 | 523 | 524 | 525 | 526 | 527 | 528 | 529 | 530 | 531 | 532 | 533 | 534 | 535 | 536 | 537 | 538 | 539 | 540 | 541 | 542 | 543 | 544 | 545 | 546 | 547 | 548 | 549 | 550 | 551 | 552 | 553 | 554 | 555 | 556 | 557 | 558 | 559 | 560 | 561 | 562 | 563 | 564 | 565 | 566 | 567 | 568 | 569 | 570 | 571 | 572 | 573 | 574 | 575 | 576 | 577 | 578 | 579 | 580 | 581 | 582 | 583 | 584 | 585 | 586 | 587 | 588 | 589 | 590 | 591 | 592 | 593 | 594 | 595 | 596 | 597 | 598 | 599 | 600 | 601 | 602 | 603 | 604 | 605 | 606 | 607 | 608 | 609 | 610 | 611 | 612 | 613 | 614 | 615 | 616 | 617 | 618 | 619 | 620 | 621 | 622 | 623 | 624 | 625 | 626 | 627 | 628 | 629 | 630 | 631 | 632 | 633 | 634 | 635 | 636 | 637 | 638 | 639 | 640 | 641 | 642 | 643 | 644 | 645 | 646 | 647 | 648 | 649 | 650 | 651 | 652 | 653 | 654 | 655 | 656 | 657 | 658 | 659 | 660 | 661 | 662 | 663 | 664 | 665 | 666 | 667 | 668 | 669 | 670 | 671 | 672 | 673 | 674 | 675 | 676 | 677 | 678 | 679 | 680 | 681 | 682 | 683 | 684 | 685 | 686 | 687 | 688 | 689 | 690 | 691 | 692 | 693 | 694 | 695 | 696 | 697 | 698 | 699 | 700 | 701 | 702 | 703 | 704 | 705 | 706 | 707 | 708 | 709 | 710 | 711 | 712 | 713 | 714 | 715 | 716 | 717 | 718 | 719 | 720 | 721 | 722 | 723 | 724 | 725 | 726 | 727 | 728 | 729 | 730 | 731 | 732 | 733 | 734 | 735 | 736 | 737 | 738 | 739 | 740 | 741 | 742 | 743 | 744 | 745 | 746 | 747 | 748 | 749 | 750 | 751 | 752 | 753 | 754 | 755 | 756 | 757 | 758 | 759 | 760 | 761 | 762 | 763 | 764 | 765 | 766 | 767 | 768 | 769 | 770 | 771 | 772 | 773 | 774 | 775 | 776 | 777 | 778 | 779 | 780 | 781 | 782 | 783 | 784 | 785 | 786 | 787 | 788 | 789 | 790 | 791 | 792 | 793 | 794 | 795 | 796 | 797 | 798 | 799 | 800 | 801 | 802 | 803 | 804 | 805 | 806 | 807 | 808 | 809 | 810 | 811 | 812 | 813 | 814 | 815 | 816 | 817 | 818 | 819 | 820 | 821 | 822 | 823 | 824 | 825 | 826 | 827 | 828 | 829 | 830 | 831 | 832 | 833 | 834 | 835 | 836 | 837 | 838 | 839 | 840 | 841 | 842 | 843 | 844 | 845 | 846 | 847 | 848 | 849 | 850 | 851 | 852 | 853 | 854 | 855 | 856 | 857 | 858 | 859 | 860 | 861 | 862 | 863 | 864 | 865 | 866 | 867 | 868 | 869 | 870 | 871 | 872 | 873 | 874 | 875 | 876 | 877 | 878 | 879 | 880 | 881 | 882 | 883 | 884 | 885 | 886 | 887 | 888 | 889 | 890 | 891 | 892 | 893 | 894 | 895 | 896 | 897 | 898 | 899 | 900 | 901 | 902 | 903 | 904 | 905 | 906 | 907 | 908 | 909 | 910 | 911 | 912 | 913 | 914 | 915 | 916 | 917 | 918 | 919 | 920 | 921 | 922 | 923 | 924 | 925 | 926 | 927 | 928 | 929 | 930 | 931 | 932 | 933 | 934 | 935 | 936 | 937 | 938 | 939 | 940 | 941 | 942 | 943 | 944 | 945 | 946 | 947 | 948 | 949 | 950 | 951 | 952 | 953 | 954 | 955 | 956 | 957 | 958 | 959 | 960 | 961 | 962 | 963 | 964 | 965 | 966 | 967 | 968 | 969 | 970 | 971 | 972 | 973 | 974 | 975 | 976 | 977 | 978 | 979 | 980 | 981 | 982 | 983 | 984 | 985 | 986 | 987 | 988 | 989 | 990 | 991 | 992 | 993 | 994 | 995 | 996 | 997 | 998 | 999 | 1000 | 1001 | 1002 | 1003 | 1004 | 1005 | 1006 | 1007 | 1008 | 1009 | 1010 | 1011 | 1012 | 1013 | 1014 | 1015 | 1016 | 1017 | 1018 | 1019 | 1020 | 1021 | 1022 | 1023 | 1024 | 1025 | 1026 | 1027 | 1028 | 1029 | 1030 | 1031 | 1032 | 1033 | 1034 | 1035 | 1036 | 1037 | 1038 | 1039 | 1040 | 1041 | 1042 | 1043 | 1044 | 1045 | 1046 | 1047 | 1048 | 1049 | 1050 | 1051 | 1052 | 1053 | 1054 | 1055 | 1056 | 1057 | 1058 | 1059 | 1060 | 1061 | 1062 | 1063 | 1064 | 1065 | 1066 | 1067 | 1068 | 1069 | 1070 | 1071 | 1072 | 1073 | 1074 | 1075 | 1076 | 1077 | 1078 | 1079 | 1080 | 1081 | 1082 | 1083 | 1084 | 1085 | 1086 | 1087 | 1088 | 1089 | 1090 | 1091 | 1092 | 1093 | 1094 | 1095 | 1096 | 1097 | 1098 | 1099 | 1100 | 1101 | 1102 | 1103 | 1104 | 1105 | 1106 | 1107 | 1108 | 1109 | 1110 | 1111 | 1112 | 1113 | 1114 | 1115 | 1116 | 1117 | 1118 | 1119 | 1120 | 1121 | 1122 | 1123 | 1124 | 1125 | 1126 | 1127 | 1128 | 1129 | 1130 | 1131 | 1132 | 1133 | 1134 | 1135 | 1136 | 1137 | 1138 | 1139 | 1140 | 1141 | 1142 | 1143 | 1144 | 1145 | 1146 | 1147 | 1148 | 1149 | 1150 | 1151 | 1152 | 1153 | 1154 | 1155 | 1156 | 1157 | 1158 | 1159 | 1160 | 1161 | 1162 | 1163 | 1164 | 1165 | 1166 | 1167 | 1168 | 1169 | 1170 | 1171 | 1172 | 1173 | 1174 | 1175 | 1176 | 1177 | 1178 | 1179 | 1180 | 1181 | 1182 | 1183 | 1184 | 1185 | 1186 | 1187 | 1188 | 1189 | 1190 | 1191 | 1192 | 1193 | 1194 | 1195 | 1196 | 1197 | 1198 | 1199 | 1200 | 1201 | 1202 | 1203 | 1204 | 1205 | 1206 | 1207 | 1208 | 1209 | 1210 | 1211 | 1212 | 1213 | 1214 | 1215 | 1216 | 1217 | 1218 | 1219 | 1220 | 1221 | 1222 | 1223 | 1224 | 1225 | 1226 | 1227 | 1228 | 1229 | 1230 | 1231 | 1232 | 1233 | 1234 | 1235 | 1236 | 1237 | 1238 | 1239 | 1240 | 1241 | 1242 | 1243 | 1244 | 1245 | 1246 | 1247 | 1248 | 1249 | 1250 | 1251 | 1252 | 1253 | 1254 | 1255 | 1256 | 1257 | 1258 | 1259 | 1260 | 1261 | 1262 | 1263 | 1264 | 1265 | 1266 | 1267 | 1268 | 1269 | 1270 | 1271 | 1272 | 1273 | 1274 | 1275 | 1276 | 1277 | 1278 | 1279 | 1280 | 1281 | 1282 | 1283 | 1284 | 1285 | 1286 | 1287 | 1288 | 1289 | 1290 | 1291 | 1292 | 1293 | 1294 | 1295 | 1296 | 1297 | 1298 | 1299 | 1300 | 1301 | 1302 | 1303 | 1304 | 1305 | 1306 | 1307 | 1308 | 1309 | 1310 | 1311 | 1312 | 1313 | 1314 | 1315 | 1316 | 1317 | 1318 | 1319 | 1320 | 1321 | 1322 | 1323 | 1324 | 1325 | 1326 | 1327 | 1328 | 1329 | 1330 | 1331 | 1332 | 1333 | 1334 | 1335 | 1336 | 1337 | 1338 | 1339 | 1340 | 1341 | 1342 | 1343 | 1344 | 1345 | 1346 | 1347 | 1348 | 1349 | 1350 | 1351 | 1352 | 1353 | 1354 | 1355 | 1356 | 1357 | 1358 | 1359 | 1360 | 1361 | 1362 | 1363 | 1364 | 1365 | 1366 | 1367 | 1368 | 1369 | 1370 | 1371 | 1372 | 1373 | 1374 | 1375 | 1376 | 1377 | 1378 | 1379 | 1380 | 1381 | 1382 | 1383 | 1384 | 1385 | 1386 | 1387 | 1388 | 1389 | 1390 | 1391 | 1392 | 1393 | 1394 | 1395 | 1396 | 1397 | 1398 | 1399 | 1400 | 1401 | 1402 | 1403 | 1404 | 1405 | 1406 | 1407 | 1408 | 1409 | 1410 | 1411 | 1412 | 1413 | 1414 | 1415 | 1416 | 1417 | 1418 | 1419 | 1420 | 1421 | 1422 | 1423 | 1424 | 1425 | 1426 | 1427 | 1428 | 1429 | 1430 | 1431 | 1432 | 1433 | 1434 | 1435 | 1436 | 1437 | 1438 | 1439 | 1440 | 1441 | 1442 | 1443 | 1444 | 1445 | 1446 | 1447 | 1448 | 1449 | 1450 | 1451 | 1452 | 1453 | 1454 | 1455 | 1456 | 1457 | 1458 | 1459 | 1460 | 1461 | 1462 | 1463 | 1464 | 1465 | 1466 | 1467 | 1468 | 1469 | 1470 | 1471 | 1472 | 1473 | 1474 | 1475 | 1476 | 1477 | 1478 | 1479 | 1480 | 1481 | 1482 | 1483 | 1484 | 1485 | 1486 | 1487 | 1488 | 1489 | 1490 | 1491 | 1492 | 1493 | 1494 | 1495 | 1496 | 1497 | 1498 | 1499 | 1500 | 1501 | 1502 | 1503 | 1504 | 1505 | 1506 | 1507 | 1508 | 1509 | 1510 | 1511 | 1512 | 1513 | 1514 | 1515 | 1516 | 1517 | 1518 | 1519 | 1520 | 1521 | 1522 | 1523 | 1524 | 1525 | 1526 | 1527 | 1528 | 1529 | 1530 | 1531 | 1532 | 1533 | 1534 | 1535 | 1536 | 1537 | 1538 | 1539 | 1540 | 1541 | 1542 | 1543 | 1544 | 1545 | 1546 | 1547 | 1548 | 1549 | 1550 | 1551 | 1552 | 1553 | 1554 | 1555 | 1556 | 1557 | 1558 | 1559 | 1560 | 1561 | 1562 | 1563 | 1564 | 1565 | 1566 | 1567 | 1568 | 1569 | 1570 | 1571 | 1572 | 1573 | 1574 | 1575 | 1576 | 1577 | 1578 | 1579 | 1580 | 1581 | 1582 | 1583 | 1584 | 1585 | 1586 | 1587 | 1588 | 1589 | 1590 | 1591 | 1592 | 1593 | 1594 | 1595 | 1596 | 1597 | 1598 | 1599 | 1600 | 1601 | 1602 | 1603 | 1604 | 1605 | 1606 | 1607 | 1608 | 1609 | 1610 | 1611 | 1612 | 1613 | 1614 | 1615 | 1616 | 1617 | 1618 | 1619 | 1620 | 1621 | 1622 | 1623 | 1624 | 1625 | 1626 | 1627 | 1628 | 1629 | 1630 | 1631 | 1632 | 1633 | 1634 | 1635 | 1636 | 1637 | 1638 | 1639 | 1640 | 1641 | 1642 | 1643 | 1644 | 1645 | 1646 | 1647 | 1648 | 1649 | 1650 | 1651 | 1652 | 1653 | 1654 | 1655 | 1656 | 1657 | 1658 | 1659 | 1660 | 1661 | 1662 | 1663 | 1664 | 1665 | 1666 | 1667 | 1668 | 1669 | 1670 | 1671 | 1672 | 1673 | 1674 | 1675 | 1676 | 1677 | 1678 | 1679 | 1680 | 1681 | 1682 | 1683 | 1684 | 1685 | 1686 | 1687 | 1688 | 1689 | 1690 | 1691 | 1692 | 1693 | 1694 | 1695 | 1696 | 1697 | 1698 | 1699 | 1700 | 1701 | 1702 | 1703 | 1704 | 1705 | 1706 | 1707 | 1708 | 1709 | 1710 | 1711 | 1712 | 1713 | 1714 | 1715 | 1716 | 1717 | 1718 | 1719 | 1720 | 1721 | 1722 | 1723 | 1724 | 1725 | 1726 | 1727 | 1728 | 1729 | 1730 | 1731 | 1732 | 1733 | 1734 | 1735 | 1736 | 1737 | 1738 | 1739 | 1740 | 1741 | 1742 | 1743 | 1744 | 1745 | 1746 | 1747 | 1748 | 1749 | 1750 | 1751 | 1752 | 1753 | 1754 | 1755 | 1756 | 1757 | 1758 | 1759 | 1760 | 1761 | 1762 | 1763 | 1764 | 1765 | 1766 | 1767 | 1768 | 1769 | 1770 | 1771 | 1772 | 1773 | 1774 | 1775 | 1776 | 1777 | 1778 | 1779 | 1780 | 1781 | 1782 | 1783 | 1784 | 1785 | 1786 | 1787 | 1788 | 1789 | 1790 | 1791 | 1792 | 1793 | 1794 | 1795 | 1796 | 1797 | 1798 | 1799 | 1800 | 1801 | 1802 | 1803 | 1804 | 1805 | 1806 | 1807 | 1808 | 1809 | 1810 | 1811 | 1812 | 1813 | 1814 | 1815 | 1816 | 1817 | 1818 | 1819 | 1820 | 1821 | 1822 | 1823 | 1824 | 1825 | 1826 | 1827 | 1828 | 1829 | 1830 | 1831 | 1832 | 1833 | 1834 | 1835 | 1836 | 1837 | 1838 | 1839 | 1840 | 1841 | 1842 | 1843 | 1844 | 1845 | 1846 | 1847 | 1848 | 1849 | 1850 | 1851 | 1852 | 1853 | 1854 | 1855 | 1856 | 1857 | 1858 | 1859 | 1860 | 1861 | 1862 | 1863 | 1864 | 1865 | 1866 | 1867 | 1868 | 1869 | 1870 | 1871 | 1872 | 1873 | 1874 | 1875 | 1876 | 1877 | 1878 | 1879 | 1880 | 1881 | 1882 | 1883 | 1884 | 1885 | 1886 | 1887 | 1888 | 1889 | 1890 | 1891 | 1892 | 1893 | 1894 | 1895 | 1896 | 1897 | 1898 | 1899 | 1900 | 1901 | 1902 | 1903 | 1904 | 1905 | 1906 | 1907 | 1908 | 1909 | 1910 | 1911 | 1912 | 1913 | 1914 | 1915 | 1916 | 1917 | 1918 | 1919 | 1920 | 1921 | 1922 | 1923 | 1924 | 1925 | 1926 | 1927 | 1928 | 1929 | 1930 | 1931 | 1932 | 1933 | 1934 | 1935 | 1936 | 1937 | 1938 | 1939 | 1940 | 1941 | 1942 | 1943 | 1944 | 1945 | 1946 | 1947 | 1948 | 1949 | 1950 | 1951 | 1952 | 1953 | 1954 | 1955 | 1956 | 1957 | 1958 | 1959 | 1960 | 1961 | 1962 | 1963 | 1964 | 1965 | 1966 | 1967 | 1968 | 1969 | 1970 | 1971 | 1972 | 1973 | 1974 | 1975 | 1976 | 1977 | 1978 | 1979 | 1980 | 1981 | 1982 | 1983 | 1984 | 1985 | 1986 | 1987 | 1988 | 1989 | 1990 | 1991 | 1992 | 1993 | 1994 | 1995 | 1996 | 1997 | 1998 | 1999 | 2000 | 2001 | 2002 | 2003 | 2004 | 2005 | 2006 | 2007 | 2008 | 2009 | 2010 | 2011 | 2012 | 2013 | 2014 | 2015 | 2016 | 2017 | 2018 | 2019 | 2020 | 2021 | 2022 | 2023 | 2024 | 2025 | 2026 | 2027 | 2028 | 2029 | 2030 | 2031 | 2032 | 2033 | 2034 | 2035 | 2036 | 2037 | 2038 | 2039 | 2040 | 2041 | 2042 | 2043 | 2044 | 2045 | 2046 | 2047
1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 2 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 2 | 2 | 2 | 2 | 1 | 2 | 0 | 4 | 0 | 4 | 3 | 3 | 3 | 2 | 2 | 0 | 2 | 2 | 2 | 2 | 1 | 3 | 1 | 2 | 2 | 2 | 1 | 1 | 2 | 2 | 1 | 0 | 1 | 0 | 2 | 2 | 1 | 2 | 0 | 1 | 1 | 0 | 0 | 1 | 4 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 5 | 0 | 2 | 2 | 2 | 2 | 0 | 0 | 1 | 2 | 3 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 2 | 2 | 1 | 0 | 0 | 1 | 3 | 2 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 2 | 1 | 2 | 1 | 1 | 2 | 2 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 3 | 0 | 1 | 2 | 1 | 4 | 2 | 1 | 1 | 3 | 1 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 3 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 2 | 1 | 0 | 0 | 2 | 0 | 0 | 0 | 1 | 2 | 0 | 1 | 1 | 3 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 1 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 3 | 0 | 2 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 1 | 2 | 1 | 0 | 2 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 2 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 3 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 2 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 2 | 1380 | 102 | 115 | 139 | 136 | 151 | 170 | 189 | 205 | 225 | 244 | 256 | 290 | 326 | 366 | 398 | 464 | 537 | 570 | 629 | 764 | 853 | 893 | 984 | 1101 | 1204 | 1343 | 1485 | 1574 | 1694 | 1906 | 2111 | 2234 | 2335 | 2435 | 2533 | 2574 | 2690 | 2808 | 2854 | 2881 | 3088 | 3138 | 2917 | 2941 | 3021 | 3049 | 3029 | 2988 | 3019 | 3043 | 3029 | 3228 | 3092 | 3061 | 3088 | 3136 | 3080 | 3050 | 3097 | 3079 | 3142 | 2976 | 3285 | 3077 | 3031 | 3057 | 3084 | 3074 | 3062 | 3075 | 3022 | 3074 | 3217 | 3252 | 3073 | 3033 | 3159 | 3179 | 3103 | 3097 | 3179 | 3138 | 3258 | 3269 | 3310 | 3176 | 3146 | 3181 | 3144 | 3141 | 3140 | 3164 | 3202 | 3264 | 3254 | 3129 | 3047 | 3015 | 2979 | 2960 | 2970 | 2978 | 2965 | 2970 | 3011 | 3039 | 2972 | 2925 | 2934 | 2939 | 2884 | 2888 | 2913 | 2917 | 2944 | 2959 | 2790 | 2651 | 2664 | 2649 | 2676 | 2680 | 2655 | 2634 | 2643 | 2690 | 2711 | 2728 | 2690 | 2666 | 2676 | 2698 | 2700 | 2711 | 2696 | 2703 | 2762 | 2766 | 2751 | 2756 | 2760 | 2762 | 2733 | 2663 | 2656 | 2726 | 2781 | 2820 | 2804 | 2842 | 2845 | 2811 | 2815 | 2892 | 2883 | 2914 | 2944 | 2984 | 2950 | 3043 | 3091 | 3031 | 3140 | 3174 | 3205 | 3362 | 3247 | 3454 | 3469 | 3488 | 3453 | 3512 | 3596 | 3619 | 3713 | 3812 | 3872 | 3878 | 3909 | 3990 | 4055 | 4083 | 4118 | 4053 | 4050 | 4156 | 4117 | 4098 | 4172 | 4194 | 4101 | 4027 | 4033 | 4005 | 3965 | 3946 | 3915 | 3868 | 3829 | 3788 | 3745 | 3738 | 3689 | 3630 | 3623 | 3613 | 3588 | 3567 | 3507 | 3418 | 3388 | 3321 | 3353 | 3375 | 3294 | 3226 | 3211 | 3253 | 3270 | 3218 | 3167 | 3109 | 2991 | 3006 | 3026 | 2963 | 2919 | 2912 | 2954 | 2884 | 2846 | 2833 | 2838 | 2781 | 2732 | 2762 | 2734 | 2710 | 2694 | 2658 | 2675 | 2717 | 2706 | 2609 | 2578 | 2635 | 2581 | 2462 | 2410 | 2459 | 2518 | 2449 | 2402 | 2414 | 2422 | 2390 | 2371 | 2341 | 2354 | 2343 | 2349 | 2279 | 2349 | 2264 | 2211 | 2284 | 2181 | 2254 | 2271 | 2145 | 2197 | 2158 | 2163 | 2162 | 2111 | 2080 | 2144 | 2087 | 2126 | 2100 | 2113 | 2075 | 2047 | 2045 | 2026 | 1997 | 1978 | 1942 | 1912 | 1892 | 1884 | 1922 | 1921 | 1911 | 1925 | 1909 | 1904 | 1901 | 1863 | 1860 | 1896 | 1891 | 1868 | 1840 | 1833 | 1867 | 1864 | 1826 | 1802 | 1799 | 1839 | 1832 | 1806 | 1781 | 1763 | 1798 | 1763 | 1715 | 1711 | 1681 | 1612 | 1651 | 1671 | 1634 | 1628 | 1641 | 1639 | 1597 | 1537 | 1516 | 1505 | 1462 | 1506 | 1501 | 1485 | 1457 | 1430 | 1436 | 1441 | 1386 | 1390 | 1367 | 1341 | 1333 | 1351 | 1336 | 1306 | 1311 | 1311 | 1301 | 1254 | 1234 | 1306 | 1279 | 1224 | 1256 | 1265 | 1242 | 1226 | 1215 | 1230 | 1237 | 1219 | 1198 | 1178 | 1165 | 1163 | 1175 | 1155 | 1161 | 1165 | 1186 | 1186 | 1175 | 1186 | 1182 | 1124 | 1084 | 1094 | 1096 | 1090 | 1104 | 1116 | 1101 | 1039 | 1067 | 1042 | 1030 | 1051 | 1083 | 1084 | 1045 | 1019 | 1052 | 1084 | 1053 | 1018 | 1036 | 1052 | 1032 | 1016 | 1015 | 1021 | 1035 | 1012 | 970 | 967 | 1005 | 1047 | 1024 | 996 | 1006 | 1009 | 999 | 995 | 996 | 1003 | 1042 | 1010 | 1009 | 1003 | 1017 | 1038 | 955 | 975 | 1059 | 1026 | 999 | 1032 | 1075 | 1008 | 1027 | 1109 | 1078 | 1116 | 1119 | 1141 | 1160 | 1190 | 1209 | 1232 | 1292 | 1331 | 1344 | 1358 | 1425 | 1486 | 1564 | 1642 | 1711 | 1798 | 1904 | 1993 | 2168 | 2348 | 2455 | 2630 | 2772 | 2950 | 3200 | 3410 | 3649 | 3876 | 4264 | 4589 | 4946 | 5294 | 5639 | 6068 | 6574 | 7087 | 7567 | 8077 | 8479 | 9147 | 9624 | 10376 | 11316 | 11804 | 12301 | 12927 | 13665 | 14416 | 15028 | 15720 | 16299 | 16936 | 17565 | 18241 | 18925 | 19324 | 19751 | 20362 | 20714 | 21078 | 21294 | 21381 | 21704 | 21939 | 21943 | 21880 | 21724 | 21479 | 21227 | 20944 | 20513 | 19982 | 19487 | 18975 | 18418 | 17832 | 17178 | 16405 | 15714 | 14983 | 14153 | 13349 | 12685 | 11725 | 11034 | 10363 | 9561 | 8816 | 8218 | 7602 | 6913 | 6455 | 5868 | 5484 | 4905 | 4500 | 4086 | 3691 | 3538 | 3193 | 2933 | 2702 | 2478 | 2325 | 2160 | 2023 | 1896 | 1810 | 1731 | 1650 | 1607 | 1525 | 1439 | 1456 | 1457 | 1393 | 1409 | 1383 | 1368 | 1416 | 1366 | 1295 | 1339 | 1371 | 1339 | 1350 | 1338 | 1310 | 1327 | 1308 | 1282 | 1274 | 1254 | 1292 | 1276 | 1304 | 1311 | 1281 | 1275 | 1267 | 1231 | 1257 | 1272 | 1240 | 1244 | 1285 | 1293 | 1264 | 1205 | 1162 | 1224 | 1232 | 1222 | 1265 | 1274 | 1230 | 1218 | 1172 | 1138 | 1160 | 1175 | 1197 | 1200 | 1195 | 1178 | 1145 | 1145 | 1178 | 1186 | 1170 | 1178 | 1183 | 1159 | 1143 | 1143 | 1138 | 1138 | 1147 | 1166 | 1150 | 1140 | 1137 | 1119 | 1147 | 1156 | 1171 | 1139 | 1111 | 1092 | 1134 | 1109 | 1079 | 1083 | 1128 | 1097 | 1148 | 1099 | 1177 | 1140 | 1107 | 1131 | 1139 | 1111 | 1102 | 1115 | 1129 | 1136 | 1130 | 1127 | 1146 | 1174 | 1143 | 1113 | 1169 | 1129 | 1139 | 1185 | 1226 | 1177 | 1200 | 1186 | 1203 | 1217 | 1260 | 1270 | 1258 | 1285 | 1311 | 1286 | 1269 | 1313 | 1365 | 1349 | 1274 | 1335 | 1384 | 1405 | 1411 | 1410 | 1441 | 1417 | 1420 | 1399 | 1418 | 1417 | 1416 | 1434 | 1446 | 1436 | 1451 | 1484 | 1462 | 1418 | 1451 | 1481 | 1439 | 1463 | 1493 | 1445 | 1444 | 1437 | 1411 | 1422 | 1412 | 1389 | 1390 | 1390 | 1349 | 1294 | 1285 | 1292 | 1300 | 1309 | 1325 | 1307 | 1289 | 1256 | 1280 | 1240 | 1254 | 1220 | 1179 | 1171 | 1131 | 1108 | 1104 | 1092 | 1108 | 1130 | 1106 | 1100 | 1124 | 1097 | 992 | 1001 | 1080 | 1074 | 996 | 959 | 988 | 987 | 970 | 945 | 899 | 914 | 912 | 879 | 876 | 881 | 888 | 890 | 856 | 879 | 869 | 829 | 813 | 804 | 788 | 784 | 777 | 766 | 761 | 771 | 772 | 776 | 791 | 781 | 762 | 753 | 758 | 766 | 736 | 725 | 702 | 698 | 718 | 731 | 702 | 660 | 657 | 657 | 645 | 645 | 658 | 659 | 640 | 634 | 624 | 602 | 599 | 609 | 626 | 633 | 614 | 589 | 590 | 603 | 583 | 568 | 580 | 587 | 579 | 562 | 541 | 540 | 550 | 550 | 543 | 537 | 551 | 552 | 550 | 527 | 502 | 535 | 516 | 519 | 517 | 520 | 530 | 508 | 483 | 471 | 452 | 454 | 457 | 449 | 456 | 452 | 447 | 457 | 438 | 440 | 460 | 465 | 455 | 415 | 458 | 433 | 420 | 446 | 429 | 420 | 411 | 417 | 408 | 404 | 403 | 368 | 375 | 350 | 351 | 381 | 379 | 402 | 381 | 396 | 377 | 376 | 356 | 352 | 343 | 335 | 336 | 324 | 328 | 342 | 346 | 348 | 328 | 319 | 345 | 339 | 307 | 311 | 332 | 327 | 314 | 311 | 311 | 301 | 309 | 316 | 302 | 291 | 296 | 297 | 280 | 271 | 283 | 298 | 291 | 290 | 294 | 284 | 272 | 277 | 270 | 263 | 269 | 279 | 287 | 282 | 276 | 278 | 284 | 289 | 279 | 281 | 295 | 274 | 272 | 294 | 288 | 278 | 272 | 269 | 278 | 258 | 265 | 261 | 260 | 267 | 275 | 279 | 275 | 273 | 272 | 274 | 288 | 272 | 279 | 252 | 290 | 293 | 257 | 269 | 259 | 279 | 303 | 296 | 281 | 299 | 306 | 295 | 292 | 312 | 337 | 328 | 340 | 360 | 346 | 368 | 396 | 380 | 392 | 391 | 430 | 468 | 452 | 449 | 486 | 492 | 503 | 535 | 559 | 592 | 628 | 644 | 670 | 701 | 704 | 741 | 821 | 849 | 861 | 915 | 957 | 972 | 1002 | 1057 | 1102 | 1165 | 1219 | 1236 | 1271 | 1339 | 1417 | 1486 | 1513 | 1577 | 1620 | 1702 | 1771 | 1793 | 1803 | 1813 | 1887 | 1970 | 2001 | 2078 | 2151 | 2160 | 2170 | 2194 | 2212 | 2231 | 2285 | 2312 | 2295 | 2333 | 2303 | 2507 | 2417 | 2365 | 2408 | 2416 | 2338 | 2314 | 2273 | 2244 | 2225 | 2194 | 2184 | 2147 | 2053 | 2033 | 2053 | 2018 | 1906 | 1800 | 1764 | 1912 | 1696 | 1603 | 1570 | 1466 | 1406 | 1338 | 1266 | 1223 | 1159 | 1110 | 1075 | 1034 | 959 | 858 | 851 | 798 | 734 | 726 | 689 | 663 | 581 | 526 | 494 | 442 | 432 | 413 | 392 | 361 | 334 | 295 | 260 | 243 | 238 | 233 | 214 | 188 | 172 | 176 | 168 | 154 | 138 | 124 | 121 | 121 | 119 | 116 | 114 | 107 | 92 | 88 | 95 | 99 | 95 | 83 | 79 | 74 | 76 | 80 | 88 | 83 | 73 | 77 | 79 | 70 | 66 | 68 | 74 | 72 | 70 | 76 | 71 | 65 | 65 | 58 | 55 | 69 | 65 | 68 | 64 | 65 | 65 | 66 | 71 | 71 | 71 | 68 | 59 | 54 | 62 | 73 | 77 | 67 | 65 | 62 | 77 | 62 | 58 | 69 | 80 | 67 | 51 | 62 | 49 | 57 | 76 | 69 | 56 | 52 | 58 | 61 | 60 | 59 | 56 | 59 | 67 | 59 | 61 | 48 | 55 | 54 | 63 | 64 | 50 | 54 | 56 | 54 | 62 | 58 | 50 | 55 | 59 | 59 | 64 | 63 | 61 | 57 | 54 | 65 | 65 | 58 | 55 | 47 | 43 | 55 | 61 | 55 | 57 | 61 | 57 | 58 | 58 | 49 | 49 | 58 | 57 | 52 | 60 | 54 | 52 | 49 | 51 | 54 | 54 | 54 | 54 | 55 | 54 | 51 | 50 | 50 | 48 | 54 | 50 | 48 | 53 | 59 | 55 | 63 | 45 | 53 | 40 | 53 | 58 | 55 | 52 | 51 | 51 | 46 | 45 | 46 | 43 | 42 | 42 | 43 | 45 | 46 | 43 | 45 | 43 | 44 | 52 | 47 | 45 | 46 | 48 | 42 | 37 | 39 | 44 | 47 | 45 | 39 | 40 | 40 | 38 | 34 | 37 | 37 | 42 | 45 | 41 | 39 | 44 | 40 | 35 | 43 | 46 | 41 | 39 | 42 | 40 | 38 | 37 | 39 | 37 | 31 | 28 | 25 | 28 | 31 | 30 | 32 | 33 | 27 | 23 | 29 | 35 | 35 | 32 | 29 | 29 | 32 | 35 | 33 | 28 | 30 | 29 | 34 | 39 | 33 | 32 | 32 | 34 | 32 | 31 | 28 | 28 | 30 | 32 | 32 | 32 | 32 | 24 | 26 | 30 | 34 | 27 | 24 | 23 | 29 | 27 | 24 | 18 | 21 | 25 | 25 | 22 | 23 | 27 | 24 | 25 | 23 | 22 | 29 | 31 | 24 | 26 | 30 | 18 | 15 | 16 | 22 | 21 | 23 | 25 | 29 | 27 | 25 | 28 | 27 | 25 | 28 | 32 | 32 | 24 | 21 | 21 | 23 | 20 | 24 | 23 | 23 | 23 | 24 | 22 | 19 | 23 | 27 | 27 | 25 | 22 | 22 | 23 | 30 | 33 | 27 | 24 | 23 | 19 | 23 | 25 | 23 | 23 | 23 | 23 | 22 | 19 | 19 | 23 | 22 | 22 | 22 | 23 | 24 | 20 | 21 | 24 | 19 | 15 | 24 | 25 | 20 | 17 | 20 | 22 | 24 | 22 | 23 | 29 | 22 | 21 | 20 | 24 | 17 | 18 | 15 | 16 | 20 | 31 | 27 | 22 | 20 | 21 | 22 | 22 | 22 | 21 | 20 | 20 | 18 | 27 | 23 | 28 | 17 | 28 | 24 | 26 | 27 | 26 | 27 | 28 | 25 | 24 | 22 | 22 | 22 | 19 | 26 | 24 | 30 | 35 | 30 | 26 | 21 | 23 | 20 | 25 | 31 | 28 | 31 | 25 | 23 | 28 | 35 | 23 | 22 | 26 | 34 | 37 | 28 | 24 | 25 | 27 | 32 | 28 | 27 | 31 | 29 | 30 | 30 | 31 | 33 | 32 | 32 | 27 | 27 | 30 | 27 | 32 | 31 | 31 | 26 | 26 | 35 | 33 | 32 | 35 | 23 | 34 | 29 | 35 | 29 | 24 | 33 | 29 | 29 | 36 | 20 | 39 | 38 | 32 | 41 | 38 | 38 | 26 | 39 | 34 | 29 | 35 | 35 | 33 | 30 | 29 | 30 | 31 | 31 | 39 | 33 | 32 | 19 | 19 | 32 | 23 | 24 | 19 | 26 | 28 | 33 | 30 | 27 | 27 | 15 | 17 | 27 | 19 | 15 | 13 | 15 | 22 | 19 | 12 | 15 | 17 | 15 | 11 | 10 | 17 | 15 | 14 | 14 | 13 | 16 | 15 | 12 | 10 | 10 | 11 | 13 | 13 | 12 | 11 | 9 | 8 | 11 | 9 | 7 | 6 | 7 | 8 | 9 | 12 | 9 | 9 | 7 | 8 | 11 | 8 | 6 | 5 | 6 | 6 | 6 | 6 | 7 | 3 | 1 | 3 | 9 | 7 | 9 | 5 | 4 | 8 | 7 | 5 | 3 | 8 | 4 | 10 | 6 | 9 | 8 | 5 | 8 | 3 | 2 | 5 | 5 | 11 | 5 | 11 | 5 | 3 | 6 | 5 | 6 | 4 | 6 | 5 | 6 | 8 | 7 | 6 | 4 | 9 | 7 | 8 | 2 | 6 | 7 | 6 | 6 | 4 | 4 | 3 | 4 | 4 | 5 | 4 | 4 | 4 | 6 | 5 | 3 | 6 | 8 | 11 | 10 | 5 | 3 | 5 | 5 | 4 | 5 | 5 | 7 | 6 | 3 | 3 | 6 | 5 | 5 | 3 | 4 | 7 | 6 | 6 | 5 | 6 | 9 | 6 | 6 | 2 | 6 | 5 | 2 | 2 | 5 | 5 | 6 | 5 | 4 | 4 | 6 | 2 | 11 | 6 | 3 | 6 | 4 | 1 | 3 | 7 | 6 | 4 | 6 | 4 | 8 | 5 | 9 | 5 | 5 | 9 | 6 | 3 | 1 | 1 | 3 | 3 | 2 | 6 | 1 | 6 | 2 | 6 | 1 | 3 | 4 | 2 | 4 | 6 | 4 | 5 | 3 | 3 | 5 | 3 | 2 | 4 | 3 | 3 | 6 | 4 | 1 | 3 | 5 | 6 | 2 | 2 | 3 | 3 | 3 | 3 | 2 | 3 | 2 | 3 | 4 | 4 | 5 | 5 | 3 | 2 | 2 | 5 | 3 | 3 | 2 | 4 | 5 | 4 | 2 | 3 | 3 | 1 | 0 | 1 | 3 | 4 | 3 | 3 | 3 | 2 | 2 | 2 | 2 | 2 | 1 | 2 | 3 | 2 | 2 | 2 | 1 | 3 | 2 | 4 | 3 | 0 | 2 | 2 | 1 | 1 | 1 | 0 | 0 | 0 | 2 | 3 | 1 | 2 | 2 | 1 | 5 | 2 | 0 | 2 | 0 | 1 | 3 | 2 | 1 | 2 | 1 | 0 | 1 | 1 | 2 | 0 | 2 | 2 | 0 | 1 | 1 | 0 | 2 | 4 | 3 | 2 | 1 | 1 | 1 | 2 | 2 | 1 | 1 | 2 | 1 | 1 | 2 | 1 | 1 | 2 | 3 | 1 | 2 | 3 | 2 | 2 | 2 | 2 | 1 | 0 | 1 | 2 | 2 | 3 | 2 | 2 | 2 | 1 | 2 | 2 | 3 | 1 | 2 | 2 | 2 | 1 | 2 | 1 | 3 | 2 | 2 | 1 | 1 | 3 | 1 | 2 | 4 | 2 | 2 | 1 | 1 | 1 | 3 | 1 | 1 | 1 | 2 | 3 | 1 | 4 | 0 | 2 | 5 | 3 | 2 | 0 | 0 | 0 | 1 | 1 | 2 | 3 | 2 | 2 | 1 | 2 | 1 | 4 | 3 | 1 | 1 | 0 | 2 | 3 | 2 | 3 | 0 | 2 | 1 | 0 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 1 | 1 | 3 | 2 | 1 | 2 | 2 | 2 | 1 | 2 | 2 | 3 | 3 | 2 | 0 | 1 | 2 | 2 | 1 | 3 | 3 | 4 | 3 | 1 | 1 | 0 | 0 | 1 | 2 | 2 | 1 | 0 | 1 | 5 | 2 | 1 | 0 | 1 | 1 | 3 | 5 | 3 | 1 | 1 | 3 | 1 | 4 | 1 | 3 | 4 | 1 | 4 | 2 | 1 | 3 | 4 | 1 | 3 | 2 | 5 | 3 | 1 | 2 | 2 | 2 | 1 | 3 | 4 | 2 | 1 | 1 | 1 | 1 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 3 | 1369 | 135 | 146 | 139 | 139 | 145 | 167 | 174 | 183 | 201 | 241 | 274 | 302 | 339 | 380 | 407 | 440 | 497 | 557 | 619 | 727 | 800 | 864 | 976 | 1094 | 1184 | 1297 | 1448 | 1554 | 1685 | 1916 | 2106 | 2177 | 2298 | 2380 | 2465 | 2548 | 2619 | 2734 | 2827 | 2933 | 2962 | 2999 | 2927 | 2975 | 3009 | 2953 | 3003 | 3047 | 3038 | 2990 | 3017 | 3226 | 3024 | 3137 | 3100 | 3064 | 3065 | 3026 | 3073 | 3000 | 3102 | 3016 | 3292 | 3063 | 2985 | 3036 | 3064 | 3087 | 3122 | 3144 | 3063 | 3046 | 3103 | 3290 | 3059 | 2998 | 3093 | 3070 | 3123 | 3160 | 3112 | 3115 | 3175 | 3196 | 3251 | 3151 | 3120 | 3098 | 3106 | 3112 | 3169 | 3220 | 3125 | 3086 | 3177 | 3149 | 3060 | 3008 | 2981 | 3006 | 3017 | 2992 | 2974 | 2919 | 2978 | 3063 | 2938 | 2908 | 2899 | 2918 | 2975 | 2971 | 2942 | 2937 | 2968 | 2965 | 2766 | 2612 | 2599 | 2626 | 2688 | 2697 | 2675 | 2660 | 2668 | 2687 | 2720 | 2715 | 2712 | 2675 | 2674 | 2696 | 2659 | 2631 | 2635 | 2689 | 2662 | 2676 | 2738 | 2757 | 2743 | 2750 | 2775 | 2779 | 2770 | 2759 | 2755 | 2754 | 2847 | 2878 | 2809 | 2777 | 2805 | 2863 | 2920 | 2959 | 2910 | 2922 | 2893 | 2845 | 3032 | 3056 | 3044 | 3149 | 3238 | 3215 | 3203 | 3235 | 3475 | 3493 | 3574 | 3536 | 3552 | 3574 | 3744 | 3748 | 3764 | 3765 | 3783 | 3816 | 3961 | 3986 | 4050 | 4115 | 4156 | 4070 | 4022 | 4083 | 4129 | 4054 | 4034 | 4035 | 4007 | 3995 | 3960 | 3877 | 3815 | 3798 | 3788 | 3762 | 3710 | 3681 | 3646 | 3585 | 3538 | 3523 | 3488 | 3460 | 3492 | 3437 | 3383 | 3383 | 3355 | 3332 | 3302 | 3217 | 3166 | 3130 | 3166 | 3169 | 3102 | 3117 | 3112 | 3021 | 2997 | 3004 | 2949 | 2865 | 2861 | 2864 | 2828 | 2802 | 2799 | 2786 | 2779 | 2757 | 2688 | 2683 | 2714 | 2694 | 2648 | 2642 | 2643 | 2578 | 2574 | 2633 | 2605 | 2548 | 2513 | 2477 | 2445 | 2415 | 2423 | 2416 | 2396 | 2396 | 2390 | 2380 | 2387 | 2428 | 2375 | 2322 | 2282 | 2299 | 2326 | 2262 | 2166 | 2210 | 2215 | 2223 | 2219 | 2173 | 2123 | 2038 | 2110 | 2088 | 2045 | 2009 | 2063 | 2032 | 2053 | 2061 | 2101 | 2036 | 2022 | 2015 | 1994 | 1943 | 1966 | 1944 | 1915 | 1920 | 1941 | 1943 | 1936 | 1900 | 1859 | 1860 | 1892 | 1890 | 1866 | 1846 | 1803 | 1772 | 1786 | 1819 | 1832 | 1826 | 1815 | 1784 | 1757 | 1755 | 1709 | 1691 | 1714 | 1724 | 1716 | 1697 | 1700 | 1644 | 1623 | 1631 | 1641 | 1652 | 1607 | 1571 | 1575 | 1538 | 1512 | 1539 | 1497 | 1481 | 1464 | 1462 | 1416 | 1393 | 1405 | 1429 | 1448 | 1412 | 1374 | 1352 | 1327 | 1310 | 1311 | 1341 | 1314 | 1264 | 1286 | 1306 | 1303 | 1302 | 1264 | 1236 | 1225 | 1195 | 1208 | 1249 | 1255 | 1229 | 1200 | 1209 | 1191 | 1168 | 1132 | 1130 | 1196 | 1136 | 1129 | 1118 | 1184 | 1163 | 1138 | 1175 | 1160 | 1147 | 1157 | 1138 | 1116 | 1097 | 1137 | 1088 | 1108 | 1092 | 1094 | 1122 | 1091 | 1065 | 1073 | 1087 | 1069 | 1055 | 1040 | 1056 | 1039 | 1016 | 1006 | 1003 | 1034 | 1035 | 1014 | 1036 | 1048 | 1044 | 1039 | 1013 | 991 | 1000 | 998 | 993 | 1014 | 1046 | 1035 | 1005 | 985 | 996 | 1035 | 1014 | 995 | 1010 | 999 | 1007 | 1005 | 1004 | 1024 | 1045 | 1007 | 1039 | 1014 | 1069 | 1054 | 1043 | 1074 | 1096 | 1098 | 1097 | 1118 | 1141 | 1200 | 1232 | 1264 | 1263 | 1283 | 1365 | 1392 | 1393 | 1421 | 1521 | 1611 | 1679 | 1777 | 1897 | 1992 | 2126 | 2247 | 2360 | 2577 | 2762 | 2917 | 3130 | 3364 | 3635 | 3868 | 4140 | 4464 | 4830 | 5243 | 5623 | 6002 | 6469 | 6927 | 7357 | 7903 | 8385 | 9051 | 9678 | 10276 | 11005 | 11688 | 12167 | 12832 | 13372 | 14096 | 14904 | 15515 | 16160 | 16667 | 17247 | 17838 | 18354 | 18964 | 19496 | 19990 | 20401 | 20880 | 21188 | 21266 | 21563 | 21800 | 21684 | 21618 | 21628 | 21412 | 20964 | 20593 | 20256 | 19945 | 19592 | 19072 | 18426 | 17778 | 17117 | 16306 | 15614 | 14922 | 14057 | 13396 | 12720 | 11860 | 11099 | 10411 | 9582 | 8845 | 8284 | 7647 | 7084 | 6428 | 5902 | 5438 | 4964 | 4448 | 4210 | 3789 | 3544 | 3260 | 2989 | 2741 | 2579 | 2317 | 2136 | 2009 | 1921 | 1804 | 1700 | 1647 | 1596 | 1563 | 1512 | 1454 | 1423 | 1401 | 1395 | 1400 | 1381 | 1329 | 1273 | 1285 | 1322 | 1337 | 1320 | 1297 | 1296 | 1275 | 1321 | 1332 | 1357 | 1317 | 1279 | 1265 | 1257 | 1266 | 1268 | 1277 | 1311 | 1285 | 1227 | 1236 | 1230 | 1229 | 1221 | 1236 | 1317 | 1308 | 1261 | 1234 | 1235 | 1200 | 1224 | 1237 | 1229 | 1209 | 1212 | 1223 | 1193 | 1171 | 1191 | 1205 | 1196 | 1169 | 1154 | 1163 | 1169 | 1157 | 1159 | 1160 | 1143 | 1136 | 1138 | 1143 | 1152 | 1144 | 1140 | 1149 | 1140 | 1127 | 1110 | 1128 | 1129 | 1146 | 1165 | 1149 | 1127 | 1128 | 1149 | 1134 | 1107 | 1087 | 1092 | 1112 | 1123 | 1163 | 1195 | 1102 | 1049 | 1050 | 1046 | 1061 | 1085 | 1059 | 1062 | 1090 | 1117 | 1117 | 1120 | 1113 | 1084 | 1126 | 1161 | 1134 | 1140 | 1142 | 1178 | 1237 | 1228 | 1205 | 1147 | 1200 | 1214 | 1202 | 1249 | 1271 | 1298 | 1299 | 1292 | 1322 | 1327 | 1302 | 1310 | 1328 | 1349 | 1383 | 1386 | 1413 | 1383 | 1391 | 1432 | 1403 | 1416 | 1425 | 1434 | 1470 | 1459 | 1458 | 1449 | 1453 | 1462 | 1442 | 1452 | 1439 | 1428 | 1456 | 1496 | 1480 | 1476 | 1472 | 1415 | 1374 | 1427 | 1444 | 1417 | 1409 | 1400 | 1381 | 1330 | 1314 | 1332 | 1347 | 1351 | 1341 | 1260 | 1240 | 1269 | 1277 | 1208 | 1192 | 1186 | 1165 | 1148 | 1165 | 1158 | 1143 | 1127 | 1110 | 1116 | 1109 | 1081 | 1028 | 1054 | 1023 | 1022 | 984 | 1041 | 1019 | 996 | 969 | 999 | 967 | 973 | 954 | 928 | 936 | 923 | 876 | 852 | 863 | 855 | 831 | 871 | 824 | 794 | 872 | 831 | 756 | 822 | 795 | 804 | 784 | 748 | 708 | 718 | 719 | 723 | 751 | 761 | 731 | 702 | 712 | 690 | 717 | 726 | 735 | 712 | 702 | 697 | 684 | 700 | 706 | 697 | 677 | 683 | 670 | 644 | 632 | 627 | 616 | 607 | 614 | 634 | 641 | 641 | 659 | 636 | 597 | 597 | 608 | 609 | 614 | 604 | 582 | 570 | 562 | 553 | 544 | 536 | 532 | 537 | 530 | 522 | 515 | 525 | 538 | 513 | 504 | 523 | 512 | 495 | 493 | 501 | 486 | 472 | 472 | 476 | 455 | 425 | 464 | 459 | 456 | 446 | 434 | 418 | 446 | 432 | 450 | 432 | 428 | 407 | 364 | 380 | 416 | 407 | 385 | 380 | 388 | 401 | 376 | 360 | 376 | 350 | 356 | 372 | 400 | 350 | 367 | 357 | 352 | 352 | 354 | 341 | 343 | 346 | 331 | 318 | 314 | 318 | 322 | 312 | 312 | 299 | 316 | 313 | 302 | 290 | 316 | 319 | 307 | 293 | 295 | 297 | 305 | 310 | 296 | 294 | 309 | 306 | 285 | 284 | 293 | 280 | 288 | 299 | 288 | 285 | 273 | 265 | 282 | 285 | 283 | 289 | 293 | 278 | 265 | 275 | 272 | 263 | 275 | 273 | 289 | 275 | 255 | 261 | 262 | 247 | 241 | 264 | 280 | 278 | 281 | 276 | 261 | 253 | 256 | 262 | 268 | 277 | 266 | 265 | 248 | 268 | 301 | 282 | 264 | 267 | 295 | 296 | 273 | 267 | 301 | 306 | 319 | 327 | 315 | 327 | 332 | 336 | 338 | 331 | 347 | 359 | 386 | 379 | 396 | 392 | 395 | 424 | 422 | 487 | 496 | 503 | 516 | 525 | 553 | 584 | 599 | 625 | 672 | 671 | 700 | 748 | 823 | 803 | 845 | 857 | 909 | 996 | 1061 | 1078 | 1125 | 1174 | 1211 | 1264 | 1305 | 1322 | 1360 | 1426 | 1507 | 1566 | 1618 | 1697 | 1722 | 1731 | 1832 | 1893 | 1894 | 1971 | 2034 | 2082 | 2155 | 2172 | 2148 | 2179 | 2267 | 2319 | 2376 | 2346 | 2345 | 2354 | 2416 | 2460 | 2369 | 2375 | 2392 | 2322 | 2312 | 2289 | 2277 | 2278 | 2270 | 2230 | 2161 | 2096 | 2057 | 2050 | 2007 | 1952 | 1867 | 1796 | 1733 | 1805 | 1667 | 1600 | 1524 | 1444 | 1397 | 1329 | 1263 | 1240 | 1149 | 1070 | 1059 | 1036 | 975 | 905 | 859 | 808 | 733 | 674 | 654 | 670 | 552 | 515 | 468 | 448 | 433 | 392 | 399 | 372 | 318 | 291 | 278 | 266 | 250 | 228 | 206 | 171 | 171 | 177 | 162 | 158 | 138 | 136 | 119 | 109 | 105 | 109 | 111 | 107 | 99 | 98 | 97 | 89 | 87 | 86 | 79 | 75 | 78 | 80 | 69 | 66 | 71 | 67 | 66 | 69 | 64 | 64 | 71 | 72 | 63 | 58 | 65 | 69 | 71 | 74 | 80 | 80 | 71 | 77 | 84 | 73 | 63 | 62 | 63 | 69 | 78 | 68 | 70 | 79 | 73 | 62 | 61 | 61 | 54 | 52 | 57 | 72 | 68 | 67 | 63 | 58 | 51 | 59 | 55 | 69 | 59 | 55 | 54 | 53 | 61 | 60 | 54 | 55 | 55 | 66 | 66 | 74 | 61 | 61 | 53 | 66 | 60 | 64 | 65 | 63 | 63 | 58 | 54 | 60 | 61 | 54 | 56 | 65 | 66 | 62 | 63 | 61 | 61 | 58 | 53 | 56 | 60 | 55 | 51 | 50 | 53 | 57 | 60 | 59 | 56 | 54 | 53 | 51 | 46 | 46 | 51 | 55 | 52 | 59 | 64 | 55 | 52 | 52 | 56 | 60 | 55 | 53 | 57 | 54 | 50 | 45 | 44 | 51 | 57 | 60 | 46 | 48 | 46 | 41 | 49 | 55 | 52 | 44 | 52 | 55 | 46 | 44 | 48 | 51 | 44 | 40 | 46 | 43 | 42 | 45 | 51 | 44 | 50 | 48 | 53 | 48 | 51 | 49 | 48 | 53 | 48 | 41 | 38 | 40 | 43 | 47 | 51 | 51 | 34 | 37 | 44 | 36 | 37 | 41 | 39 | 36 | 35 | 33 | 38 | 33 | 34 | 40 | 38 | 33 | 30 | 28 | 29 | 30 | 28 | 33 | 35 | 31 | 35 | 37 | 34 | 34 | 34 | 33 | 32 | 38 | 36 | 30 | 31 | 34 | 34 | 31 | 30 | 32 | 29 | 28 | 33 | 35 | 35 | 32 | 34 | 35 | 30 | 27 | 36 | 40 | 31 | 26 | 30 | 33 | 33 | 31 | 29 | 25 | 23 | 25 | 22 | 26 | 24 | 27 | 27 | 30 | 30 | 25 | 26 | 29 | 28 | 27 | 27 | 25 | 24 | 23 | 28 | 28 | 25 | 23 | 19 | 20 | 29 | 20 | 25 | 25 | 23 | 23 | 26 | 34 | 28 | 23 | 27 | 26 | 20 | 22 | 19 | 15 | 26 | 28 | 30 | 29 | 15 | 23 | 28 | 23 | 20 | 22 | 25 | 31 | 29 | 27 | 28 | 26 | 26 | 25 | 22 | 23 | 25 | 25 | 24 | 26 | 27 | 27 | 23 | 27 | 29 | 23 | 22 | 21 | 20 | 20 | 21 | 23 | 22 | 19 | 16 | 16 | 18 | 22 | 20 | 17 | 16 | 18 | 19 | 21 | 23 | 18 | 16 | 19 | 23 | 22 | 21 | 23 | 20 | 19 | 18 | 20 | 22 | 24 | 25 | 22 | 23 | 22 | 17 | 15 | 19 | 19 | 21 | 19 | 22 | 26 | 24 | 27 | 27 | 22 | 18 | 23 | 22 | 22 | 20 | 16 | 24 | 26 | 29 | 22 | 26 | 25 | 28 | 27 | 28 | 26 | 18 | 32 | 20 | 24 | 28 | 23 | 24 | 24 | 23 | 28 | 27 | 29 | 23 | 30 | 30 | 22 | 18 | 22 | 27 | 27 | 25 | 26 | 34 | 36 | 31 | 29 | 29 | 25 | 25 | 25 | 32 | 30 | 32 | 35 | 27 | 26 | 29 | 28 | 31 | 28 | 32 | 35 | 33 | 33 | 31 | 32 | 30 | 27 | 27 | 22 | 22 | 28 | 34 | 37 | 35 | 32 | 39 | 35 | 26 | 28 | 44 | 39 | 32 | 31 | 32 | 45 | 30 | 28 | 43 | 35 | 26 | 23 | 26 | 29 | 24 | 29 | 33 | 39 | 30 | 31 | 29 | 26 | 36 | 23 | 16 | 30 | 26 | 24 | 26 | 35 | 22 | 27 | 18 | 21 | 26 | 17 | 23 | 22 | 15 | 21 | 26 | 18 | 20 | 15 | 19 | 11 | 17 | 18 | 15 | 17 | 21 | 16 | 13 | 14 | 16 | 14 | 9 | 17 | 15 | 10 | 10 | 8 | 10 | 12 | 10 | 9 | 10 | 10 | 11 | 9 | 6 | 7 | 9 | 9 | 10 | 8 | 5 | 4 | 7 | 6 | 5 | 5 | 6 | 5 | 4 | 5 | 5 | 6 | 6 | 8 | 7 | 6 | 7 | 10 | 7 | 6 | 5 | 3 | 5 | 7 | 6 | 5 | 4 | 7 | 6 | 7 | 8 | 6 | 8 | 9 | 7 | 6 | 6 | 5 | 9 | 6 | 5 | 5 | 6 | 7 | 8 | 9 | 5 | 6 | 4 | 3 | 6 | 3 | 7 | 6 | 10 | 4 | 8 | 8 | 6 | 7 | 8 | 7 | 3 | 6 | 7 | 4 | 9 | 7 | 5 | 4 | 5 | 6 | 4 | 3 | 2 | 5 | 6 | 7 | 10 | 5 | 5 | 4 | 6 | 5 | 7 | 5 | 3 | 4 | 6 | 4 | 3 | 5 | 5 | 2 | 6 | 11 | 7 | 4 | 5 | 6 | 6 | 6 | 5 | 5 | 3 | 2 | 7 | 4 | 3 | 4 | 6 | 5 | 5 | 4 | 4 | 5 | 6 | 7 | 5 | 3 | 6 | 4 | 3 | 2 | 7 | 4 | 5 | 5 | 3 | 5 | 4 | 6 | 5 | 5 | 7 | 3 | 5 | 6 | 3 | 4 | 5 | 8 | 3 | 3 | 4 | 4 | 4 | 3 | 4 | 7 | 1 | 6 | 4 | 3 | 4 | 3 | 4 | 6 | 3 | 3 | 6 | 4 | 2 | 4 | 5 | 6 | 1 | 2 | 2 | 1 | 3 | 4 | 4 | 6 | 5 | 1 | 1 | 1 | 4 | 4 | 4 | 5 | 3 | 3 | 0 | 1 | 2 | 4 | 5 | 4 | 4 | 5 | 3 | 6 | 6 | 6 | 4 | 1 | 1 | 3 | 3 | 1 | 3 | 5 | 3 | 4 | 5 | 6 | 3 | 3 | 2 | 2 | 3 | 1 | 1 | 3 | 2 | 2 | 4 | 2 | 3 | 2 | 2 | 1 | 1 | 3 | 2 | 1 | 1 | 1 | 0 | 4 | 1 | 3 | 2 | 2 | 2 | 0 | 1 | 0 | 0 | 3 | 3 | 2 | 0 | 1 | 2 | 3 | 3 | 4 | 1 | 4 | 0 | 1 | 1 | 0 | 1 | 3 | 4 | 0 | 1 | 2 | 1 | 1 | 3 | 3 | 1 | 2 | 2 | 2 | 2 | 1 | 1 | 2 | 2 | 0 | 1 | 3 | 2 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 2 | 3 | 3 | 2 | 1 | 1 | 1 | 1 | 2 | 1 | 1 | 2 | 1 | 1 | 1 | 2 | 1 | 0 | 1 | 1 | 2 | 1 | 1 | 1 | 2 | 1 | 2 | 2 | 1 | 0 | 0 | 0 | 2 | 3 | 1 | 0 | 2 | 2 | 0 | 1 | 0 | 2 | 1 | 1 | 1 | 1 | 2 | 4 | 0 | 2 | 1 | 2 | 0 | 2 | 0 | 2 | 3 | 0 | 2 | 1 | 2 | 2 | 1 | 1 | 4 | 3 | 0 | 0 | 1 | 1 | 2 | 1 | 1 | 3 | 4 | 2 | 1 | 1 | 2 | 3 | 3 | 0 | 1 | 1 | 1 | 2 | 2 | 2 | 1 | 2 | 3 | 2 | 1 | 1 | 2 | 0 | 2 | 4 | 3 | 1 | 0 | 4 | 4 | 3 | 2 | 2 | 2 | 3 | 2 | 2 | 1 | 2 | 1 | 1 | 2 | 2 | 3 | 2 | 1 | 0 | 1 | 3 | 2 | 3 | 1 | 1 | 3 | 2 | 3 | 1 | 0 | 0 | 0 | 0 | 2 | 1 | 1 | 1 | 0 | 2 | 0 | 1 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 4 | 1326 | 91 | 108 | 137 | 151 | 143 | 166 | 167 | 169 | 197 | 230 | 265 | 290 | 316 | 341 | 379 | 424 | 474 | 526 | 582 | 692 | 797 | 853 | 948 | 1051 | 1139 | 1268 | 1377 | 1528 | 1667 | 1868 | 2092 | 2194 | 2228 | 2362 | 2473 | 2522 | 2596 | 2657 | 2794 | 2912 | 2967 | 2997 | 2974 | 2943 | 2939 | 2922 | 2966 | 2984 | 2975 | 2977 | 2963 | 3160 | 3041 | 3012 | 3077 | 3057 | 3053 | 3023 | 2988 | 3007 | 3041 | 2945 | 3336 | 3053 | 3065 | 3036 | 3034 | 3034 | 3009 | 3014 | 2978 | 3075 | 3104 | 3355 | 2987 | 3059 | 3095 | 3028 | 2985 | 3023 | 3094 | 3136 | 3062 | 3229 | 3199 | 3096 | 3100 | 3119 | 3080 | 3062 | 3062 | 3131 | 3056 | 3094 | 3170 | 3094 | 3106 | 3036 | 2951 | 2932 | 2926 | 2927 | 2942 | 2934 | 2964 | 3029 | 2906 | 2875 | 2867 | 2876 | 2865 | 2847 | 2833 | 2840 | 2885 | 2896 | 2736 | 2589 | 2559 | 2592 | 2639 | 2637 | 2633 | 2586 | 2609 | 2661 | 2660 | 2672 | 2694 | 2646 | 2641 | 2660 | 2661 | 2657 | 2645 | 2599 | 2610 | 2688 | 2688 | 2640 | 2619 | 2644 | 2737 | 2778 | 2702 | 2669 | 2678 | 2638 | 2623 | 2668 | 2724 | 2746 | 2798 | 2792 | 2791 | 2781 | 2842 | 2869 | 2896 | 2950 | 3017 | 3047 | 3083 | 3093 | 3122 | 3185 | 3232 | 3307 | 3351 | 3356 | 3394 | 3517 | 3478 | 3482 | 3582 | 3660 | 3712 | 3772 | 3785 | 3915 | 3942 | 3993 | 4024 | 4077 | 3992 | 4024 | 4083 | 4056 | 4026 | 3994 | 3918 | 3879 | 3914 | 3954 | 3916 | 3877 | 3866 | 3809 | 3769 | 3764 | 3704 | 3671 | 3664 | 3572 | 3531 | 3479 | 3446 | 3439 | 3397 | 3335 | 3376 | 3397 | 3243 | 3199 | 3233 | 3231 | 3177 | 3085 | 3088 | 3062 | 2993 | 2997 | 2985 | 2974 | 3025 | 2973 | 2898 | 2903 | 2873 | 2847 | 2832 | 2813 | 2823 | 2768 | 2715 | 2694 | 2627 | 2620 | 2636 | 2618 | 2614 | 2620 | 2610 | 2569 | 2528 | 2497 | 2486 | 2490 | 2464 | 2439 | 2414 | 2395 | 2416 | 2401 | 2367 | 2396 | 2377 | 2292 | 2274 | 2305 | 2294 | 2240 | 2241 | 2291 | 2290 | 2284 | 2132 | 2212 | 2146 | 2134 | 2095 | 2131 | 2221 | 2035 | 2145 | 2166 | 2009 | 1975 | 2006 | 1978 | 1964 | 2016 | 1960 | 1938 | 1909 | 1858 | 1935 | 1924 | 1936 | 1949 | 1923 | 1894 | 1919 | 1912 | 1878 | 1833 | 1808 | 1835 | 1842 | 1823 | 1820 | 1840 | 1875 | 1842 | 1798 | 1810 | 1771 | 1747 | 1786 | 1807 | 1766 | 1743 | 1712 | 1685 | 1690 | 1705 | 1719 | 1692 | 1665 | 1611 | 1566 | 1599 | 1614 | 1585 | 1599 | 1598 | 1581 | 1576 | 1552 | 1530 | 1482 | 1484 | 1465 | 1424 | 1429 | 1410 | 1401 | 1409 | 1401 | 1390 | 1371 | 1329 | 1312 | 1280 | 1290 | 1342 | 1324 | 1277 | 1259 | 1251 | 1253 | 1255 | 1237 | 1216 | 1245 | 1252 | 1207 | 1203 | 1155 | 1168 | 1210 | 1155 | 1149 | 1173 | 1150 | 1197 | 1161 | 1145 | 1132 | 1100 | 1167 | 1089 | 1127 | 1126 | 1169 | 1150 | 1119 | 1078 | 1128 | 1066 | 1064 | 1099 | 1077 | 1069 | 1060 | 1055 | 1015 | 1034 | 1064 | 1041 | 1061 | 1064 | 1071 | 1079 | 1071 | 1055 | 1066 | 1033 | 1025 | 1024 | 1014 | 1012 | 997 | 990 | 988 | 987 | 1001 | 1032 | 1065 | 1053 | 1009 | 982 | 960 | 974 | 998 | 968 | 960 | 978 | 992 | 1025 | 1030 | 992 | 978 | 981 | 990 | 1015 | 996 | 993 | 981 | 1015 | 1029 | 1032 | 1065 | 1088 | 1064 | 1042 | 1065 | 1095 | 1129 | 1179 | 1226 | 1251 | 1285 | 1299 | 1343 | 1381 | 1461 | 1521 | 1587 | 1653 | 1673 | 1753 | 1912 | 2030 | 2147 | 2306 | 2443 | 2637 | 2863 | 3049 | 3290 | 3525 | 3734 | 3981 | 4383 | 4718 | 5085 | 5521 | 5872 | 6265 | 6552 | 7171 | 7694 | 8053 | 8570 | 9234 | 9938 | 10512 | 11224 | 11970 | 12395 | 12981 | 13519 | 14397 | 15197 | 15682 | 16430 | 17025 | 17573 | 18127 | 18696 | 19070 | 19493 | 20061 | 20628 | 20868 | 20879 | 21202 | 21301 | 21380 | 21387 | 21206 | 21099 | 20967 | 20602 | 20156 | 19846 | 19464 | 18910 | 18264 | 17639 | 17053 | 16321 | 15592 | 14892 | 14098 | 13335 | 12641 | 11877 | 11073 | 10319 | 9663 | 9013 | 8191 | 7577 | 7007 | 6392 | 5919 | 5460 | 4961 | 4592 | 4129 | 3724 | 3416 | 3145 | 2960 | 2710 | 2508 | 2315 | 2178 | 2057 | 1921 | 1808 | 1731 | 1671 | 1638 | 1566 | 1484 | 1413 | 1431 | 1455 | 1415 | 1378 | 1341 | 1310 | 1300 | 1285 | 1258 | 1281 | 1309 | 1284 | 1320 | 1316 | 1310 | 1297 | 1290 | 1277 | 1297 | 1313 | 1218 | 1292 | 1225 | 1222 | 1267 | 1291 | 1266 | 1274 | 1223 | 1211 | 1217 | 1194 | 1251 | 1233 | 1245 | 1246 | 1218 | 1201 | 1206 | 1246 | 1233 | 1214 | 1184 | 1205 | 1208 | 1184 | 1162 | 1156 | 1158 | 1209 | 1204 | 1163 | 1142 | 1138 | 1141 | 1136 | 1145 | 1168 | 1169 | 1148 | 1133 | 1133 | 1132 | 1112 | 1099 | 1096 | 1092 | 1092 | 1109 | 1147 | 1141 | 1113 | 1075 | 1070 | 1086 | 1109 | 1121 | 1117 | 1116 | 1078 | 1087 | 1114 | 1123 | 1142 | 1116 | 1082 | 1046 | 1055 | 1068 | 1089 | 1101 | 1106 | 1118 | 1143 | 1133 | 1092 | 1100 | 1134 | 1140 | 1113 | 1098 | 1127 | 1155 | 1233 | 1177 | 1232 | 1216 | 1212 | 1204 | 1199 | 1196 | 1227 | 1247 | 1257 | 1295 | 1309 | 1302 | 1319 | 1327 | 1359 | 1302 | 1345 | 1395 | 1460 | 1414 | 1420 | 1423 | 1393 | 1421 | 1419 | 1443 | 1449 | 1458 | 1453 | 1459 | 1473 | 1462 | 1473 | 1452 | 1424 | 1476 | 1486 | 1502 | 1512 | 1419 | 1420 | 1466 | 1419 | 1419 | 1407 | 1390 | 1356 | 1339 | 1384 | 1397 | 1321 | 1284 | 1301 | 1283 | 1245 | 1233 | 1255 | 1243 | 1200 | 1178 | 1225 | 1224 | 1165 | 1164 | 1172 | 1159 | 1165 | 1158 | 1114 | 1086 | 1089 | 1102 | 1080 | 1054 | 1010 | 1001 | 983 | 1027 | 992 | 989 | 994 | 982 | 979 | 979 | 955 | 926 | 907 | 881 | 873 | 903 | 909 | 866 | 860 | 867 | 923 | 824 | 849 | 908 | 847 | 826 | 787 | 756 | 800 | 784 | 759 | 765 | 766 | 772 | 761 | 743 | 762 | 730 | 681 | 689 | 734 | 710 | 698 | 709 | 683 | 663 | 697 | 655 | 668 | 653 | 639 | 646 | 644 | 636 | 620 | 598 | 597 | 599 | 599 | 625 | 631 | 603 | 583 | 609 | 612 | 612 | 594 | 588 | 590 | 571 | 562 | 540 | 525 | 536 | 548 | 547 | 543 | 535 | 525 | 510 | 516 | 536 | 529 | 519 | 515 | 502 | 489 | 489 | 503 | 501 | 488 | 484 | 483 | 478 | 473 | 474 | 463 | 456 | 462 | 442 | 427 | 410 | 416 | 428 | 400 | 416 | 420 | 432 | 421 | 405 | 403 | 395 | 381 | 383 | 388 | 373 | 371 | 391 | 362 | 367 | 325 | 355 | 353 | 333 | 330 | 344 | 354 | 364 | 352 | 360 | 353 | 341 | 330 | 322 | 348 | 348 | 304 | 321 | 314 | 305 | 318 | 300 | 327 | 316 | 301 | 310 | 300 | 298 | 296 | 297 | 300 | 291 | 277 | 281 | 295 | 299 | 301 | 289 | 288 | 291 | 290 | 287 | 281 | 258 | 287 | 276 | 280 | 282 | 277 | 261 | 256 | 270 | 279 | 275 | 271 | 269 | 264 | 259 | 268 | 268 | 261 | 272 | 268 | 259 | 253 | 261 | 271 | 261 | 260 | 259 | 261 | 265 | 260 | 268 | 279 | 276 | 280 | 285 | 283 | 261 | 278 | 273 | 282 | 287 | 257 | 280 | 294 | 279 | 275 | 290 | 298 | 310 | 329 | 324 | 310 | 320 | 322 | 352 | 345 | 333 | 379 | 371 | 369 | 373 | 419 | 422 | 418 | 421 | 478 | 489 | 515 | 551 | 547 | 584 | 605 | 630 | 653 | 671 | 689 | 707 | 790 | 807 | 865 | 868 | 912 | 975 | 1043 | 1082 | 1109 | 1108 | 1157 | 1236 | 1294 | 1367 | 1422 | 1429 | 1479 | 1560 | 1610 | 1632 | 1735 | 1772 | 1803 | 1875 | 1879 | 1928 | 2000 | 2022 | 2047 | 2089 | 2141 | 2166 | 2184 | 2226 | 2270 | 2294 | 2278 | 2311 | 2354 | 2393 | 2417 | 2333 | 2321 | 2347 | 2340 | 2302 | 2262 | 2254 | 2255 | 2230 | 2165 | 2097 | 2081 | 2079 | 2004 | 1976 | 1913 | 1821 | 1783 | 1805 | 1719 | 1588 | 1532 | 1515 | 1457 | 1354 | 1289 | 1251 | 1189 | 1147 | 1118 | 1051 | 971 | 908 | 863 | 799 | 727 | 670 | 672 | 626 | 559 | 530 | 517 | 472 | 414 | 396 | 387 | 363 | 323 | 298 | 283 | 247 | 221 | 230 | 219 | 202 | 186 | 165 | 153 | 155 | 143 | 133 | 126 | 125 | 112 | 112 | 99 | 93 | 97 | 90 | 84 | 81 | 72 | 70 | 66 | 71 | 76 | 75 | 70 | 68 | 68 | 67 | 70 | 68 | 66 | 71 | 71 | 70 | 71 | 69 | 66 | 65 | 66 | 69 | 73 | 71 | 62 | 55 | 61 | 73 | 73 | 62 | 55 | 61 | 62 | 62 | 67 | 65 | 65 | 68 | 70 | 72 | 65 | 69 | 67 | 54 | 48 | 58 | 68 | 65 | 65 | 61 | 59 | 62 | 67 | 64 | 54 | 52 | 51 | 56 | 63 | 57 | 61 | 70 | 59 | 56 | 64 | 64 | 66 | 63 | 57 | 55 | 54 | 50 | 49 | 55 | 51 | 47 | 52 | 56 | 57 | 62 | 66 | 61 | 72 | 68 | 63 | 56 | 60 | 50 | 65 | 62 | 58 | 59 | 63 | 57 | 53 | 50 | 55 | 57 | 56 | 55 | 54 | 58 | 56 | 45 | 44 | 56 | 61 | 56 | 54 | 52 | 58 | 60 | 52 | 46 | 47 | 48 | 43 | 40 | 38 | 42 | 52 | 56 | 52 | 45 | 41 | 53 | 61 | 56 | 51 | 46 | 46 | 46 | 42 | 40 | 43 | 44 | 45 | 51 | 47 | 42 | 42 | 39 | 42 | 42 | 42 | 41 | 38 | 43 | 39 | 35 | 42 | 45 | 45 | 43 | 40 | 42 | 41 | 34 | 31 | 34 | 39 | 46 | 40 | 42 | 35 | 36 | 31 | 31 | 35 | 34 | 34 | 37 | 34 | 31 | 32 | 32 | 31 | 36 | 40 | 37 | 38 | 36 | 41 | 43 | 25 | 33 | 28 | 35 | 38 | 34 | 30 | 29 | 33 | 32 | 26 | 29 | 34 | 36 | 41 | 43 | 38 | 28 | 28 | 36 | 40 | 32 | 29 | 35 | 32 | 31 | 33 | 36 | 30 | 24 | 24 | 26 | 26 | 23 | 22 | 24 | 22 | 23 | 26 | 28 | 31 | 28 | 27 | 31 | 29 | 24 | 24 | 27 | 30 | 30 | 26 | 26 | 30 | 28 | 24 | 25 | 22 | 24 | 24 | 28 | 26 | 26 | 30 | 23 | 22 | 25 | 27 | 27 | 25 | 23 | 22 | 25 | 28 | 27 | 21 | 21 | 19 | 28 | 23 | 28 | 17 | 25 | 22 | 19 | 24 | 27 | 24 | 25 | 24 | 26 | 27 | 23 | 18 | 23 | 24 | 17 | 21 | 20 | 26 | 24 | 22 | 16 | 24 | 26 | 26 | 22 | 24 | 26 | 21 | 18 | 19 | 18 | 20 | 20 | 18 | 19 | 22 | 20 | 18 | 17 | 15 | 22 | 20 | 20 | 23 | 23 | 25 | 21 | 20 | 19 | 18 | 21 | 21 | 21 | 21 | 21 | 24 | 25 | 23 | 19 | 25 | 28 | 25 | 22 | 19 | 22 | 23 | 21 | 21 | 20 | 20 | 22 | 22 | 22 | 20 | 21 | 24 | 27 | 25 | 22 | 26 | 24 | 26 | 27 | 22 | 26 | 16 | 25 | 18 | 18 | 22 | 22 | 25 | 29 | 30 | 23 | 25 | 34 | 24 | 23 | 29 | 24 | 24 | 33 | 29 | 25 | 25 | 23 | 26 | 28 | 26 | 27 | 24 | 23 | 20 | 21 | 19 | 23 | 32 | 32 | 33 | 30 | 29 | 34 | 31 | 30 | 34 | 29 | 29 | 38 | 31 | 27 | 34 | 29 | 30 | 33 | 23 | 24 | 24 | 22 | 32 | 33 | 27 | 32 | 23 | 32 | 40 | 32 | 37 | 41 | 36 | 31 | 25 | 31 | 31 | 32 | 34 | 30 | 34 | 38 | 33 | 28 | 33 | 33 | 34 | 31 | 26 | 23 | 24 | 24 | 21 | 23 | 25 | 23 | 27 | 23 | 21 | 25 | 22 | 25 | 29 | 30 | 30 | 26 | 23 | 22 | 26 | 23 | 27 | 19 | 20 | 27 | 17 | 16 | 17 | 20 | 13 | 24 | 17 | 12 | 12 | 16 | 23 | 13 | 14 | 10 | 9 | 7 | 10 | 11 | 9 | 16 | 9 | 6 | 10 | 8 | 7 | 6 | 4 | 6 | 9 | 4 | 5 | 7 | 7 | 6 | 6 | 6 | 7 | 5 | 13 | 11 | 10 | 7 | 10 | 7 | 6 | 7 | 8 | 8 | 6 | 6 | 6 | 5 | 7 | 4 | 4 | 6 | 7 | 9 | 7 | 6 | 6 | 4 | 7 | 5 | 5 | 5 | 3 | 4 | 7 | 6 | 6 | 5 | 5 | 5 | 6 | 6 | 4 | 7 | 6 | 5 | 4 | 5 | 6 | 5 | 7 | 5 | 3 | 7 | 4 | 6 | 5 | 3 | 3 | 5 | 7 | 9 | 9 | 4 | 4 | 6 | 8 | 6 | 3 | 3 | 6 | 3 | 2 | 5 | 6 | 5 | 9 | 3 | 2 | 0 | 6 | 2 | 4 | 2 | 2 | 4 | 3 | 6 | 3 | 7 | 9 | 4 | 7 | 4 | 5 | 5 | 3 | 3 | 7 | 8 | 2 | 6 | 4 | 6 | 6 | 3 | 5 | 4 | 6 | 8 | 3 | 3 | 4 | 5 | 6 | 6 | 3 | 4 | 6 | 5 | 3 | 4 | 4 | 3 | 4 | 4 | 3 | 4 | 4 | 5 | 6 | 5 | 3 | 3 | 5 | 4 | 5 | 4 | 2 | 4 | 4 | 4 | 2 | 2 | 3 | 3 | 3 | 5 | 2 | 2 | 3 | 1 | 1 | 3 | 3 | 6 | 6 | 4 | 5 | 4 | 1 | 6 | 4 | 1 | 2 | 2 | 3 | 1 | 2 | 3 | 3 | 4 | 1 | 2 | 2 | 1 | 4 | 1 | 1 | 5 | 3 | 1 | 2 | 1 | 2 | 2 | 2 | 4 | 1 | 2 | 1 | 4 | 1 | 2 | 1 | 4 | 1 | 0 | 1 | 5 | 4 | 2 | 2 | 1 | 3 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 3 | 5 | 2 | 2 | 3 | 4 | 4 | 3 | 3 | 1 | 1 | 3 | 3 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 3 | 4 | 2 | 2 | 2 | 1 | 2 | 2 | 2 | 1 | 1 | 3 | 3 | 3 | 2 | 2 | 4 | 3 | 1 | 3 | 1 | 3 | 1 | 2 | 3 | 3 | 2 | 1 | 2 | 1 | 2 | 4 | 1 | 3 | 3 | 0 | 1 | 2 | 0 | 1 | 2 | 1 | 3 | 0 | 1 | 1 | 2 | 2 | 0 | 2 | 3 | 1 | 2 | 2 | 1 | 2 | 3 | 0 | 2 | 2 | 0 | 1 | 2 | 1 | 1 | 3 | 2 | 0 | 1 | 3 | 2 | 1 | 1 | 0 | 0 | 1 | 2 | 1 | 0 | 1 | 2 | 1 | 1 | 0 | 1 | 2 | 2 | 2 | 2 | 2 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 4 | 3 | 1 | 2 | 2 | 1 | 2 | 1 | 0 | 1 | 0 | 1 | 1 | 3 | 2 | 2 | 3 | 1 | 1 | 1 | 1 | 2 | 3 | 2 | 1 | 1 | 1 | 1 | 2 | 1 | 1 | 3 | 2 | 1 | 1 | 3 | 4 | 2 | 1 | 0 | 1 | 2 | 1 | 1 | 2 | 2 | 3 | 2 | 4 | 4 | 2 | 1 | 0 | 0 | 1 | 2 | 1 | 1 | 4 | 1 | 2 | 1 | 0 | 2 | 1 | 1 | 2 | 2 | 1 | 1 | 3 | 5 | 3 | 0 | 1 | 2 | 2 | 2 | 1 | 2 | 2 | 2 | 2 | 1 | 3 | 2 | 3 | 1 | 0 | 2 | 3 | 1 | 1 | 2 | 1 | 2 | 3 | 1 | 1 | 3 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 5 | 1351 | 102 | 113 | 123 | 133 | 161 | 156 | 165 | 176 | 199 | 225 | 255 | 287 | 319 | 342 | 377 | 434 | 504 | 548 | 605 | 698 | 802 | 875 | 963 | 1093 | 1217 | 1332 | 1422 | 1526 | 1673 | 1900 | 2076 | 2146 | 2290 | 2422 | 2536 | 2625 | 2690 | 2724 | 2755 | 2842 | 2973 | 2988 | 2940 | 2919 | 2927 | 2923 | 2918 | 2986 | 2989 | 3000 | 3028 | 3218 | 3023 | 2944 | 3001 | 2966 | 2984 | 2991 | 3006 | 3002 | 3056 | 2953 | 3176 | 3047 | 3075 | 2938 | 2982 | 3039 | 3038 | 3071 | 2998 | 3017 | 3039 | 3334 | 3104 | 3052 | 3073 | 3043 | 3088 | 3055 | 3109 | 3070 | 3075 | 3193 | 3296 | 3152 | 3144 | 3124 | 3097 | 3095 | 3078 | 3035 | 3030 | 3059 | 3201 | 3052 | 3036 | 2965 | 2977 | 2944 | 2892 | 2929 | 2941 | 2920 | 2905 | 2973 | 2911 | 2899 | 2855 | 2858 | 2864 | 2833 | 2851 | 2865 | 2877 | 2905 | 2758 | 2587 | 2594 | 2577 | 2579 | 2612 | 2642 | 2679 | 2645 | 2600 | 2592 | 2647 | 2672 | 2682 | 2642 | 2608 | 2610 | 2590 | 2614 | 2653 | 2643 | 2643 | 2667 | 2689 | 2651 | 2659 | 2709 | 2685 | 2699 | 2730 | 2711 | 2723 | 2797 | 2783 | 2761 | 2758 | 2740 | 2759 | 2843 | 2877 | 2912 | 2920 | 2918 | 2893 | 3012 | 3029 | 3055 | 3128 | 3171 | 3181 | 3204 | 3372 | 3394 | 3408 | 3449 | 3529 | 3608 | 3543 | 3726 | 3704 | 3777 | 3782 | 3778 | 3825 | 3947 | 4012 | 3952 | 3995 | 4079 | 4036 | 4050 | 4042 | 4041 | 4006 | 3968 | 3990 | 3975 | 3911 | 3870 | 3847 | 3819 | 3794 | 3768 | 3704 | 3614 | 3591 | 3628 | 3618 | 3587 | 3525 | 3467 | 3468 | 3434 | 3433 | 3440 | 3337 | 3289 | 3273 | 3241 | 3198 | 3127 | 3127 | 3125 | 3112 | 3079 | 3025 | 3038 | 3026 | 3031 | 3026 | 2988 | 2857 | 2807 | 2835 | 2831 | 2804 | 2728 | 2703 | 2701 | 2700 | 2703 | 2667 | 2646 | 2651 | 2644 | 2601 | 2575 | 2552 | 2534 | 2513 | 2460 | 2457 | 2493 | 2454 | 2396 | 2441 | 2440 | 2365 | 2338 | 2311 | 2347 | 2328 | 2328 | 2276 | 2236 | 2319 | 2213 | 2192 | 2171 | 2165 | 2147 | 2218 | 2174 | 2126 | 2153 | 2082 | 2128 | 2043 | 2093 | 2085 | 2061 | 2092 | 2123 | 2045 | 2079 | 2060 | 2019 | 1986 | 2007 | 1986 | 1968 | 1948 | 1960 | 1971 | 1970 | 1922 | 1869 | 1865 | 1846 | 1854 | 1902 | 1884 | 1864 | 1877 | 1836 | 1815 | 1845 | 1827 | 1821 | 1836 | 1816 | 1793 | 1764 | 1764 | 1792 | 1791 | 1752 | 1740 | 1708 | 1712 | 1744 | 1738 | 1707 | 1636 | 1562 | 1588 | 1664 | 1612 | 1545 | 1543 | 1552 | 1552 | 1548 | 1547 | 1516 | 1481 | 1484 | 1469 | 1431 | 1412 | 1371 | 1381 | 1414 | 1402 | 1357 | 1342 | 1312 | 1285 | 1321 | 1325 | 1286 | 1258 | 1241 | 1264 | 1280 | 1269 | 1231 | 1173 | 1202 | 1198 | 1178 | 1170 | 1177 | 1206 | 1199 | 1194 | 1203 | 1178 | 1169 | 1194 | 1097 | 1131 | 1127 | 1152 | 1117 | 1172 | 1055 | 1112 | 1104 | 1145 | 1145 | 1117 | 1077 | 1074 | 1066 | 1037 | 1091 | 1059 | 1054 | 1090 | 1024 | 1016 | 1039 | 1057 | 1034 | 1034 | 1056 | 1076 | 1061 | 1039 | 1010 | 1020 | 1031 | 1000 | 990 | 1009 | 1021 | 1042 | 1038 | 1001 | 976 | 972 | 990 | 1003 | 1003 | 996 | 988 | 991 | 995 | 987 | 968 | 988 | 1022 | 1032 | 1028 | 984 | 962 | 988 | 1006 | 1006 | 1054 | 1028 | 996 | 1036 | 1037 | 1041 | 1027 | 1030 | 1072 | 1119 | 1123 | 1127 | 1139 | 1169 | 1198 | 1220 | 1222 | 1248 | 1310 | 1368 | 1418 | 1491 | 1569 | 1647 | 1712 | 1809 | 1944 | 2070 | 2198 | 2376 | 2576 | 2698 | 2837 | 3093 | 3310 | 3508 | 3819 | 4109 | 4439 | 4668 | 5062 | 5439 | 5950 | 6451 | 6796 | 7314 | 7950 | 8381 | 8975 | 9524 | 10180 | 10932 | 11577 | 12207 | 12722 | 13329 | 13947 | 14847 | 15279 | 15794 | 16531 | 17134 | 17822 | 18340 | 19022 | 19515 | 19917 | 20246 | 20668 | 20932 | 21027 | 21201 | 21409 | 21462 | 21364 | 21227 | 21077 | 20818 | 20458 | 20022 | 19545 | 19017 | 18559 | 18123 | 17525 | 16883 | 16096 | 15230 | 14458 | 13754 | 13092 | 12384 | 11567 | 10781 | 10097 | 9316 | 8596 | 7977 | 7401 | 6839 | 6239 | 5633 | 5253 | 4892 | 4497 | 4102 | 3762 | 3351 | 3095 | 2905 | 2645 | 2436 | 2295 | 2157 | 2035 | 1915 | 1797 | 1720 | 1662 | 1587 | 1524 | 1464 | 1421 | 1419 | 1367 | 1339 | 1328 | 1341 | 1333 | 1300 | 1283 | 1282 | 1285 | 1285 | 1335 | 1332 | 1307 | 1270 | 1285 | 1296 | 1285 | 1272 | 1279 | 1241 | 1209 | 1231 | 1239 | 1197 | 1262 | 1208 | 1252 | 1225 | 1257 | 1236 | 1280 | 1281 | 1216 | 1202 | 1161 | 1220 | 1215 | 1185 | 1170 | 1207 | 1212 | 1159 | 1147 | 1179 | 1182 | 1125 | 1174 | 1200 | 1180 | 1160 | 1171 | 1153 | 1107 | 1091 | 1112 | 1141 | 1150 | 1128 | 1105 | 1112 | 1126 | 1144 | 1149 | 1103 | 1079 | 1114 | 1117 | 1082 | 1086 | 1101 | 1106 | 1073 | 1071 | 1080 | 1085 | 1112 | 1102 | 1080 | 1102 | 1112 | 1139 | 1167 | 1120 | 1074 | 1041 | 1065 | 1092 | 1099 | 1096 | 1076 | 1064 | 1067 | 1096 | 1117 | 1105 | 1110 | 1108 | 1121 | 1132 | 1125 | 1136 | 1168 | 1231 | 1201 | 1161 | 1204 | 1223 | 1161 | 1235 | 1242 | 1185 | 1199 | 1242 | 1253 | 1256 | 1270 | 1312 | 1354 | 1311 | 1305 | 1281 | 1315 | 1457 | 1455 | 1408 | 1389 | 1371 | 1357 | 1383 | 1414 | 1404 | 1427 | 1433 | 1396 | 1378 | 1394 | 1416 | 1433 | 1431 | 1386 | 1321 | 1385 | 1430 | 1392 | 1402 | 1420 | 1374 | 1402 | 1384 | 1371 | 1359 | 1325 | 1308 | 1306 | 1307 | 1320 | 1311 | 1278 | 1261 | 1275 | 1271 | 1258 | 1258 | 1202 | 1194 | 1200 | 1160 | 1166 | 1160 | 1111 | 1073 | 1068 | 1080 | 1087 | 1072 | 1054 | 1008 | 972 | 951 | 984 | 1022 | 1002 | 1015 | 1009 | 964 | 920 | 924 | 948 | 914 | 896 | 904 | 903 | 907 | 907 | 883 | 866 | 855 | 840 | 854 | 843 | 781 | 803 | 800 | 780 | 772 | 786 | 815 | 793 | 783 | 759 | 732 | 735 | 739 | 701 | 686 | 725 | 693 | 691 | 705 | 686 | 728 | 746 | 699 | 656 | 631 | 646 | 681 | 666 | 676 | 654 | 634 | 633 | 622 | 630 | 636 | 616 | 633 | 591 | 581 | 593 | 631 | 618 | 612 | 630 | 611 | 581 | 561 | 526 | 557 | 580 | 561 | 539 | 539 | 531 | 520 | 535 | 532 | 517 | 522 | 520 | 533 | 553 | 512 | 476 | 487 | 491 | 467 | 457 | 476 | 483 | 474 | 465 | 458 | 452 | 431 | 429 | 441 | 447 | 413 | 412 | 410 | 412 | 420 | 410 | 396 | 381 | 392 | 394 | 405 | 418 | 400 | 385 | 392 | 382 | 365 | 364 | 374 | 350 | 365 | 376 | 361 | 365 | 335 | 350 | 358 | 326 | 333 | 368 | 339 | 302 | 309 | 318 | 313 | 309 | 320 | 311 | 315 | 329 | 314 | 316 | 313 | 323 | 294 | 318 | 301 | 293 | 298 | 290 | 295 | 300 | 301 | 298 | 276 | 276 | 307 | 290 | 288 | 288 | 303 | 295 | 272 | 261 | 254 | 268 | 269 | 278 | 283 | 280 | 268 | 260 | 270 | 288 | 293 | 295 | 285 | 255 | 248 | 249 | 262 | 275 | 263 | 267 | 266 | 256 | 268 | 267 | 252 | 246 | 258 | 268 | 265 | 258 | 264 | 280 | 295 | 276 | 266 | 272 | 270 | 267 | 256 | 274 | 312 | 324 | 300 | 270 | 257 | 275 | 293 | 299 | 301 | 299 | 313 | 308 | 340 | 350 | 318 | 354 | 369 | 385 | 375 | 373 | 420 | 420 | 443 | 466 | 460 | 472 | 510 | 517 | 513 | 545 | 567 | 637 | 673 | 667 | 712 | 742 | 825 | 869 | 835 | 875 | 915 | 969 | 1012 | 1054 | 1086 | 1181 | 1271 | 1299 | 1302 | 1347 | 1420 | 1441 | 1502 | 1539 | 1594 | 1646 | 1763 | 1794 | 1796 | 1874 | 1874 | 1959 | 2040 | 2095 | 2140 | 2146 | 2139 | 2172 | 2221 | 2235 | 2220 | 2228 | 2299 | 2334 | 2326 | 2395 | 2368 | 2317 | 2324 | 2275 | 2267 | 2254 | 2238 | 2235 | 2222 | 2179 | 2142 | 2108 | 2042 | 1969 | 1913 | 1892 | 1848 | 1802 | 1715 | 1736 | 1613 | 1528 | 1506 | 1422 | 1417 | 1340 | 1275 | 1203 | 1127 | 1074 | 1004 | 956 | 927 | 848 | 823 | 782 | 717 | 701 | 639 | 602 | 588 | 522 | 463 | 441 | 422 | 382 | 339 | 312 | 292 | 283 | 269 | 251 | 243 | 207 | 196 | 179 | 171 | 149 | 159 | 154 | 111 | 122 | 126 | 118 | 121 | 119 | 96 | 91 | 96 | 96 | 88 | 79 | 74 | 72 | 74 | 87 | 83 | 76 | 80 | 69 | 69 | 65 | 64 | 61 | 65 | 74 | 78 | 67 | 60 | 63 | 64 | 63 | 68 | 69 | 61 | 58 | 60 | 60 | 72 | 64 | 58 | 56 | 57 | 64 | 64 | 62 | 59 | 61 | 63 | 63 | 65 | 73 | 67 | 64 | 63 | 59 | 64 | 66 | 58 | 57 | 55 | 60 | 54 | 59 | 62 | 61 | 57 | 58 | 60 | 54 | 55 | 62 | 62 | 61 | 51 | 56 | 60 | 63 | 59 | 50 | 44 | 60 | 62 | 55 | 55 | 58 | 59 | 60 | 53 | 51 | 57 | 60 | 54 | 50 | 59 | 57 | 47 | 58 | 58 | 58 | 53 | 66 | 62 | 61 | 64 | 71 | 61 | 53 | 53 | 55 | 54 | 51 | 56 | 60 | 56 | 48 | 54 | 55 | 60 | 53 | 47 | 44 | 53 | 51 | 47 | 49 | 57 | 56 | 58 | 64 | 61 | 54 | 56 | 57 | 55 | 47 | 43 | 49 | 49 | 47 | 46 | 52 | 46 | 36 | 39 | 47 | 52 | 49 | 43 | 44 | 43 | 39 | 44 | 48 | 44 | 52 | 51 | 51 | 50 | 48 | 55 | 53 | 50 | 44 | 36 | 35 | 37 | 40 | 44 | 45 | 39 | 36 | 40 | 37 | 31 | 38 | 39 | 40 | 32 | 38 | 43 | 41 | 34 | 40 | 33 | 29 | 33 | 36 | 34 | 30 | 31 | 37 | 37 | 36 | 39 | 32 | 32 | 36 | 32 | 35 | 36 | 31 | 36 | 37 | 28 | 28 | 30 | 31 | 31 | 26 | 24 | 21 | 25 | 27 | 28 | 27 | 28 | 26 | 31 | 34 | 34 | 32 | 24 | 25 | 30 | 33 | 32 | 28 | 26 | 26 | 27 | 28 | 26 | 27 | 26 | 20 | 22 | 29 | 26 | 25 | 28 | 28 | 26 | 27 | 30 | 28 | 24 | 24 | 24 | 27 | 30 | 26 | 24 | 24 | 26 | 29 | 24 | 17 | 20 | 24 | 28 | 28 | 24 | 22 | 21 | 23 | 21 | 20 | 26 | 28 | 20 | 24 | 29 | 25 | 21 | 20 | 23 | 32 | 23 | 24 | 24 | 21 | 30 | 30 | 25 | 22 | 23 | 28 | 30 | 20 | 23 | 20 | 20 | 18 | 9 | 33 | 20 | 16 | 24 | 26 | 22 | 21 | 20 | 14 | 16 | 22 | 28 | 27 | 22 | 23 | 19 | 17 | 16 | 21 | 23 | 24 | 18 | 22 | 18 | 20 | 22 | 24 | 28 | 28 | 25 | 20 | 16 | 15 | 17 | 18 | 16 | 19 | 21 | 23 | 27 | 23 | 15 | 19 | 21 | 19 | 18 | 18 | 18 | 23 | 25 | 24 | 24 | 21 | 18 | 20 | 25 | 25 | 17 | 18 | 21 | 21 | 22 | 23 | 24 | 25 | 22 | 18 | 21 | 23 | 24 | 24 | 26 | 29 | 30 | 31 | 24 | 21 | 28 | 30 | 15 | 22 | 31 | 34 | 27 | 26 | 33 | 19 | 22 | 22 | 22 | 25 | 24 | 30 | 16 | 25 | 22 | 30 | 30 | 22 | 24 | 22 | 27 | 28 | 26 | 26 | 27 | 28 | 35 | 29 | 32 | 32 | 21 | 28 | 35 | 28 | 34 | 33 | 36 | 37 | 39 | 37 | 35 | 27 | 37 | 40 | 28 | 34 | 38 | 32 | 26 | 26 | 33 | 37 | 33 | 31 | 32 | 33 | 31 | 30 | 30 | 30 | 30 | 27 | 30 | 33 | 33 | 29 | 27 | 26 | 27 | 27 | 32 | 30 | 26 | 28 | 31 | 29 | 27 | 25 | 27 | 24 | 23 | 24 | 19 | 22 | 19 | 15 | 13 | 13 | 20 | 22 | 18 | 17 | 13 | 15 | 18 | 20 | 10 | 15 | 22 | 17 | 8 | 12 | 12 | 15 | 16 | 18 | 15 | 10 | 11 | 8 | 17 | 9 | 11 | 9 | 5 | 7 | 5 | 14 | 8 | 10 | 6 | 9 | 11 | 3 | 4 | 5 | 7 | 12 | 8 | 6 | 10 | 7 | 8 | 8 | 8 | 6 | 5 | 8 | 6 | 8 | 8 | 6 | 4 | 6 | 7 | 8 | 10 | 7 | 5 | 5 | 5 | 3 | 4 | 7 | 8 | 5 | 6 | 7 | 6 | 8 | 6 | 4 | 6 | 7 | 6 | 4 | 4 | 6 | 6 | 7 | 7 | 7 | 8 | 8 | 7 | 5 | 6 | 8 | 8 | 5 | 3 | 3 | 6 | 6 | 7 | 4 | 4 | 5 | 5 | 5 | 6 | 7 | 6 | 3 | 3 | 5 | 3 | 4 | 5 | 6 | 6 | 7 | 10 | 5 | 2 | 5 | 7 | 7 | 6 | 6 | 6 | 9 | 3 | 6 | 3 | 5 | 8 | 1 | 4 | 5 | 2 | 4 | 2 | 6 | 5 | 0 | 2 | 5 | 4 | 7 | 7 | 8 | 5 | 3 | 2 | 1 | 3 | 5 | 5 | 3 | 3 | 4 | 6 | 5 | 1 | 2 | 1 | 4 | 6 | 5 | 5 | 4 | 3 | 2 | 4 | 8 | 5 | 1 | 2 | 4 | 3 | 3 | 3 | 2 | 1 | 2 | 5 | 4 | 3 | 3 | 2 | 5 | 6 | 5 | 3 | 3 | 3 | 2 | 2 | 3 | 2 | 2 | 2 | 5 | 4 | 2 | 1 | 2 | 2 | 2 | 3 | 1 | 3 | 2 | 4 | 1 | 1 | 2 | 4 | 1 | 0 | 4 | 3 | 3 | 4 | 0 | 3 | 1 | 4 | 0 | 4 | 6 | 6 | 4 | 2 | 4 | 2 | 3 | 2 | 3 | 7 | 0 | 2 | 1 | 3 | 4 | 1 | 3 | 2 | 3 | 1 | 2 | 4 | 2 | 4 | 2 | 2 | 2 | 3 | 3 | 2 | 1 | 1 | 4 | 1 | 2 | 3 | 4 | 2 | 2 | 1 | 0 | 0 | 1 | 2 | 2 | 0 | 2 | 3 | 3 | 2 | 1 | 2 | 4 | 3 | 1 | 2 | 2 | 1 | 1 | 1 | 1 | 0 | 1 | 2 | 3 | 2 | 1 | 2 | 2 | 2 | 3 | 3 | 2 | 1 | 1 | 2 | 2 | 2 | 1 | 1 | 1 | 3 | 3 | 5 | 2 | 2 | 1 | 2 | 0 | 1 | 2 | 0 | 5 | 3 | 2 | 3 | 0 | 3 | 0 | 3 | 2 | 3 | 1 | 3 | 1 | 4 | 0 | 2 | 2 | 0 | 2 | 5 | 2 | 3 | 1 | 1 | 2 | 4 | 0 | 2 | 1 | 1 | 3 | 2 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 2 | 2 | 1 | 1 | 2 | 1 | 0 | 1 | 2 | 2 | 1 | 0 | 1 | 2 | 3 | 3 | 1 | 0 | 0 | 2 | 3 | 3 | 2 | 2 | 1 | 2 | 3 | 2 | 1 | 1 | 2 | 1 | 1 | 2 | 2 | 2 | 2 | 2 | 3 | 3 | 2 | 2 | 2 | 0 | 1 | 3 | 3 | 1 | 0 | 2 | 1 | 1 | 1 | 0 | 2 | 0 | 2 | 3 | 1 | 4 | 2 | 0 | 2 | 0 | 2 | 0 | 2 | 1 | 1 | 1 | 3 | 4 | 0 | 3 | 2 | 3 | 0 | 1 | 4 | 2 | 1 | 1 | 0 | 3 | 3 | 4 | 0 | 2 | 2 | 2 | 0 | 2 | 4 | 4 | 4 | 3 | 1 | 1 | 1 | 1 | 3 | 2 | 1 | 2 | 2 | 2 | 1 | 1 | 3 | 3 | 2 | 2 | 3 | 1 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 6 | 237 | 14 | 15 | 22 | 20 | 29 | 31 | 18 | 36 | 26 | 28 | 38 | 54 | 68 | 92 | 117 | 131 | 140 | 175 | 201 | 230 | 327 | 406 | 445 | 499 | 589 | 682 | 756 | 848 | 915 | 992 | 1011 | 1117 | 1130 | 1116 | 1117 | 1150 | 1176 | 1211 | 1223 | 1155 | 1210 | 1207 | 1227 | 1245 | 1186 | 1180 | 1199 | 1214 | 1247 | 1244 | 1194 | 1187 | 1185 | 1269 | 1178 | 1175 | 1169 | 1154 | 1176 | 1175 | 1192 | 1186 | 1157 | 1241 | 1306 | 1233 | 1236 | 1212 | 1224 | 1222 | 1206 | 1204 | 1221 | 1208 | 1207 | 1307 | 1226 | 1188 | 1243 | 1251 | 1251 | 1260 | 1247 | 1243 | 1219 | 1259 | 1290 | 1187 | 1167 | 1165 | 1187 | 1167 | 1172 | 1186 | 1172 | 1158 | 1223 | 1169 | 1136 | 1137 | 1163 | 1141 | 1112 | 1129 | 1130 | 1163 | 1161 | 1201 | 1158 | 1165 | 1155 | 1088 | 1067 | 1106 | 1149 | 1191 | 1157 | 1184 | 1154 | 1068 | 1065 | 1047 | 1052 | 1054 | 1018 | 1061 | 1030 | 1035 | 1040 | 1036 | 1051 | 1053 | 1071 | 1073 | 1067 | 1094 | 1060 | 1045 | 1055 | 1056 | 1080 | 1066 | 1045 | 1055 | 1080 | 1075 | 1052 | 1051 | 1057 | 1078 | 1089 | 1086 | 1096 | 1089 | 1096 | 1104 | 1110 | 1125 | 1124 | 1157 | 1169 | 1190 | 1203 | 1228 | 1250 | 1247 | 1289 | 1366 | 1394 | 1415 | 1460 | 1454 | 1451 | 1462 | 1491 | 1557 | 1631 | 1634 | 1624 | 1624 | 1625 | 1620 | 1599 | 1596 | 1625 | 1587 | 1557 | 1573 | 1567 | 1536 | 1502 | 1494 | 1500 | 1490 | 1478 | 1448 | 1417 | 1407 | 1409 | 1417 | 1380 | 1341 | 1349 | 1342 | 1311 | 1291 | 1298 | 1292 | 1257 | 1220 | 1232 | 1227 | 1219 | 1192 | 1178 | 1178 | 1138 | 1122 | 1150 | 1144 | 1135 | 1114 | 1093 | 1109 | 1154 | 1093 | 1085 | 1070 | 1052 | 1083 | 1061 | 1017 | 999 | 1031 | 1035 | 1010 | 1012 | 1002 | 986 | 960 | 945 | 932 | 938 | 944 | 944 | 943 | 923 | 911 | 915 | 914 | 931 | 938 | 903 | 894 | 904 | 896 | 877 | 883 | 871 | 844 | 823 | 816 | 828 | 805 | 793 | 820 | 791 | 770 | 788 | 763 | 773 | 779 | 766 | 767 | 751 | 743 | 710 | 736 | 738 | 703 | 714 | 682 | 707 | 713 | 730 | 725 | 714 | 696 | 680 | 696 | 716 | 714 | 684 | 689 | 708 | 693 | 671 | 674 | 675 | 682 | 705 | 679 | 654 | 673 | 668 | 646 | 616 | 621 | 637 | 642 | 611 | 578 | 586 | 599 | 622 | 591 | 599 | 598 | 583 | 606 | 587 | 581 | 585 | 580 | 591 | 552 | 534 | 539 | 544 | 502 | 520 | 510 | 528 | 524 | 489 | 491 | 478 | 501 | 478 | 487 | 492 | 512 | 494 | 474 | 470 | 453 | 461 | 470 | 473 | 473 | 444 | 428 | 453 | 465 | 449 | 436 | 430 | 432 | 448 | 460 | 444 | 433 | 437 | 439 | 427 | 424 | 424 | 401 | 403 | 410 | 398 | 418 | 399 | 385 | 387 | 383 | 356 | 409 | 410 | 424 | 396 | 398 | 416 | 400 | 394 | 382 | 381 | 390 | 387 | 384 | 368 | 379 | 385 | 385 | 396 | 376 | 377 | 375 | 348 | 353 | 388 | 373 | 347 | 366 | 376 | 360 | 365 | 382 | 384 | 382 | 363 | 333 | 316 | 328 | 353 | 345 | 354 | 352 | 361 | 376 | 342 | 332 | 334 | 351 | 371 | 363 | 347 | 353 | 355 | 352 | 353 | 368 | 366 | 387 | 388 | 373 | 367 | 371 | 375 | 390 | 405 | 394 | 385 | 424 | 451 | 454 | 466 | 497 | 518 | 568 | 612 | 635 | 680 | 735 | 792 | 860 | 922 | 1006 | 1110 | 1206 | 1315 | 1467 | 1667 | 1864 | 2011 | 2151 | 2327 | 2544 | 2756 | 2989 | 3265 | 3536 | 3853 | 4083 | 4336 | 4700 | 4996 | 5334 | 5744 | 6064 | 6410 | 6866 | 7194 | 7472 | 7795 | 7966 | 8226 | 8581 | 8958 | 9173 | 9372 | 9496 | 9503 | 9583 | 9710 | 9672 | 9593 | 9480 | 9406 | 9339 | 9219 | 9056 | 8823 | 8570 | 8280 | 7941 | 7616 | 7312 | 6925 | 6500 | 6172 | 5879 | 5508 | 5080 | 4702 | 4390 | 4070 | 3695 | 3372 | 3094 | 2845 | 2637 | 2390 | 2179 | 1953 | 1788 | 1636 | 1467 | 1367 | 1255 | 1158 | 1088 | 990 | 937 | 911 | 881 | 809 | 749 | 710 | 688 | 681 | 672 | 658 | 661 | 628 | 619 | 628 | 632 | 630 | 605 | 605 | 619 | 610 | 605 | 594 | 594 | 589 | 577 | 602 | 627 | 625 | 622 | 612 | 597 | 592 | 587 | 577 | 571 | 563 | 574 | 604 | 605 | 615 | 605 | 589 | 576 | 576 | 552 | 566 | 558 | 591 | 588 | 585 | 583 | 568 | 578 | 587 | 568 | 535 | 555 | 569 | 559 | 553 | 573 | 537 | 530 | 538 | 562 | 548 | 523 | 545 | 536 | 539 | 520 | 511 | 500 | 507 | 505 | 497 | 499 | 515 | 520 | 503 | 501 | 520 | 493 | 461 | 476 | 480 | 508 | 514 | 513 | 535 | 512 | 525 | 502 | 511 | 521 | 513 | 546 | 521 | 524 | 507 | 535 | 512 | 516 | 522 | 501 | 513 | 554 | 502 | 488 | 498 | 525 | 514 | 518 | 523 | 501 | 484 | 497 | 515 | 530 | 545 | 549 | 570 | 573 | 546 | 557 | 598 | 561 | 603 | 584 | 580 | 590 | 605 | 597 | 604 | 622 | 627 | 619 | 622 | 640 | 653 | 666 | 638 | 638 | 657 | 666 | 713 | 706 | 668 | 649 | 668 | 679 | 660 | 648 | 649 | 660 | 652 | 616 | 595 | 614 | 633 | 628 | 613 | 607 | 602 | 612 | 635 | 607 | 584 | 579 | 589 | 576 | 568 | 559 | 556 | 561 | 555 | 540 | 541 | 541 | 531 | 536 | 536 | 526 | 520 | 508 | 510 | 484 | 468 | 486 | 492 | 469 | 462 | 459 | 462 | 444 | 426 | 422 | 419 | 424 | 435 | 430 | 423 | 432 | 417 | 416 | 450 | 412 | 374 | 371 | 362 | 361 | 352 | 353 | 353 | 352 | 351 | 355 | 359 | 347 | 336 | 337 | 342 | 345 | 346 | 333 | 315 | 326 | 328 | 317 | 325 | 331 | 339 | 329 | 296 | 286 | 297 | 297 | 289 | 286 | 290 | 284 | 283 | 284 | 285 | 289 | 296 | 282 | 255 | 261 | 268 | 254 | 253 | 261 | 263 | 256 | 245 | 246 | 255 | 257 | 247 | 236 | 233 | 241 | 249 | 260 | 238 | 223 | 231 | 231 | 233 | 230 | 230 | 217 | 210 | 217 | 211 | 202 | 209 | 213 | 209 | 209 | 199 | 203 | 201 | 196 | 201 | 204 | 199 | 181 | 178 | 176 | 175 | 174 | 177 | 172 | 158 | 150 | 148 | 149 | 146 | 156 | 148 | 156 | 168 | 154 | 138 | 148 | 146 | 156 | 147 | 139 | 143 | 146 | 144 | 143 | 150 | 149 | 139 | 143 | 139 | 132 | 124 | 120 | 113 | 117 | 135 | 125 | 112 | 122 | 121 | 116 | 124 | 120 | 116 | 120 | 125 | 126 | 123 | 129 | 128 | 124 | 131 | 103 | 112 | 112 | 127 | 116 | 111 | 107 | 104 | 107 | 106 | 104 | 107 | 105 | 99 | 98 | 107 | 112 | 117 | 123 | 98 | 108 | 95 | 106 | 85 | 97 | 96 | 105 | 104 | 97 | 94 | 93 | 85 | 92 | 102 | 104 | 108 | 108 | 96 | 98 | 114 | 108 | 90 | 87 | 93 | 101 | 97 | 97 | 96 | 102 | 98 | 94 | 101 | 109 | 104 | 98 | 100 | 103 | 113 | 116 | 109 | 91 | 133 | 108 | 121 | 117 | 112 | 121 | 119 | 117 | 128 | 139 | 147 | 146 | 145 | 150 | 152 | 167 | 178 | 186 | 198 | 190 | 222 | 244 | 254 | 269 | 285 | 296 | 324 | 340 | 380 | 403 | 393 | 386 | 410 | 453 | 494 | 512 | 536 | 571 | 584 | 677 | 707 | 703 | 689 | 723 | 809 | 841 | 885 | 940 | 973 | 981 | 1033 | 1081 | 1072 | 1073 | 1126 | 1168 | 1178 | 1237 | 1232 | 1261 | 1328 | 1295 | 1316 | 1327 | 1335 | 1343 | 1372 | 1371 | 1361 | 1335 | 1326 | 1347 | 1371 | 1359 | 1316 | 1297 | 1293 | 1275 | 1253 | 1251 | 1206 | 1139 | 1125 | 1080 | 1056 | 991 | 938 | 924 | 895 | 844 | 801 | 779 | 745 | 708 | 687 | 651 | 593 | 565 | 548 | 522 | 487 | 434 | 384 | 352 | 348 | 327 | 310 | 283 | 249 | 239 | 229 | 199 | 170 | 163 | 154 | 124 | 112 | 109 | 103 | 105 | 94 | 84 | 72 | 62 | 59 | 54 | 55 | 52 | 41 | 34 | 32 | 34 | 37 | 36 | 32 | 29 | 21 | 16 | 19 | 21 | 20 | 23 | 21 | 14 | 15 | 17 | 16 | 18 | 20 | 18 | 15 | 14 | 16 | 20 | 16 | 14 | 16 | 18 | 14 | 13 | 18 | 21 | 16 | 15 | 16 | 14 | 14 | 15 | 13 | 11 | 11 | 13 | 13 | 13 | 12 | 14 | 14 | 17 | 18 | 17 | 17 | 10 | 13 | 17 | 13 | 11 | 13 | 13 | 13 | 13 | 15 | 13 | 10 | 11 | 12 | 11 | 15 | 15 | 11 | 10 | 14 | 13 | 14 | 16 | 18 | 12 | 11 | 14 | 11 | 10 | 11 | 11 | 9 | 9 | 8 | 13 | 10 | 10 | 9 | 15 | 11 | 8 | 9 | 8 | 10 | 8 | 9 | 10 | 16 | 16 | 9 | 11 | 10 | 9 | 11 | 12 | 17 | 17 | 15 | 14 | 8 | 15 | 12 | 7 | 9 | 10 | 11 | 11 | 12 | 11 | 10 | 13 | 14 | 12 | 13 | 9 | 10 | 9 | 13 | 9 | 7 | 11 | 16 | 9 | 15 | 8 | 11 | 12 | 9 | 12 | 14 | 10 | 12 | 14 | 9 | 9 | 6 | 7 | 11 | 11 | 5 | 8 | 8 | 5 | 7 | 8 | 7 | 5 | 4 | 5 | 10 | 11 | 8 | 7 | 8 | 11 | 10 | 9 | 10 | 6 | 4 | 3 | 5 | 7 | 6 | 7 | 6 | 5 | 7 | 11 | 8 | 7 | 7 | 3 | 6 | 7 | 8 | 10 | 8 | 6 | 8 | 7 | 9 | 10 | 8 | 8 | 8 | 7 | 6 | 7 | 7 | 6 | 8 | 9 | 7 | 4 | 5 | 6 | 5 | 4 | 4 | 4 | 6 | 5 | 5 | 6 | 8 | 6 | 3 | 3 | 4 | 5 | 7 | 5 | 3 | 3 | 5 | 6 | 6 | 6 | 6 | 8 | 7 | 6 | 5 | 3 | 3 | 3 | 4 | 5 | 5 | 6 | 5 | 4 | 7 | 7 | 4 | 5 | 7 | 8 | 7 | 7 | 7 | 5 | 4 | 3 | 4 | 4 | 4 | 3 | 3 | 5 | 5 | 6 | 5 | 4 | 4 | 6 | 6 | 7 | 10 | 9 | 5 | 4 | 4 | 3 | 3 | 3 | 4 | 5 | 5 | 4 | 5 | 4 | 5 | 9 | 6 | 3 | 5 | 5 | 6 | 7 | 5 | 6 | 6 | 5 | 6 | 6 | 4 | 4 | 4 | 4 | 5 | 3 | 3 | 2 | 3 | 4 | 4 | 3 | 2 | 5 | 4 | 3 | 4 | 5 | 4 | 8 | 8 | 5 | 5 | 4 | 2 | 2 | 2 | 4 | 2 | 5 | 1 | 2 | 2 | 5 | 4 | 3 | 5 | 7 | 5 | 2 | 3 | 3 | 5 | 3 | 4 | 3 | 5 | 4 | 2 | 5 | 6 | 6 | 4 | 0 | 3 | 6 | 4 | 5 | 5 | 4 | 4 | 2 | 5 | 7 | 5 | 3 | 3 | 2 | 2 | 6 | 3 | 9 | 5 | 4 | 5 | 10 | 7 | 8 | 2 | 6 | 1 | 5 | 5 | 7 | 4 | 6 | 5 | 2 | 4 | 4 | 6 | 3 | 6 | 4 | 1 | 2 | 6 | 5 | 3 | 7 | 7 | 9 | 8 | 4 | 3 | 6 | 8 | 7 | 6 | 8 | 9 | 4 | 11 | 6 | 6 | 8 | 5 | 8 | 6 | 7 | 11 | 6 | 6 | 7 | 8 | 4 | 8 | 6 | 11 | 7 | 10 | 5 | 7 | 6 | 8 | 8 | 5 | 6 | 5 | 6 | 3 | 0 | 2 | 3 | 8 | 5 | 4 | 4 | 4 | 3 | 3 | 5 | 4 | 7 | 6 | 4 | 4 | 6 | 4 | 4 | 5 | 3 | 4 | 2 | 1 | 3 | 4 | 2 | 4 | 4 | 1 | 4 | 5 | 3 | 2 | 1 | 2 | 2 | 3 | 3 | 3 | 2 | 3 | 2 | 2 | 2 | 2 | 2 | 1 | 2 | 2 | 2 | 1 | 1 | 1 | 3 | 1 | 0 | 1 | 2 | 2 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 2 | 3 | 3 | 2 | 1 | 2 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 1 | 2 | 1 | 0 | 3 | 4 | 2 | 3 | 1 | 0 | 0 | 2 | 1 | 0 | 1 | 1 | 0 | 2 | 2 | 1 | 0 | 0 | 1 | 2 | 3 | 2 | 1 | 0 | 2 | 2 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 1 | 2 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 0 | 2 | 2 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 2 | 2 | 1 | 1 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 0 | 2 | 4 | 2 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 2 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 2 | 0 | 0 | 2 | 1 | 0 | 0 | 2 | 2 | 0 | 1 | 1 | 0 | 3 | 0 | 1 | 0 | 0 | 2 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 2 | 1 | 0 | 3 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 2 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 1 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 2 | 1 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 7 | 240 | 17 | 27 | 15 | 16 | 19 | 23 | 26 | 30 | 45 | 41 | 50 | 59 | 73 | 75 | 101 | 128 | 137 | 167 | 216 | 237 | 274 | 383 | 429 | 520 | 625 | 678 | 749 | 845 | 926 | 969 | 996 | 1066 | 1152 | 1134 | 1101 | 1177 | 1157 | 1113 | 1146 | 1141 | 1086 | 1137 | 1174 | 1192 | 1125 | 1143 | 1134 | 1148 | 1177 | 1160 | 1172 | 1154 | 1138 | 1191 | 1133 | 1117 | 1143 | 1152 | 1201 | 1178 | 1180 | 1130 | 1147 | 1185 | 1230 | 1129 | 1178 | 1146 | 1130 | 1133 | 1135 | 1092 | 1115 | 1188 | 1240 | 1249 | 1138 | 1184 | 1202 | 1230 | 1199 | 1170 | 1156 | 1139 | 1181 | 1211 | 1232 | 1160 | 1186 | 1179 | 1145 | 1186 | 1177 | 1174 | 1188 | 1190 | 1187 | 1158 | 1122 | 1155 | 1128 | 1125 | 1131 | 1135 | 1118 | 1085 | 1101 | 1189 | 1137 | 1131 | 1131 | 1098 | 1098 | 1123 | 1127 | 1125 | 1120 | 1104 | 1120 | 1043 | 984 | 981 | 1025 | 1054 | 1005 | 996 | 1018 | 1022 | 1024 | 973 | 992 | 998 | 997 | 1011 | 1017 | 1037 | 1050 | 1084 | 1042 | 1021 | 1049 | 1034 | 992 | 993 | 1007 | 1028 | 1050 | 1051 | 1043 | 1031 | 1029 | 1034 | 1037 | 1025 | 1028 | 1045 | 1060 | 1109 | 1120 | 1128 | 1156 | 1189 | 1193 | 1209 | 1235 | 1301 | 1302 | 1278 | 1321 | 1345 | 1372 | 1412 | 1436 | 1458 | 1515 | 1488 | 1522 | 1595 | 1617 | 1545 | 1509 | 1531 | 1532 | 1518 | 1485 | 1479 | 1525 | 1524 | 1489 | 1495 | 1493 | 1471 | 1473 | 1474 | 1448 | 1392 | 1370 | 1389 | 1386 | 1366 | 1345 | 1307 | 1271 | 1267 | 1296 | 1301 | 1254 | 1226 | 1219 | 1199 | 1177 | 1164 | 1148 | 1150 | 1148 | 1134 | 1099 | 1071 | 1076 | 1088 | 1076 | 1051 | 1070 | 1062 | 1009 | 992 | 1002 | 1043 | 1032 | 1011 | 1000 | 962 | 964 | 947 | 964 | 948 | 905 | 913 | 938 | 942 | 944 | 928 | 905 | 905 | 911 | 907 | 899 | 889 | 876 | 867 | 878 | 875 | 839 | 831 | 850 | 842 | 833 | 823 | 797 | 795 | 785 | 805 | 809 | 806 | 789 | 786 | 800 | 772 | 775 | 756 | 746 | 749 | 732 | 718 | 733 | 690 | 665 | 702 | 657 | 668 | 677 | 658 | 673 | 697 | 664 | 639 | 669 | 694 | 698 | 669 | 668 | 683 | 655 | 648 | 641 | 620 | 626 | 639 | 638 | 635 | 638 | 635 | 616 | 612 | 615 | 610 | 603 | 586 | 585 | 595 | 586 | 563 | 575 | 561 | 576 | 593 | 587 | 589 | 580 | 589 | 580 | 559 | 578 | 531 | 542 | 533 | 542 | 524 | 533 | 528 | 487 | 487 | 499 | 492 | 479 | 478 | 488 | 484 | 458 | 456 | 439 | 463 | 467 | 446 | 466 | 460 | 437 | 434 | 454 | 461 | 463 | 443 | 425 | 427 | 448 | 467 | 448 | 437 | 450 | 447 | 428 | 413 | 408 | 436 | 417 | 386 | 416 | 396 | 396 | 430 | 419 | 419 | 385 | 423 | 404 | 383 | 372 | 376 | 391 | 372 | 385 | 394 | 357 | 372 | 369 | 372 | 359 | 362 | 348 | 383 | 392 | 400 | 363 | 347 | 374 | 375 | 354 | 339 | 329 | 325 | 353 | 378 | 383 | 398 | 375 | 337 | 347 | 353 | 353 | 367 | 361 | 354 | 366 | 378 | 356 | 341 | 354 | 343 | 334 | 329 | 346 | 340 | 340 | 333 | 358 | 339 | 352 | 349 | 346 | 348 | 341 | 338 | 339 | 357 | 362 | 375 | 384 | 335 | 358 | 357 | 376 | 390 | 415 | 416 | 417 | 439 | 473 | 490 | 515 | 559 | 582 | 595 | 646 | 694 | 761 | 839 | 895 | 971 | 1083 | 1197 | 1306 | 1412 | 1546 | 1711 | 1875 | 2065 | 2237 | 2445 | 2660 | 2832 | 3097 | 3426 | 3741 | 4063 | 4369 | 4627 | 4971 | 5293 | 5528 | 5814 | 6197 | 6606 | 6940 | 7160 | 7450 | 7820 | 8085 | 8294 | 8517 | 8762 | 8893 | 9011 | 9187 | 9351 | 9462 | 9436 | 9314 | 9295 | 9211 | 9097 | 8953 | 8733 | 8508 | 8232 | 7954 | 7671 | 7322 | 7026 | 6661 | 6204 | 5844 | 5550 | 5178 | 4769 | 4423 | 4121 | 3822 | 3494 | 3197 | 2950 | 2733 | 2515 | 2253 | 2004 | 1823 | 1694 | 1513 | 1369 | 1307 | 1249 | 1087 | 990 | 944 | 907 | 857 | 765 | 737 | 737 | 684 | 690 | 669 | 649 | 648 | 637 | 632 | 606 | 600 | 604 | 600 | 597 | 589 | 570 | 556 | 587 | 612 | 589 | 580 | 579 | 577 | 586 | 575 | 570 | 586 | 578 | 558 | 561 | 565 | 575 | 576 | 536 | 534 | 582 | 592 | 577 | 580 | 572 | 565 | 562 | 544 | 531 | 532 | 503 | 558 | 579 | 547 | 527 | 535 | 513 | 517 | 529 | 540 | 524 | 518 | 542 | 543 | 538 | 528 | 541 | 541 | 507 | 528 | 495 | 479 | 473 | 487 | 513 | 535 | 523 | 495 | 482 | 476 | 490 | 501 | 493 | 505 | 518 | 514 | 500 | 488 | 501 | 490 | 507 | 515 | 488 | 450 | 466 | 489 | 495 | 522 | 501 | 484 | 471 | 472 | 458 | 513 | 532 | 490 | 523 | 507 | 546 | 501 | 516 | 502 | 500 | 524 | 514 | 526 | 535 | 516 | 505 | 503 | 511 | 533 | 530 | 570 | 512 | 573 | 551 | 587 | 527 | 610 | 592 | 593 | 573 | 603 | 592 | 579 | 601 | 623 | 617 | 604 | 610 | 619 | 633 | 639 | 635 | 597 | 655 | 628 | 634 | 660 | 638 | 667 | 662 | 663 | 620 | 620 | 633 | 634 | 634 | 620 | 604 | 601 | 605 | 631 | 628 | 622 | 580 | 612 | 560 | 558 | 585 | 537 | 565 | 567 | 544 | 519 | 524 | 510 | 521 | 548 | 524 | 522 | 505 | 496 | 489 | 510 | 461 | 476 | 479 | 460 | 465 | 456 | 456 | 427 | 434 | 438 | 420 | 413 | 423 | 426 | 417 | 403 | 392 | 394 | 376 | 370 | 383 | 395 | 367 | 372 | 364 | 375 | 365 | 356 | 361 | 363 | 352 | 345 | 347 | 347 | 341 | 335 | 331 | 322 | 307 | 328 | 345 | 321 | 315 | 319 | 317 | 309 | 305 | 297 | 289 | 295 | 300 | 295 | 287 | 288 | 284 | 269 | 265 | 278 | 293 | 294 | 288 | 281 | 264 | 260 | 267 | 261 | 254 | 254 | 271 | 263 | 251 | 253 | 249 | 231 | 226 | 230 | 243 | 246 | 242 | 235 | 216 | 212 | 204 | 204 | 213 | 214 | 216 | 217 | 204 | 190 | 195 | 216 | 221 | 211 | 202 | 191 | 185 | 190 | 196 | 184 | 177 | 180 | 190 | 179 | 177 | 184 | 179 | 179 | 182 | 176 | 169 | 165 | 159 | 166 | 168 | 159 | 154 | 156 | 149 | 149 | 161 | 155 | 152 | 150 | 157 | 160 | 157 | 143 | 132 | 131 | 128 | 123 | 130 | 132 | 127 | 137 | 149 | 133 | 125 | 137 | 132 | 145 | 116 | 114 | 118 | 119 | 111 | 114 | 124 | 119 | 120 | 124 | 110 | 106 | 104 | 109 | 115 | 116 | 111 | 118 | 121 | 117 | 102 | 111 | 106 | 95 | 89 | 105 | 121 | 122 | 114 | 104 | 99 | 102 | 110 | 99 | 100 | 105 | 101 | 91 | 83 | 88 | 97 | 90 | 87 | 93 | 87 | 91 | 90 | 95 | 97 | 93 | 91 | 96 | 96 | 95 | 110 | 91 | 85 | 95 | 104 | 99 | 91 | 92 | 98 | 81 | 88 | 97 | 97 | 89 | 85 | 90 | 92 | 93 | 93 | 103 | 108 | 103 | 121 | 124 | 87 | 114 | 112 | 125 | 116 | 117 | 122 | 125 | 130 | 124 | 135 | 159 | 162 | 163 | 161 | 178 | 191 | 186 | 194 | 229 | 207 | 260 | 243 | 294 | 330 | 302 | 333 | 343 | 361 | 391 | 400 | 403 | 423 | 460 | 507 | 530 | 539 | 562 | 629 | 690 | 708 | 724 | 724 | 764 | 797 | 865 | 879 | 929 | 967 | 984 | 1014 | 1047 | 1081 | 1117 | 1149 | 1165 | 1166 | 1211 | 1273 | 1315 | 1272 | 1314 | 1316 | 1333 | 1353 | 1328 | 1346 | 1339 | 1337 | 1321 | 1305 | 1316 | 1319 | 1323 | 1324 | 1277 | 1262 | 1190 | 1181 | 1202 | 1147 | 1075 | 1045 | 997 | 984 | 918 | 908 | 870 | 829 | 787 | 735 | 705 | 705 | 687 | 634 | 611 | 570 | 521 | 493 | 438 | 415 | 417 | 383 | 350 | 310 | 284 | 264 | 240 | 224 | 206 | 184 | 165 | 151 | 142 | 134 | 124 | 114 | 109 | 102 | 89 | 83 | 72 | 66 | 62 | 57 | 48 | 37 | 31 | 28 | 27 | 27 | 29 | 30 | 25 | 20 | 19 | 21 | 23 | 21 | 18 | 15 | 13 | 14 | 17 | 17 | 17 | 16 | 16 | 15 | 13 | 14 | 13 | 13 | 16 | 14 | 10 | 10 | 13 | 13 | 13 | 16 | 14 | 11 | 11 | 12 | 12 | 10 | 13 | 17 | 16 | 18 | 16 | 14 | 14 | 13 | 16 | 15 | 11 | 14 | 15 | 12 | 11 | 13 | 13 | 10 | 13 | 14 | 12 | 13 | 17 | 16 | 12 | 11 | 11 | 10 | 12 | 9 | 8 | 9 | 12 | 14 | 12 | 16 | 17 | 13 | 10 | 9 | 11 | 11 | 10 | 12 | 10 | 10 | 10 | 9 | 12 | 13 | 10 | 9 | 8 | 12 | 12 | 10 | 15 | 13 | 13 | 15 | 12 | 10 | 11 | 9 | 8 | 7 | 9 | 9 | 7 | 5 | 9 | 8 | 11 | 11 | 10 | 10 | 8 | 7 | 9 | 8 | 9 | 9 | 11 | 11 | 9 | 10 | 13 | 11 | 12 | 8 | 6 | 9 | 9 | 8 | 12 | 10 | 7 | 10 | 10 | 8 | 10 | 11 | 8 | 5 | 8 | 8 | 9 | 6 | 10 | 10 | 8 | 6 | 12 | 6 | 8 | 10 | 7 | 10 | 11 | 7 | 5 | 7 | 7 | 8 | 12 | 11 | 8 | 9 | 6 | 6 | 3 | 7 | 5 | 5 | 8 | 6 | 6 | 5 | 5 | 5 | 3 | 3 | 3 | 6 | 8 | 8 | 3 | 5 | 9 | 8 | 5 | 3 | 6 | 4 | 7 | 12 | 10 | 7 | 5 | 6 | 5 | 5 | 6 | 6 | 8 | 9 | 10 | 8 | 4 | 6 | 8 | 2 | 4 | 7 | 8 | 5 | 3 | 6 | 5 | 3 | 3 | 5 | 3 | 4 | 5 | 7 | 4 | 2 | 5 | 7 | 5 | 4 | 6 | 7 | 8 | 7 | 3 | 5 | 6 | 5 | 4 | 4 | 4 | 6 | 7 | 5 | 4 | 6 | 9 | 6 | 6 | 5 | 4 | 4 | 6 | 5 | 5 | 7 | 7 | 5 | 5 | 7 | 7 | 6 | 6 | 5 | 5 | 3 | 4 | 4 | 5 | 4 | 5 | 6 | 4 | 1 | 1 | 2 | 3 | 4 | 4 | 5 | 6 | 5 | 4 | 4 | 5 | 3 | 3 | 3 | 5 | 4 | 4 | 4 | 3 | 3 | 4 | 4 | 5 | 5 | 4 | 3 | 3 | 2 | 2 | 4 | 6 | 4 | 3 | 5 | 4 | 3 | 3 | 3 | 3 | 4 | 5 | 6 | 5 | 2 | 4 | 4 | 3 | 1 | 2 | 3 | 3 | 4 | 3 | 2 | 2 | 5 | 4 | 2 | 1 | 6 | 7 | 3 | 5 | 8 | 6 | 6 | 4 | 4 | 5 | 3 | 2 | 4 | 5 | 4 | 6 | 5 | 4 | 6 | 6 | 5 | 4 | 3 | 2 | 3 | 5 | 3 | 3 | 5 | 5 | 5 | 4 | 4 | 3 | 5 | 5 | 4 | 6 | 4 | 2 | 3 | 3 | 5 | 5 | 5 | 7 | 3 | 5 | 5 | 3 | 3 | 8 | 7 | 5 | 7 | 7 | 5 | 6 | 3 | 5 | 4 | 8 | 10 | 5 | 5 | 4 | 1 | 5 | 7 | 9 | 5 | 6 | 5 | 4 | 6 | 6 | 8 | 6 | 4 | 5 | 5 | 4 | 6 | 7 | 9 | 6 | 4 | 3 | 9 | 7 | 6 | 8 | 4 | 4 | 5 | 4 | 3 | 6 | 7 | 4 | 5 | 7 | 5 | 6 | 5 | 8 | 8 | 7 | 3 | 6 | 3 | 6 | 6 | 3 | 6 | 6 | 5 | 6 | 5 | 7 | 8 | 2 | 6 | 4 | 4 | 6 | 3 | 6 | 1 | 2 | 2 | 9 | 2 | 4 | 4 | 4 | 2 | 3 | 0 | 7 | 5 | 5 | 2 | 1 | 4 | 2 | 1 | 3 | 0 | 2 | 1 | 2 | 3 | 4 | 2 | 1 | 4 | 1 | 4 | 0 | 1 | 1 | 2 | 0 | 1 | 0 | 1 | 0 | 2 | 3 | 1 | 3 | 1 | 4 | 1 | 1 | 0 | 0 | 0 | 1 | 2 | 1 | 2 | 2 | 0 | 2 | 2 | 3 | 0 | 0 | 1 | 3 | 2 | 3 | 1 | 1 | 2 | 2 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 2 | 2 | 1 | 1 | 2 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 2 | 2 | 2 | 1 | 1 | 2 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 2 | 1 | 2 | 3 | 2 | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 2 | 3 | 3 | 3 | 1 | 2 | 2 | 2 | 1 | 2 | 1 | 2 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 2 | 3 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 0 | 2 | 1 | 1 | 1 | 1 | 1 | 1 | 2 | 1 | 0 | 1 | 2 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 2 | 1 | 1 | 1 | 0 | 3 | 2 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 2 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 1 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 8 | 209 | 19 | 18 | 18 | 17 | 26 | 25 | 22 | 19 | 28 | 35 | 38 | 55 | 76 | 80 | 90 | 111 | 133 | 158 | 194 | 248 | 291 | 380 | 431 | 481 | 559 | 629 | 728 | 814 | 867 | 932 | 975 | 1061 | 1089 | 1114 | 1054 | 1083 | 1139 | 1123 | 1066 | 1092 | 1145 | 1112 | 1225 | 1239 | 1141 | 1155 | 1085 | 1114 | 1159 | 1172 | 1181 | 1139 | 1104 | 1157 | 1158 | 1114 | 1197 | 1132 | 1142 | 1172 | 1100 | 1106 | 1074 | 1075 | 1193 | 1132 | 1134 | 1131 | 1127 | 1104 | 1090 | 1113 | 1144 | 1197 | 1143 | 1262 | 1140 | 1134 | 1175 | 1210 | 1173 | 1177 | 1171 | 1179 | 1200 | 1248 | 1282 | 1128 | 1136 | 1120 | 1174 | 1156 | 1111 | 1100 | 1074 | 1065 | 1119 | 1050 | 1088 | 1051 | 1077 | 1108 | 1077 | 1086 | 1096 | 1089 | 1055 | 1124 | 1096 | 1096 | 1135 | 1099 | 1080 | 1094 | 1068 | 1062 | 1111 | 1103 | 1085 | 987 | 1017 | 1031 | 998 | 974 | 983 | 1021 | 1022 | 992 | 1019 | 995 | 972 | 981 | 1025 | 1016 | 1003 | 1027 | 1033 | 1036 | 1056 | 1035 | 994 | 964 | 960 | 964 | 1014 | 1040 | 1016 | 999 | 976 | 983 | 992 | 1010 | 1038 | 1035 | 1050 | 1049 | 1064 | 1082 | 1093 | 1092 | 1085 | 1099 | 1130 | 1164 | 1203 | 1225 | 1247 | 1265 | 1279 | 1374 | 1445 | 1415 | 1433 | 1427 | 1471 | 1495 | 1494 | 1489 | 1525 | 1545 | 1544 | 1537 | 1541 | 1539 | 1548 | 1514 | 1473 | 1453 | 1463 | 1460 | 1459 | 1463 | 1450 | 1420 | 1385 | 1374 | 1370 | 1338 | 1316 | 1324 | 1303 | 1258 | 1256 | 1252 | 1216 | 1210 | 1225 | 1224 | 1219 | 1185 | 1130 | 1143 | 1150 | 1114 | 1107 | 1123 | 1106 | 1085 | 1102 | 1082 | 1055 | 1074 | 1083 | 1080 | 1043 | 995 | 1004 | 1007 | 977 | 953 | 948 | 932 | 943 | 945 | 956 | 957 | 939 | 938 | 922 | 915 | 917 | 903 | 893 | 877 | 850 | 850 | 886 | 910 | 901 | 886 | 869 | 851 | 838 | 828 | 823 | 820 | 804 | 813 | 813 | 798 | 794 | 787 | 755 | 736 | 726 | 719 | 747 | 729 | 693 | 743 | 752 | 754 | 738 | 711 | 723 | 711 | 695 | 678 | 693 | 696 | 696 | 711 | 681 | 687 | 680 | 667 | 659 | 640 | 622 | 630 | 637 | 627 | 646 | 671 | 666 | 651 | 631 | 628 | 628 | 621 | 610 | 587 | 594 | 614 | 599 | 591 | 600 | 610 | 621 | 602 | 599 | 596 | 594 | 584 | 594 | 600 | 595 | 568 | 536 | 549 | 552 | 544 | 529 | 534 | 531 | 537 | 514 | 496 | 497 | 493 | 499 | 497 | 497 | 498 | 484 | 475 | 486 | 453 | 459 | 458 | 460 | 481 | 459 | 423 | 438 | 461 | 445 | 434 | 439 | 412 | 398 | 431 | 439 | 438 | 443 | 430 | 430 | 430 | 413 | 412 | 425 | 407 | 418 | 418 | 429 | 402 | 380 | 405 | 418 | 406 | 411 | 420 | 385 | 405 | 379 | 399 | 401 | 375 | 374 | 371 | 399 | 374 | 371 | 359 | 378 | 353 | 363 | 355 | 369 | 350 | 331 | 345 | 349 | 355 | 346 | 368 | 356 | 333 | 347 | 366 | 354 | 349 | 357 | 353 | 356 | 348 | 327 | 325 | 347 | 364 | 348 | 336 | 332 | 332 | 342 | 362 | 341 | 315 | 316 | 328 | 341 | 324 | 324 | 341 | 330 | 316 | 319 | 343 | 312 | 344 | 358 | 327 | 350 | 358 | 354 | 363 | 365 | 343 | 368 | 351 | 372 | 384 | 379 | 406 | 406 | 422 | 446 | 468 | 484 | 517 | 576 | 615 | 649 | 701 | 760 | 824 | 905 | 990 | 1063 | 1152 | 1245 | 1347 | 1515 | 1698 | 1836 | 1978 | 2150 | 2353 | 2598 | 2839 | 3065 | 3332 | 3589 | 3878 | 4188 | 4409 | 4777 | 5097 | 5377 | 5758 | 6074 | 6522 | 6857 | 7031 | 7424 | 7706 | 7920 | 8130 | 8345 | 8589 | 8767 | 8947 | 9055 | 9089 | 9054 | 9050 | 9118 | 9158 | 9115 | 8991 | 8781 | 8607 | 8409 | 8065 | 7722 | 7458 | 7176 | 6879 | 6576 | 6246 | 5923 | 5586 | 5169 | 4736 | 4382 | 4049 | 3762 | 3526 | 3247 | 2952 | 2695 | 2481 | 2254 | 2030 | 1874 | 1689 | 1511 | 1404 | 1314 | 1200 | 1067 | 989 | 906 | 834 | 799 | 776 | 768 | 719 | 683 | 684 | 673 | 644 | 657 | 644 | 618 | 602 | 609 | 611 | 584 | 591 | 589 | 567 | 568 | 580 | 579 | 573 | 578 | 586 | 562 | 555 | 573 | 572 | 568 | 566 | 565 | 561 | 560 | 574 | 574 | 553 | 543 | 550 | 558 | 570 | 567 | 560 | 568 | 537 | 526 | 541 | 521 | 534 | 531 | 557 | 532 | 560 | 599 | 529 | 510 | 516 | 503 | 506 | 536 | 536 | 533 | 520 | 500 | 508 | 518 | 506 | 502 | 496 | 486 | 484 | 501 | 502 | 498 | 500 | 503 | 527 | 510 | 504 | 514 | 499 | 490 | 469 | 473 | 485 | 475 | 490 | 487 | 485 | 468 | 486 | 506 | 502 | 505 | 493 | 493 | 514 | 471 | 473 | 508 | 477 | 464 | 471 | 511 | 505 | 495 | 523 | 494 | 489 | 471 | 475 | 488 | 483 | 500 | 535 | 524 | 489 | 509 | 542 | 511 | 538 | 540 | 521 | 545 | 557 | 591 | 512 | 562 | 593 | 573 | 602 | 591 | 625 | 600 | 574 | 590 | 609 | 634 | 660 | 647 | 600 | 574 | 601 | 551 | 642 | 655 | 646 | 644 | 650 | 639 | 606 | 620 | 627 | 633 | 637 | 633 | 641 | 639 | 603 | 592 | 610 | 623 | 612 | 638 | 626 | 600 | 555 | 584 | 560 | 571 | 552 | 552 | 541 | 532 | 529 | 553 | 559 | 526 | 515 | 508 | 477 | 481 | 498 | 489 | 488 | 464 | 468 | 486 | 473 | 451 | 449 | 434 | 432 | 437 | 434 | 428 | 428 | 405 | 370 | 380 | 393 | 391 | 394 | 418 | 401 | 411 | 396 | 372 | 369 | 357 | 350 | 379 | 364 | 348 | 340 | 331 | 336 | 341 | 328 | 325 | 335 | 319 | 327 | 320 | 316 | 313 | 321 | 319 | 315 | 299 | 309 | 295 | 277 | 270 | 281 | 292 | 287 | 288 | 296 | 289 | 270 | 261 | 261 | 247 | 259 | 273 | 268 | 268 | 257 | 251 | 250 | 252 | 241 | 235 | 236 | 244 | 240 | 234 | 239 | 229 | 210 | 213 | 222 | 230 | 234 | 215 | 216 | 211 | 203 | 207 | 215 | 219 | 205 | 192 | 187 | 193 | 203 | 201 | 195 | 188 | 191 | 197 | 197 | 189 | 193 | 186 | 170 | 174 | 167 | 156 | 159 | 167 | 176 | 183 | 171 | 160 | 158 | 162 | 162 | 150 | 139 | 149 | 162 | 156 | 145 | 141 | 143 | 146 | 144 | 143 | 142 | 145 | 140 | 131 | 130 | 132 | 131 | 121 | 118 | 121 | 133 | 135 | 131 | 130 | 134 | 128 | 125 | 136 | 126 | 117 | 117 | 115 | 116 | 111 | 110 | 120 | 120 | 107 | 101 | 101 | 94 | 96 | 94 | 104 | 119 | 119 | 110 | 108 | 106 | 104 | 106 | 102 | 101 | 98 | 88 | 89 | 93 | 87 | 91 | 93 | 82 | 82 | 90 | 112 | 94 | 106 | 101 | 86 | 87 | 97 | 102 | 101 | 100 | 90 | 84 | 95 | 100 | 96 | 95 | 83 | 100 | 93 | 93 | 98 | 95 | 88 | 94 | 90 | 94 | 94 | 93 | 93 | 96 | 94 | 93 | 95 | 99 | 104 | 106 | 102 | 111 | 108 | 114 | 105 | 113 | 124 | 120 | 116 | 114 | 125 | 120 | 123 | 127 | 131 | 147 | 157 | 158 | 159 | 142 | 141 | 171 | 201 | 210 | 230 | 221 | 232 | 247 | 278 | 286 | 314 | 337 | 330 | 370 | 400 | 405 | 429 | 458 | 464 | 484 | 527 | 538 | 576 | 658 | 706 | 646 | 693 | 756 | 795 | 824 | 820 | 826 | 889 | 933 | 972 | 1028 | 1058 | 1058 | 1092 | 1161 | 1159 | 1152 | 1190 | 1260 | 1313 | 1293 | 1272 | 1316 | 1304 | 1299 | 1307 | 1341 | 1338 | 1323 | 1335 | 1333 | 1306 | 1271 | 1227 | 1248 | 1275 | 1201 | 1201 | 1176 | 1145 | 1118 | 1059 | 1043 | 1017 | 1018 | 971 | 909 | 892 | 831 | 783 | 766 | 733 | 698 | 665 | 633 | 580 | 535 | 508 | 473 | 489 | 423 | 401 | 357 | 338 | 300 | 265 | 277 | 269 | 236 | 203 | 183 | 177 | 170 | 153 | 136 | 125 | 108 | 100 | 98 | 81 | 73 | 72 | 66 | 48 | 47 | 46 | 40 | 32 | 31 | 34 | 33 | 27 | 27 | 26 | 22 | 23 | 24 | 19 | 22 | 19 | 14 | 19 | 17 | 13 | 12 | 13 | 20 | 20 | 15 | 13 | 14 | 14 | 14 | 14 | 14 | 12 | 15 | 15 | 12 | 16 | 14 | 9 | 12 | 13 | 13 | 13 | 11 | 8 | 11 | 14 | 11 | 11 | 14 | 16 | 14 | 10 | 10 | 13 | 12 | 10 | 12 | 14 | 14 | 16 | 16 | 14 | 15 | 16 | 16 | 14 | 11 | 13 | 12 | 13 | 11 | 11 | 12 | 16 | 12 | 8 | 8 | 9 | 11 | 11 | 10 | 10 | 13 | 16 | 15 | 15 | 13 | 10 | 10 | 13 | 10 | 10 | 14 | 14 | 8 | 8 | 16 | 15 | 12 | 10 | 8 | 9 | 9 | 9 | 10 | 12 | 12 | 10 | 9 | 7 | 9 | 11 | 15 | 10 | 9 | 11 | 13 | 14 | 7 | 6 | 11 | 11 | 12 | 12 | 12 | 11 | 7 | 6 | 8 | 6 | 7 | 4 | 9 | 9 | 9 | 7 | 10 | 9 | 7 | 12 | 11 | 8 | 10 | 10 | 9 | 10 | 11 | 10 | 10 | 8 | 11 | 11 | 8 | 8 | 11 | 11 | 6 | 7 | 9 | 9 | 8 | 8 | 8 | 7 | 6 | 7 | 12 | 13 | 9 | 9 | 5 | 8 | 7 | 7 | 8 | 3 | 5 | 6 | 8 | 6 | 6 | 4 | 5 | 6 | 5 | 5 | 7 | 7 | 8 | 7 | 3 | 3 | 3 | 4 | 7 | 4 | 3 | 7 | 7 | 9 | 11 | 9 | 6 | 6 | 9 | 11 | 8 | 1 | 2 | 6 | 8 | 8 | 10 | 7 | 7 | 2 | 4 | 4 | 3 | 6 | 5 | 3 | 5 | 5 | 6 | 5 | 5 | 6 | 5 | 5 | 6 | 4 | 6 | 4 | 7 | 7 | 4 | 5 | 2 | 5 | 5 | 5 | 3 | 3 | 6 | 6 | 5 | 4 | 5 | 5 | 3 | 4 | 4 | 9 | 4 | 5 | 5 | 5 | 5 | 5 | 6 | 4 | 4 | 5 | 6 | 6 | 3 | 6 | 8 | 5 | 3 | 5 | 5 | 3 | 6 | 5 | 6 | 5 | 4 | 3 | 2 | 4 | 5 | 5 | 4 | 3 | 2 | 4 | 4 | 4 | 6 | 4 | 5 | 5 | 2 | 2 | 3 | 2 | 3 | 4 | 4 | 3 | 3 | 3 | 4 | 3 | 3 | 3 | 3 | 3 | 5 | 4 | 2 | 4 | 4 | 2 | 3 | 3 | 2 | 2 | 3 | 4 | 4 | 4 | 4 | 4 | 2 | 1 | 2 | 2 | 2 | 2 | 3 | 3 | 1 | 3 | 3 | 2 | 1 | 2 | 4 | 3 | 2 | 3 | 3 | 4 | 4 | 3 | 2 | 3 | 2 | 5 | 5 | 3 | 5 | 5 | 4 | 3 | 4 | 5 | 5 | 3 | 3 | 4 | 3 | 4 | 5 | 6 | 5 | 3 | 3 | 4 | 6 | 6 | 3 | 2 | 2 | 5 | 5 | 2 | 3 | 4 | 3 | 5 | 5 | 5 | 3 | 6 | 6 | 3 | 3 | 6 | 6 | 4 | 3 | 2 | 5 | 7 | 7 | 5 | 3 | 4 | 6 | 7 | 4 | 4 | 3 | 7 | 4 | 5 | 6 | 5 | 8 | 6 | 5 | 4 | 4 | 4 | 5 | 9 | 6 | 12 | 7 | 6 | 8 | 7 | 3 | 5 | 5 | 6 | 8 | 5 | 4 | 3 | 7 | 6 | 6 | 8 | 10 | 8 | 3 | 3 | 5 | 6 | 11 | 10 | 4 | 6 | 8 | 11 | 8 | 4 | 4 | 2 | 1 | 4 | 5 | 4 | 8 | 4 | 2 | 2 | 7 | 4 | 4 | 5 | 6 | 4 | 3 | 1 | 4 | 3 | 3 | 0 | 6 | 3 | 1 | 2 | 7 | 2 | 0 | 3 | 5 | 2 | 2 | 2 | 1 | 0 | 1 | 4 | 1 | 2 | 1 | 1 | 1 | 2 | 3 | 2 | 0 | 2 | 0 | 0 | 3 | 3 | 0 | 2 | 0 | 1 | 0 | 2 | 1 | 0 | 2 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 2 | 1 | 3 | 0 | 1 | 0 | 1 | 2 | 2 | 1 | 1 | 3 | 0 | 2 | 0 | 2 | 1 | 1 | 2 | 2 | 2 | 2 | 2 | 0 | 2 | 1 | 1 | 1 | 2 | 2 | 0 | 2 | 0 | 1 | 1 | 0 | 1 | 2 | 2 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 2 | 2 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 2 | 2 | 0 | 1 | 1 | 1 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 2 | 1 | 2 | 2 | 1 | 0 | 1 | 1 | 1 | 2 | 2 | 1 | 1 | 1 | 2 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 3 | 2 | 1 | 1 | 2 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 2 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 9 | 221 | 25 | 20 | 17 | 18 | 17 | 20 | 22 | 28 | 33 | 33 | 34 | 48 | 63 | 71 | 93 | 115 | 142 | 165 | 195 | 242 | 281 | 374 | 429 | 493 | 579 | 659 | 748 | 830 | 930 | 950 | 957 | 1065 | 1077 | 1036 | 1089 | 1091 | 1064 | 1112 | 1098 | 1101 | 1133 | 1098 | 1102 | 1184 | 1131 | 1149 | 1134 | 1158 | 1148 | 1137 | 1138 | 1129 | 1170 | 1178 | 1152 | 1124 | 1109 | 1116 | 1152 | 1075 | 1091 | 1141 | 1136 | 1109 | 1223 | 1077 | 1139 | 1149 | 1066 | 1108 | 1157 | 1186 | 1135 | 1153 | 1123 | 1294 | 1194 | 1127 | 1206 | 1179 | 1153 | 1170 | 1183 | 1171 | 1161 | 1199 | 1170 | 1097 | 1088 | 1141 | 1125 | 1145 | 1148 | 1092 | 1099 | 1122 | 1124 | 1099 | 1139 | 1127 | 1051 | 1036 | 1053 | 1024 | 1092 | 1100 | 1066 | 1133 | 1093 | 1074 | 1085 | 1124 | 1093 | 1083 | 1082 | 1030 | 1067 | 1107 | 1091 | 1027 | 985 | 991 | 990 | 1000 | 982 | 1004 | 965 | 977 | 971 | 971 | 957 | 990 | 1009 | 1009 | 1002 | 985 | 962 | 959 | 963 | 989 | 1029 | 1013 | 988 | 1010 | 1014 | 1009 | 1009 | 1022 | 1035 | 1027 | 1047 | 1052 | 1048 | 1063 | 1044 | 1037 | 1055 | 1052 | 1051 | 1051 | 1051 | 1091 | 1158 | 1214 | 1236 | 1225 | 1266 | 1266 | 1276 | 1355 | 1420 | 1447 | 1417 | 1424 | 1488 | 1480 | 1490 | 1528 | 1532 | 1513 | 1497 | 1530 | 1534 | 1459 | 1450 | 1504 | 1499 | 1462 | 1473 | 1451 | 1416 | 1414 | 1406 | 1385 | 1366 | 1326 | 1322 | 1351 | 1374 | 1372 | 1327 | 1278 | 1263 | 1246 | 1187 | 1170 | 1192 | 1197 | 1182 | 1165 | 1165 | 1144 | 1108 | 1106 | 1093 | 1049 | 1074 | 1097 | 1071 | 1063 | 1065 | 1075 | 1038 | 1010 | 1002 | 999 | 1000 | 979 | 963 | 957 | 983 | 967 | 962 | 984 | 954 | 926 | 926 | 934 | 927 | 924 | 913 | 889 | 880 | 868 | 854 | 856 | 861 | 865 | 874 | 856 | 818 | 807 | 814 | 815 | 807 | 785 | 782 | 797 | 812 | 816 | 811 | 770 | 730 | 773 | 777 | 752 | 753 | 762 | 732 | 710 | 705 | 714 | 715 | 725 | 710 | 682 | 685 | 668 | 689 | 709 | 705 | 679 | 661 | 690 | 663 | 639 | 650 | 645 | 647 | 648 | 629 | 655 | 666 | 640 | 625 | 621 | 629 | 637 | 626 | 606 | 595 | 595 | 591 | 575 | 585 | 609 | 592 | 593 | 589 | 600 | 616 | 585 | 558 | 579 | 600 | 579 | 583 | 575 | 569 | 524 | 530 | 525 | 540 | 510 | 523 | 534 | 546 | 498 | 492 | 510 | 490 | 480 | 479 | 504 | 475 | 446 | 478 | 463 | 490 | 474 | 468 | 455 | 448 | 466 | 427 | 439 | 450 | 430 | 426 | 434 | 439 | 438 | 422 | 410 | 415 | 413 | 422 | 429 | 413 | 386 | 377 | 401 | 393 | 403 | 372 | 414 | 413 | 418 | 398 | 391 | 379 | 381 | 417 | 351 | 396 | 389 | 388 | 353 | 390 | 401 | 374 | 369 | 371 | 379 | 359 | 359 | 382 | 385 | 381 | 369 | 356 | 356 | 341 | 354 | 368 | 365 | 350 | 350 | 361 | 352 | 337 | 350 | 344 | 333 | 332 | 336 | 342 | 341 | 341 | 333 | 327 | 334 | 337 | 317 | 318 | 318 | 324 | 334 | 332 | 331 | 371 | 337 | 318 | 327 | 322 | 353 | 357 | 340 | 320 | 345 | 341 | 321 | 321 | 337 | 343 | 327 | 317 | 350 | 361 | 357 | 376 | 396 | 421 | 414 | 403 | 407 | 463 | 491 | 499 | 497 | 536 | 598 | 646 | 716 | 785 | 828 | 884 | 963 | 1054 | 1175 | 1292 | 1397 | 1522 | 1642 | 1811 | 2017 | 2217 | 2410 | 2597 | 2852 | 3068 | 3284 | 3538 | 3868 | 4206 | 4485 | 4828 | 5180 | 5424 | 5802 | 6077 | 6357 | 6799 | 7049 | 7338 | 7578 | 7732 | 8137 | 8526 | 8631 | 8740 | 8915 | 9062 | 9133 | 9197 | 9339 | 9192 | 8999 | 8945 | 8882 | 8698 | 8482 | 8282 | 8068 | 7807 | 7502 | 7110 | 6784 | 6532 | 6183 | 5787 | 5421 | 5075 | 4806 | 4499 | 4082 | 3755 | 3476 | 3168 | 2917 | 2687 | 2463 | 2227 | 2007 | 1850 | 1673 | 1513 | 1380 | 1258 | 1159 | 1063 | 982 | 932 | 878 | 820 | 792 | 773 | 713 | 689 | 679 | 673 | 667 | 620 | 614 | 635 | 613 | 595 | 593 | 587 | 568 | 555 | 563 | 564 | 571 | 585 | 597 | 595 | 567 | 557 | 573 | 571 | 554 | 544 | 547 | 544 | 542 | 551 | 559 | 574 | 559 | 547 | 559 | 546 | 561 | 561 | 548 | 553 | 541 | 550 | 536 | 547 | 530 | 533 | 525 | 536 | 514 | 552 | 537 | 528 | 545 | 520 | 516 | 523 | 521 | 490 | 508 | 500 | 493 | 490 | 489 | 513 | 504 | 501 | 492 | 500 | 493 | 496 | 499 | 502 | 496 | 485 | 484 | 481 | 481 | 486 | 481 | 486 | 492 | 475 | 463 | 480 | 499 | 484 | 482 | 485 | 496 | 500 | 482 | 495 | 466 | 460 | 488 | 482 | 500 | 479 | 484 | 497 | 482 | 520 | 483 | 477 | 492 | 478 | 479 | 468 | 474 | 483 | 485 | 482 | 488 | 503 | 508 | 501 | 497 | 531 | 528 | 533 | 485 | 578 | 548 | 530 | 563 | 598 | 560 | 572 | 552 | 548 | 556 | 576 | 597 | 608 | 613 | 608 | 612 | 647 | 644 | 625 | 654 | 668 | 646 | 588 | 585 | 617 | 639 | 624 | 602 | 604 | 617 | 626 | 623 | 639 | 662 | 619 | 589 | 595 | 624 | 633 | 615 | 590 | 574 | 580 | 591 | 557 | 567 | 554 | 548 | 541 | 539 | 535 | 520 | 514 | 525 | 525 | 504 | 476 | 502 | 482 | 464 | 464 | 458 | 428 | 455 | 430 | 429 | 452 | 441 | 415 | 422 | 425 | 408 | 392 | 387 | 397 | 378 | 365 | 373 | 354 | 365 | 375 | 354 | 362 | 361 | 374 | 376 | 363 | 342 | 332 | 334 | 324 | 321 | 331 | 326 | 335 | 361 | 344 | 331 | 337 | 311 | 324 | 303 | 276 | 271 | 273 | 274 | 278 | 280 | 270 | 285 | 300 | 290 | 283 | 277 | 266 | 268 | 268 | 271 | 259 | 260 | 253 | 231 | 244 | 262 | 263 | 254 | 252 | 248 | 241 | 239 | 252 | 243 | 216 | 210 | 212 | 222 | 225 | 221 | 231 | 243 | 238 | 222 | 218 | 216 | 210 | 207 | 214 | 205 | 196 | 195 | 199 | 209 | 200 | 190 | 191 | 180 | 178 | 188 | 187 | 181 | 182 | 190 | 192 | 185 | 178 | 160 | 158 | 174 | 184 | 173 | 160 | 156 | 161 | 163 | 158 | 157 | 154 | 149 | 142 | 150 | 146 | 138 | 137 | 137 | 138 | 132 | 141 | 149 | 141 | 140 | 138 | 121 | 114 | 130 | 126 | 117 | 118 | 111 | 114 | 113 | 120 | 117 | 107 | 105 | 109 | 116 | 114 | 111 | 114 | 113 | 110 | 111 | 113 | 112 | 116 | 118 | 100 | 98 | 107 | 101 | 106 | 109 | 103 | 100 | 105 | 117 | 105 | 99 | 104 | 104 | 101 | 100 | 101 | 103 | 98 | 92 | 84 | 89 | 98 | 99 | 98 | 98 | 98 | 103 | 99 | 94 | 96 | 91 | 81 | 78 | 85 | 90 | 84 | 89 | 94 | 81 | 93 | 85 | 81 | 95 | 92 | 95 | 95 | 93 | 95 | 79 | 79 | 94 | 96 | 90 | 92 | 93 | 98 | 113 | 105 | 89 | 109 | 97 | 114 | 110 | 102 | 103 | 113 | 104 | 107 | 115 | 129 | 125 | 125 | 132 | 132 | 141 | 154 | 170 | 186 | 204 | 201 | 194 | 213 | 213 | 241 | 256 | 264 | 281 | 282 | 325 | 361 | 378 | 394 | 406 | 427 | 454 | 475 | 496 | 483 | 537 | 532 | 633 | 632 | 707 | 745 | 756 | 783 | 798 | 835 | 878 | 893 | 961 | 1002 | 1005 | 1047 | 1079 | 1074 | 1130 | 1170 | 1145 | 1186 | 1224 | 1280 | 1245 | 1295 | 1261 | 1290 | 1293 | 1293 | 1302 | 1322 | 1310 | 1277 | 1284 | 1297 | 1273 | 1268 | 1265 | 1207 | 1194 | 1194 | 1135 | 1136 | 1066 | 1060 | 1041 | 1003 | 968 | 919 | 904 | 863 | 816 | 779 | 767 | 753 | 704 | 646 | 595 | 559 | 534 | 502 | 502 | 463 | 416 | 401 | 374 | 340 | 308 | 287 | 265 | 232 | 202 | 184 | 173 | 167 | 153 | 136 | 128 | 124 | 116 | 107 | 95 | 72 | 59 | 54 | 51 | 51 | 45 | 44 | 38 | 37 | 36 | 35 | 35 | 30 | 27 | 26 | 24 | 20 | 20 | 18 | 12 | 16 | 19 | 18 | 18 | 17 | 19 | 21 | 16 | 14 | 15 | 16 | 16 | 13 | 12 | 15 | 16 | 15 | 15 | 10 | 10 | 14 | 12 | 14 | 13 | 15 | 16 | 13 | 15 | 17 | 14 | 12 | 12 | 12 | 12 | 11 | 9 | 11 | 12 | 13 | 16 | 14 | 10 | 11 | 12 | 17 | 15 | 11 | 10 | 11 | 11 | 13 | 13 | 10 | 11 | 12 | 8 | 8 | 9 | 13 | 12 | 13 | 14 | 11 | 10 | 11 | 8 | 7 | 9 | 13 | 16 | 14 | 11 | 9 | 10 | 13 | 13 | 11 | 13 | 10 | 10 | 10 | 9 | 9 | 12 | 12 | 10 | 9 | 9 | 9 | 10 | 12 | 11 | 12 | 14 | 17 | 14 | 13 | 11 | 10 | 7 | 8 | 7 | 11 | 13 | 12 | 11 | 9 | 7 | 7 | 8 | 10 | 11 | 12 | 11 | 6 | 7 | 7 | 9 | 12 | 10 | 8 | 6 | 7 | 13 | 14 | 15 | 10 | 9 | 12 | 10 | 8 | 10 | 10 | 6 | 4 | 6 | 12 | 9 | 11 | 8 | 9 | 11 | 9 | 10 | 8 | 6 | 9 | 7 | 7 | 7 | 7 | 6 | 5 | 5 | 6 | 9 | 11 | 4 | 6 | 5 | 4 | 3 | 5 | 6 | 8 | 6 | 7 | 5 | 4 | 6 | 7 | 6 | 7 | 8 | 7 | 5 | 7 | 4 | 4 | 7 | 5 | 6 | 5 | 4 | 7 | 5 | 8 | 6 | 4 | 6 | 6 | 6 | 6 | 4 | 6 | 13 | 5 | 7 | 5 | 8 | 3 | 3 | 6 | 8 | 4 | 5 | 9 | 8 | 7 | 5 | 5 | 3 | 4 | 4 | 7 | 5 | 3 | 4 | 6 | 8 | 7 | 6 | 4 | 6 | 3 | 3 | 6 | 7 | 6 | 5 | 5 | 6 | 3 | 2 | 3 | 5 | 2 | 4 | 4 | 6 | 5 | 5 | 6 | 6 | 5 | 4 | 3 | 4 | 4 | 2 | 4 | 6 | 4 | 3 | 5 | 3 | 4 | 6 | 7 | 4 | 6 | 4 | 2 | 2 | 4 | 5 | 6 | 5 | 2 | 3 | 4 | 4 | 4 | 4 | 5 | 4 | 4 | 3 | 2 | 2 | 4 | 7 | 5 | 4 | 3 | 3 | 3 | 3 | 4 | 3 | 3 | 5 | 6 | 6 | 4 | 3 | 3 | 2 | 5 | 8 | 4 | 4 | 4 | 3 | 3 | 4 | 3 | 3 | 2 | 2 | 3 | 4 | 3 | 2 | 4 | 5 | 3 | 2 | 3 | 5 | 5 | 4 | 6 | 4 | 3 | 4 | 5 | 4 | 3 | 4 | 4 | 3 | 4 | 5 | 5 | 3 | 3 | 4 | 2 | 2 | 4 | 4 | 4 | 5 | 5 | 5 | 4 | 5 | 5 | 5 | 5 | 4 | 2 | 2 | 3 | 5 | 7 | 6 | 3 | 3 | 5 | 7 | 5 | 3 | 4 | 8 | 6 | 2 | 3 | 3 | 2 | 3 | 3 | 3 | 4 | 3 | 1 | 2 | 3 | 3 | 4 | 3 | 4 | 6 | 5 | 7 | 7 | 4 | 3 | 7 | 8 | 4 | 5 | 5 | 5 | 7 | 5 | 4 | 6 | 7 | 3 | 7 | 5 | 4 | 7 | 5 | 5 | 7 | 8 | 4 | 3 | 4 | 5 | 3 | 5 | 10 | 7 | 5 | 7 | 5 | 5 | 10 | 7 | 8 | 6 | 6 | 4 | 5 | 7 | 4 | 5 | 5 | 4 | 6 | 4 | 3 | 10 | 6 | 3 | 4 | 3 | 5 | 4 | 7 | 3 | 5 | 4 | 3 | 3 | 3 | 3 | 5 | 2 | 3 | 0 | 3 | 3 | 3 | 3 | 1 | 5 | 3 | 3 | 4 | 4 | 2 | 2 | 3 | 4 | 3 | 3 | 3 | 4 | 1 | 0 | 0 | 4 | 2 | 1 | 0 | 0 | 2 | 2 | 2 | 1 | 2 | 1 | 5 | 1 | 4 | 1 | 0 | 2 | 3 | 1 | 2 | 3 | 0 | 3 | 1 | 2 | 2 | 0 | 1 | 3 | 4 | 0 | 1 | 2 | 2 | 0 | 1 | 1 | 1 | 2 | 5 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 3 | 1 | 2 | 0 | 1 | 4 | 1 | 1 | 4 | 1 | 2 | 0 | 2 | 1 | 0 | 1 | 2 | 3 | 3 | 1 | 0 | 2 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 2 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 2 | 0 | 1 | 3 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 2 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 2 | 2 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 2 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 2 | 1 | 1 | 1 | 2 | 2 | 1 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 2 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 2 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 2 | 2 | 0 | 0 | 1 | 2 | 3 | 0 | 0 | 1 | 1 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 3 | 0 | 2 | 2 | 1 | 1 | 1 | 2 | 4 | 4 | 3 | 3 | 4 | 2 | 1 | 3 | 0 | 2 | 2 | 0 | 1 | 3 | 2 | 1 | 1 | 3 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 2 | 0 | 0 | 2 | 2 | 2 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 3 | 2 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 2 | 3 | 2 | 2 | 0 | 2 | 1 | 0 | 1 | 3 | 2 | 1 | 1 | 3 | 1 | 0 | 0 | 2 | 0 | 1 | 1 | 2 | 2 | 2 | 0 | 1 | 0 | 2 | 3 | 0 | 1 | 2 | 1 | 0 | 1 | 2 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 2 | 3 | 2 | 0 | 1 | 1 | 1 | 0 | 0 | 2 | 3 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 2 | 1 | 2 | 0 | 0 | 0 | 1 | 0 | 2 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 2 | 0 | 1 | 0 | 3 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 2 | 1 | 1 | 2 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 2 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0 | 1 | 0 | 0 | 2 | 0 | 1 | 0 | 1 | 0 | 2 | 1 | 2 | 0 | 1 | 1 | 2 | 1 | 0 | 1 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 3 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 2 | 2 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0