## Results
Hidex Automatic Gamma Counter
Counting 18F
Run
Template | Counting 18F
Run status | Finished
Run started | 2021-09-01 20:35
Run type | Count only, single rack
Software version | 1.8.0.0
AMG serial number | 2210227
Firmware versions | 0: 50, 1: 50, F: 100
Counting time | 60
Result normalization | 2021-09-01 11:48
Comment
Counting windows
Window start | Window end | Additional windows | Isotope | Counting efficiency | Half-life
400 | 650 |  | ¹⁸F | 1 | 1.82951 h
Racks
Rack | Time | Barcode | Slot count | Latest QC check | QC status
1 | 2021-09-01 20:35 | 006 | 10 | 2021-09-01 10:39 | Calibration suggested
Results
Rack | Vial | Time | Counted time (s) | Dead time factor | ¹⁸F (counts) | ¹⁸F (CPM) | Normalized ¹⁸F (CPM) | Normalized ¹⁸F (Bq) | Tags
1 | 1 | 2021-09-01 20:37 | 60 | 1.000 | 45 | 45 | 1267 | 21
1 | 2 | 2021-09-01 20:38 | 60 | 1.184 | 783897 | 927766 | 26362722 | 439379
1 | 3 | 2021-09-01 20:40 | 60 | 1.181 | 777187 | 918171 | 26334630 | 438910
1 | 4 | 2021-09-01 20:41 | 60 | 1.178 | 765797 | 902358 | 26123450 | 435391
1 | 5 | 2021-09-01 20:42 | 60 | 1.177 | 763678 | 898972 | 26269526 | 437825
1 | 6 | 2021-09-01 20:44 | 60 | 1.063 | 298541 | 317254 | 9358267 | 155971
1 | 7 | 2021-09-01 20:45 | 60 | 1.060 | 286994 | 304185 | 9057398 | 150957
1 | 8 | 2021-09-01 20:47 | 60 | 1.059 | 282808 | 299600 | 9004880 | 150081
1 | 9 | 2021-09-01 20:48 | 60 | 1.059 | 280496 | 296964 | 9009778 | 150163
1 | 10 | 2021-09-01 20:50 | 60 | 1.000 | 44 | 44 | 1348 | 22

## Spectra
Hidex Automatic Gamma Counter
Counting 18F
Run
Template | Counting 18F
Run status | Finished
Run started | 2021-09-01 20:35
Run type | Count only, single rack
Software version | 1.8.0.0
AMG serial number | 2210227
Firmware versions | 0: 50, 1: 50, F: 100
Counting time | 60
Result normalization | 2021-09-01 11:48
Comment
Spectra
Channel
Rack | Vial | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66 | 67 | 68 | 69 | 70 | 71 | 72 | 73 | 74 | 75 | 76 | 77 | 78 | 79 | 80 | 81 | 82 | 83 | 84 | 85 | 86 | 87 | 88 | 89 | 90 | 91 | 92 | 93 | 94 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 117 | 118 | 119 | 120 | 121 | 122 | 123 | 124 | 125 | 126 | 127 | 128 | 129 | 130 | 131 | 132 | 133 | 134 | 135 | 136 | 137 | 138 | 139 | 140 | 141 | 142 | 143 | 144 | 145 | 146 | 147 | 148 | 149 | 150 | 151 | 152 | 153 | 154 | 155 | 156 | 157 | 158 | 159 | 160 | 161 | 162 | 163 | 164 | 165 | 166 | 167 | 168 | 169 | 170 | 171 | 172 | 173 | 174 | 175 | 176 | 177 | 178 | 179 | 180 | 181 | 182 | 183 | 184 | 185 | 186 | 187 | 188 | 189 | 190 | 191 | 192 | 193 | 194 | 195 | 196 | 197 | 198 | 199 | 200 | 201 | 202 | 203 | 204 | 205 | 206 | 207 | 208 | 209 | 210 | 211 | 212 | 213 | 214 | 215 | 216 | 217 | 218 | 219 | 220 | 221 | 222 | 223 | 224 | 225 | 226 | 227 | 228 | 229 | 230 | 231 | 232 | 233 | 234 | 235 | 236 | 237 | 238 | 239 | 240 | 241 | 242 | 243 | 244 | 245 | 246 | 247 | 248 | 249 | 250 | 251 | 252 | 253 | 254 | 255 | 256 | 257 | 258 | 259 | 260 | 261 | 262 | 263 | 264 | 265 | 266 | 267 | 268 | 269 | 270 | 271 | 272 | 273 | 274 | 275 | 276 | 277 | 278 | 279 | 280 | 281 | 282 | 283 | 284 | 285 | 286 | 287 | 288 | 289 | 290 | 291 | 292 | 293 | 294 | 295 | 296 | 297 | 298 | 299 | 300 | 301 | 302 | 303 | 304 | 305 | 306 | 307 | 308 | 309 | 310 | 311 | 312 | 313 | 314 | 315 | 316 | 317 | 318 | 319 | 320 | 321 | 322 | 323 | 324 | 325 | 326 | 327 | 328 | 329 | 330 | 331 | 332 | 333 | 334 | 335 | 336 | 337 | 338 | 339 | 340 | 341 | 342 | 343 | 344 | 345 | 346 | 347 | 348 | 349 | 350 | 351 | 352 | 353 | 354 | 355 | 356 | 357 | 358 | 359 | 360 | 361 | 362 | 363 | 364 | 365 | 366 | 367 | 368 | 369 | 370 | 371 | 372 | 373 | 374 | 375 | 376 | 377 | 378 | 379 | 380 | 381 | 382 | 383 | 384 | 385 | 386 | 387 | 388 | 389 | 390 | 391 | 392 | 393 | 394 | 395 | 396 | 397 | 398 | 399 | 400 | 401 | 402 | 403 | 404 | 405 | 406 | 407 | 408 | 409 | 410 | 411 | 412 | 413 | 414 | 415 | 416 | 417 | 418 | 419 | 420 | 421 | 422 | 423 | 424 | 425 | 426 | 427 | 428 | 429 | 430 | 431 | 432 | 433 | 434 | 435 | 436 | 437 | 438 | 439 | 440 | 441 | 442 | 443 | 444 | 445 | 446 | 447 | 448 | 449 | 450 | 451 | 452 | 453 | 454 | 455 | 456 | 457 | 458 | 459 | 460 | 461 | 462 | 463 | 464 | 465 | 466 | 467 | 468 | 469 | 470 | 471 | 472 | 473 | 474 | 475 | 476 | 477 | 478 | 479 | 480 | 481 | 482 | 483 | 484 | 485 | 486 | 487 | 488 | 489 | 490 | 491 | 492 | 493 | 494 | 495 | 496 | 497 | 498 | 499 | 500 | 501 | 502 | 503 | 504 | 505 | 506 | 507 | 508 | 509 | 510 | 511 | 512 | 513 | 514 | 515 | 516 | 517 | 518 | 519 | 520 | 521 | 522 | 523 | 524 | 525 | 526 | 527 | 528 | 529 | 530 | 531 | 532 | 533 | 534 | 535 | 536 | 537 | 538 | 539 | 540 | 541 | 542 | 543 | 544 | 545 | 546 | 547 | 548 | 549 | 550 | 551 | 552 | 553 | 554 | 555 | 556 | 557 | 558 | 559 | 560 | 561 | 562 | 563 | 564 | 565 | 566 | 567 | 568 | 569 | 570 | 571 | 572 | 573 | 574 | 575 | 576 | 577 | 578 | 579 | 580 | 581 | 582 | 583 | 584 | 585 | 586 | 587 | 588 | 589 | 590 | 591 | 592 | 593 | 594 | 595 | 596 | 597 | 598 | 599 | 600 | 601 | 602 | 603 | 604 | 605 | 606 | 607 | 608 | 609 | 610 | 611 | 612 | 613 | 614 | 615 | 616 | 617 | 618 | 619 | 620 | 621 | 622 | 623 | 624 | 625 | 626 | 627 | 628 | 629 | 630 | 631 | 632 | 633 | 634 | 635 | 636 | 637 | 638 | 639 | 640 | 641 | 642 | 643 | 644 | 645 | 646 | 647 | 648 | 649 | 650 | 651 | 652 | 653 | 654 | 655 | 656 | 657 | 658 | 659 | 660 | 661 | 662 | 663 | 664 | 665 | 666 | 667 | 668 | 669 | 670 | 671 | 672 | 673 | 674 | 675 | 676 | 677 | 678 | 679 | 680 | 681 | 682 | 683 | 684 | 685 | 686 | 687 | 688 | 689 | 690 | 691 | 692 | 693 | 694 | 695 | 696 | 697 | 698 | 699 | 700 | 701 | 702 | 703 | 704 | 705 | 706 | 707 | 708 | 709 | 710 | 711 | 712 | 713 | 714 | 715 | 716 | 717 | 718 | 719 | 720 | 721 | 722 | 723 | 724 | 725 | 726 | 727 | 728 | 729 | 730 | 731 | 732 | 733 | 734 | 735 | 736 | 737 | 738 | 739 | 740 | 741 | 742 | 743 | 744 | 745 | 746 | 747 | 748 | 749 | 750 | 751 | 752 | 753 | 754 | 755 | 756 | 757 | 758 | 759 | 760 | 761 | 762 | 763 | 764 | 765 | 766 | 767 | 768 | 769 | 770 | 771 | 772 | 773 | 774 | 775 | 776 | 777 | 778 | 779 | 780 | 781 | 782 | 783 | 784 | 785 | 786 | 787 | 788 | 789 | 790 | 791 | 792 | 793 | 794 | 795 | 796 | 797 | 798 | 799 | 800 | 801 | 802 | 803 | 804 | 805 | 806 | 807 | 808 | 809 | 810 | 811 | 812 | 813 | 814 | 815 | 816 | 817 | 818 | 819 | 820 | 821 | 822 | 823 | 824 | 825 | 826 | 827 | 828 | 829 | 830 | 831 | 832 | 833 | 834 | 835 | 836 | 837 | 838 | 839 | 840 | 841 | 842 | 843 | 844 | 845 | 846 | 847 | 848 | 849 | 850 | 851 | 852 | 853 | 854 | 855 | 856 | 857 | 858 | 859 | 860 | 861 | 862 | 863 | 864 | 865 | 866 | 867 | 868 | 869 | 870 | 871 | 872 | 873 | 874 | 875 | 876 | 877 | 878 | 879 | 880 | 881 | 882 | 883 | 884 | 885 | 886 | 887 | 888 | 889 | 890 | 891 | 892 | 893 | 894 | 895 | 896 | 897 | 898 | 899 | 900 | 901 | 902 | 903 | 904 | 905 | 906 | 907 | 908 | 909 | 910 | 911 | 912 | 913 | 914 | 915 | 916 | 917 | 918 | 919 | 920 | 921 | 922 | 923 | 924 | 925 | 926 | 927 | 928 | 929 | 930 | 931 | 932 | 933 | 934 | 935 | 936 | 937 | 938 | 939 | 940 | 941 | 942 | 943 | 944 | 945 | 946 | 947 | 948 | 949 | 950 | 951 | 952 | 953 | 954 | 955 | 956 | 957 | 958 | 959 | 960 | 961 | 962 | 963 | 964 | 965 | 966 | 967 | 968 | 969 | 970 | 971 | 972 | 973 | 974 | 975 | 976 | 977 | 978 | 979 | 980 | 981 | 982 | 983 | 984 | 985 | 986 | 987 | 988 | 989 | 990 | 991 | 992 | 993 | 994 | 995 | 996 | 997 | 998 | 999 | 1000 | 1001 | 1002 | 1003 | 1004 | 1005 | 1006 | 1007 | 1008 | 1009 | 1010 | 1011 | 1012 | 1013 | 1014 | 1015 | 1016 | 1017 | 1018 | 1019 | 1020 | 1021 | 1022 | 1023 | 1024 | 1025 | 1026 | 1027 | 1028 | 1029 | 1030 | 1031 | 1032 | 1033 | 1034 | 1035 | 1036 | 1037 | 1038 | 1039 | 1040 | 1041 | 1042 | 1043 | 1044 | 1045 | 1046 | 1047 | 1048 | 1049 | 1050 | 1051 | 1052 | 1053 | 1054 | 1055 | 1056 | 1057 | 1058 | 1059 | 1060 | 1061 | 1062 | 1063 | 1064 | 1065 | 1066 | 1067 | 1068 | 1069 | 1070 | 1071 | 1072 | 1073 | 1074 | 1075 | 1076 | 1077 | 1078 | 1079 | 1080 | 1081 | 1082 | 1083 | 1084 | 1085 | 1086 | 1087 | 1088 | 1089 | 1090 | 1091 | 1092 | 1093 | 1094 | 1095 | 1096 | 1097 | 1098 | 1099 | 1100 | 1101 | 1102 | 1103 | 1104 | 1105 | 1106 | 1107 | 1108 | 1109 | 1110 | 1111 | 1112 | 1113 | 1114 | 1115 | 1116 | 1117 | 1118 | 1119 | 1120 | 1121 | 1122 | 1123 | 1124 | 1125 | 1126 | 1127 | 1128 | 1129 | 1130 | 1131 | 1132 | 1133 | 1134 | 1135 | 1136 | 1137 | 1138 | 1139 | 1140 | 1141 | 1142 | 1143 | 1144 | 1145 | 1146 | 1147 | 1148 | 1149 | 1150 | 1151 | 1152 | 1153 | 1154 | 1155 | 1156 | 1157 | 1158 | 1159 | 1160 | 1161 | 1162 | 1163 | 1164 | 1165 | 1166 | 1167 | 1168 | 1169 | 1170 | 1171 | 1172 | 1173 | 1174 | 1175 | 1176 | 1177 | 1178 | 1179 | 1180 | 1181 | 1182 | 1183 | 1184 | 1185 | 1186 | 1187 | 1188 | 1189 | 1190 | 1191 | 1192 | 1193 | 1194 | 1195 | 1196 | 1197 | 1198 | 1199 | 1200 | 1201 | 1202 | 1203 | 1204 | 1205 | 1206 | 1207 | 1208 | 1209 | 1210 | 1211 | 1212 | 1213 | 1214 | 1215 | 1216 | 1217 | 1218 | 1219 | 1220 | 1221 | 1222 | 1223 | 1224 | 1225 | 1226 | 1227 | 1228 | 1229 | 1230 | 1231 | 1232 | 1233 | 1234 | 1235 | 1236 | 1237 | 1238 | 1239 | 1240 | 1241 | 1242 | 1243 | 1244 | 1245 | 1246 | 1247 | 1248 | 1249 | 1250 | 1251 | 1252 | 1253 | 1254 | 1255 | 1256 | 1257 | 1258 | 1259 | 1260 | 1261 | 1262 | 1263 | 1264 | 1265 | 1266 | 1267 | 1268 | 1269 | 1270 | 1271 | 1272 | 1273 | 1274 | 1275 | 1276 | 1277 | 1278 | 1279 | 1280 | 1281 | 1282 | 1283 | 1284 | 1285 | 1286 | 1287 | 1288 | 1289 | 1290 | 1291 | 1292 | 1293 | 1294 | 1295 | 1296 | 1297 | 1298 | 1299 | 1300 | 1301 | 1302 | 1303 | 1304 | 1305 | 1306 | 1307 | 1308 | 1309 | 1310 | 1311 | 1312 | 1313 | 1314 | 1315 | 1316 | 1317 | 1318 | 1319 | 1320 | 1321 | 1322 | 1323 | 1324 | 1325 | 1326 | 1327 | 1328 | 1329 | 1330 | 1331 | 1332 | 1333 | 1334 | 1335 | 1336 | 1337 | 1338 | 1339 | 1340 | 1341 | 1342 | 1343 | 1344 | 1345 | 1346 | 1347 | 1348 | 1349 | 1350 | 1351 | 1352 | 1353 | 1354 | 1355 | 1356 | 1357 | 1358 | 1359 | 1360 | 1361 | 1362 | 1363 | 1364 | 1365 | 1366 | 1367 | 1368 | 1369 | 1370 | 1371 | 1372 | 1373 | 1374 | 1375 | 1376 | 1377 | 1378 | 1379 | 1380 | 1381 | 1382 | 1383 | 1384 | 1385 | 1386 | 1387 | 1388 | 1389 | 1390 | 1391 | 1392 | 1393 | 1394 | 1395 | 1396 | 1397 | 1398 | 1399 | 1400 | 1401 | 1402 | 1403 | 1404 | 1405 | 1406 | 1407 | 1408 | 1409 | 1410 | 1411 | 1412 | 1413 | 1414 | 1415 | 1416 | 1417 | 1418 | 1419 | 1420 | 1421 | 1422 | 1423 | 1424 | 1425 | 1426 | 1427 | 1428 | 1429 | 1430 | 1431 | 1432 | 1433 | 1434 | 1435 | 1436 | 1437 | 1438 | 1439 | 1440 | 1441 | 1442 | 1443 | 1444 | 1445 | 1446 | 1447 | 1448 | 1449 | 1450 | 1451 | 1452 | 1453 | 1454 | 1455 | 1456 | 1457 | 1458 | 1459 | 1460 | 1461 | 1462 | 1463 | 1464 | 1465 | 1466 | 1467 | 1468 | 1469 | 1470 | 1471 | 1472 | 1473 | 1474 | 1475 | 1476 | 1477 | 1478 | 1479 | 1480 | 1481 | 1482 | 1483 | 1484 | 1485 | 1486 | 1487 | 1488 | 1489 | 1490 | 1491 | 1492 | 1493 | 1494 | 1495 | 1496 | 1497 | 1498 | 1499 | 1500 | 1501 | 1502 | 1503 | 1504 | 1505 | 1506 | 1507 | 1508 | 1509 | 1510 | 1511 | 1512 | 1513 | 1514 | 1515 | 1516 | 1517 | 1518 | 1519 | 1520 | 1521 | 1522 | 1523 | 1524 | 1525 | 1526 | 1527 | 1528 | 1529 | 1530 | 1531 | 1532 | 1533 | 1534 | 1535 | 1536 | 1537 | 1538 | 1539 | 1540 | 1541 | 1542 | 1543 | 1544 | 1545 | 1546 | 1547 | 1548 | 1549 | 1550 | 1551 | 1552 | 1553 | 1554 | 1555 | 1556 | 1557 | 1558 | 1559 | 1560 | 1561 | 1562 | 1563 | 1564 | 1565 | 1566 | 1567 | 1568 | 1569 | 1570 | 1571 | 1572 | 1573 | 1574 | 1575 | 1576 | 1577 | 1578 | 1579 | 1580 | 1581 | 1582 | 1583 | 1584 | 1585 | 1586 | 1587 | 1588 | 1589 | 1590 | 1591 | 1592 | 1593 | 1594 | 1595 | 1596 | 1597 | 1598 | 1599 | 1600 | 1601 | 1602 | 1603 | 1604 | 1605 | 1606 | 1607 | 1608 | 1609 | 1610 | 1611 | 1612 | 1613 | 1614 | 1615 | 1616 | 1617 | 1618 | 1619 | 1620 | 1621 | 1622 | 1623 | 1624 | 1625 | 1626 | 1627 | 1628 | 1629 | 1630 | 1631 | 1632 | 1633 | 1634 | 1635 | 1636 | 1637 | 1638 | 1639 | 1640 | 1641 | 1642 | 1643 | 1644 | 1645 | 1646 | 1647 | 1648 | 1649 | 1650 | 1651 | 1652 | 1653 | 1654 | 1655 | 1656 | 1657 | 1658 | 1659 | 1660 | 1661 | 1662 | 1663 | 1664 | 1665 | 1666 | 1667 | 1668 | 1669 | 1670 | 1671 | 1672 | 1673 | 1674 | 1675 | 1676 | 1677 | 1678 | 1679 | 1680 | 1681 | 1682 | 1683 | 1684 | 1685 | 1686 | 1687 | 1688 | 1689 | 1690 | 1691 | 1692 | 1693 | 1694 | 1695 | 1696 | 1697 | 1698 | 1699 | 1700 | 1701 | 1702 | 1703 | 1704 | 1705 | 1706 | 1707 | 1708 | 1709 | 1710 | 1711 | 1712 | 1713 | 1714 | 1715 | 1716 | 1717 | 1718 | 1719 | 1720 | 1721 | 1722 | 1723 | 1724 | 1725 | 1726 | 1727 | 1728 | 1729 | 1730 | 1731 | 1732 | 1733 | 1734 | 1735 | 1736 | 1737 | 1738 | 1739 | 1740 | 1741 | 1742 | 1743 | 1744 | 1745 | 1746 | 1747 | 1748 | 1749 | 1750 | 1751 | 1752 | 1753 | 1754 | 1755 | 1756 | 1757 | 1758 | 1759 | 1760 | 1761 | 1762 | 1763 | 1764 | 1765 | 1766 | 1767 | 1768 | 1769 | 1770 | 1771 | 1772 | 1773 | 1774 | 1775 | 1776 | 1777 | 1778 | 1779 | 1780 | 1781 | 1782 | 1783 | 1784 | 1785 | 1786 | 1787 | 1788 | 1789 | 1790 | 1791 | 1792 | 1793 | 1794 | 1795 | 1796 | 1797 | 1798 | 1799 | 1800 | 1801 | 1802 | 1803 | 1804 | 1805 | 1806 | 1807 | 1808 | 1809 | 1810 | 1811 | 1812 | 1813 | 1814 | 1815 | 1816 | 1817 | 1818 | 1819 | 1820 | 1821 | 1822 | 1823 | 1824 | 1825 | 1826 | 1827 | 1828 | 1829 | 1830 | 1831 | 1832 | 1833 | 1834 | 1835 | 1836 | 1837 | 1838 | 1839 | 1840 | 1841 | 1842 | 1843 | 1844 | 1845 | 1846 | 1847 | 1848 | 1849 | 1850 | 1851 | 1852 | 1853 | 1854 | 1855 | 1856 | 1857 | 1858 | 1859 | 1860 | 1861 | 1862 | 1863 | 1864 | 1865 | 1866 | 1867 | 1868 | 1869 | 1870 | 1871 | 1872 | 1873 | 1874 | 1875 | 1876 | 1877 | 1878 | 1879 | 1880 | 1881 | 1882 | 1883 | 1884 | 1885 | 1886 | 1887 | 1888 | 1889 | 1890 | 1891 | 1892 | 1893 | 1894 | 1895 | 1896 | 1897 | 1898 | 1899 | 1900 | 1901 | 1902 | 1903 | 1904 | 1905 | 1906 | 1907 | 1908 | 1909 | 1910 | 1911 | 1912 | 1913 | 1914 | 1915 | 1916 | 1917 | 1918 | 1919 | 1920 | 1921 | 1922 | 1923 | 1924 | 1925 | 1926 | 1927 | 1928 | 1929 | 1930 | 1931 | 1932 | 1933 | 1934 | 1935 | 1936 | 1937 | 1938 | 1939 | 1940 | 1941 | 1942 | 1943 | 1944 | 1945 | 1946 | 1947 | 1948 | 1949 | 1950 | 1951 | 1952 | 1953 | 1954 | 1955 | 1956 | 1957 | 1958 | 1959 | 1960 | 1961 | 1962 | 1963 | 1964 | 1965 | 1966 | 1967 | 1968 | 1969 | 1970 | 1971 | 1972 | 1973 | 1974 | 1975 | 1976 | 1977 | 1978 | 1979 | 1980 | 1981 | 1982 | 1983 | 1984 | 1985 | 1986 | 1987 | 1988 | 1989 | 1990 | 1991 | 1992 | 1993 | 1994 | 1995 | 1996 | 1997 | 1998 | 1999 | 2000 | 2001 | 2002 | 2003 | 2004 | 2005 | 2006 | 2007 | 2008 | 2009 | 2010 | 2011 | 2012 | 2013 | 2014 | 2015 | 2016 | 2017 | 2018 | 2019 | 2020 | 2021 | 2022 | 2023 | 2024 | 2025 | 2026 | 2027 | 2028 | 2029 | 2030 | 2031 | 2032 | 2033 | 2034 | 2035 | 2036 | 2037 | 2038 | 2039 | 2040 | 2041 | 2042 | 2043 | 2044 | 2045 | 2046 | 2047
1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 2 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 2 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 2 | 0 | 1 | 0 | 3 | 1 | 0 | 1 | 3 | 3 | 3 | 3 | 1 | 3 | 2 | 5 | 2 | 3 | 3 | 2 | 2 | 2 | 1 | 1 | 1 | 1 | 2 | 1 | 3 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 2 | 1 | 0 | 0 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 2 | 0 | 1 | 0 | 1 | 0 | 2 | 2 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 2 | 1 | 0 | 0 | 2 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 0 | 0 | 1 | 2 | 1 | 1 | 0 | 2 | 1 | 2 | 0 | 3 | 0 | 0 | 2 | 1 | 0 | 0 | 2 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 2 | 0 | 1 | 1 | 2 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 2 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 2 | 0 | 0 | 2 | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 2 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 2 | 2 | 1 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 2 | 2 | 3 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 2 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 2 | 0 | 0 | 1 | 1 | 0 | 2 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 2 | 3 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 2 | 2 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 2 | 558 | 48 | 48 | 49 | 50 | 61 | 72 | 73 | 68 | 93 | 113 | 121 | 152 | 175 | 197 | 224 | 254 | 300 | 363 | 414 | 464 | 552 | 651 | 729 | 829 | 938 | 1059 | 1201 | 1312 | 1395 | 1536 | 1743 | 1843 | 1829 | 1857 | 1888 | 1918 | 1963 | 1969 | 1932 | 1945 | 1977 | 2095 | 2148 | 2015 | 1975 | 2025 | 2037 | 2042 | 2043 | 2046 | 2030 | 2022 | 2122 | 2116 | 2035 | 2057 | 2073 | 2100 | 2052 | 2040 | 2035 | 2038 | 2073 | 2127 | 2098 | 2059 | 2029 | 2033 | 2042 | 2081 | 2067 | 2094 | 2136 | 2087 | 2190 | 2075 | 2108 | 2160 | 2120 | 2143 | 2161 | 2194 | 2205 | 2155 | 2182 | 2263 | 2174 | 2191 | 2173 | 2150 | 2149 | 2119 | 2038 | 2014 | 2029 | 2061 | 2059 | 2008 | 2009 | 2014 | 1997 | 1975 | 1964 | 1939 | 1968 | 1886 | 2118 | 1989 | 1952 | 1920 | 1923 | 1931 | 1987 | 1971 | 1929 | 1926 | 1968 | 1922 | 1769 | 1836 | 1806 | 1747 | 1772 | 1735 | 1796 | 1790 | 1788 | 1779 | 1765 | 1795 | 1797 | 1794 | 1806 | 1787 | 1754 | 1741 | 1783 | 1800 | 1835 | 1839 | 1841 | 1844 | 1813 | 1785 | 1786 | 1822 | 1857 | 1843 | 1858 | 1896 | 1879 | 1885 | 1883 | 1901 | 1914 | 1968 | 2013 | 2001 | 2041 | 1996 | 2029 | 2148 | 2197 | 2141 | 2156 | 2169 | 2267 | 2368 | 2381 | 2382 | 2379 | 2475 | 2571 | 2594 | 2622 | 2619 | 2653 | 2707 | 2709 | 2670 | 2740 | 2799 | 2804 | 2827 | 2813 | 2801 | 2772 | 2738 | 2687 | 2714 | 2726 | 2683 | 2674 | 2649 | 2582 | 2556 | 2558 | 2534 | 2488 | 2431 | 2387 | 2374 | 2359 | 2385 | 2388 | 2332 | 2305 | 2303 | 2258 | 2226 | 2213 | 2185 | 2202 | 2172 | 2111 | 2079 | 2032 | 2044 | 1964 | 1995 | 2043 | 1998 | 1946 | 1929 | 1847 | 1843 | 1876 | 1861 | 1858 | 1825 | 1849 | 1834 | 1837 | 1832 | 1770 | 1766 | 1789 | 1757 | 1767 | 1750 | 1742 | 1717 | 1725 | 1700 | 1674 | 1646 | 1627 | 1623 | 1606 | 1611 | 1602 | 1566 | 1535 | 1567 | 1597 | 1554 | 1531 | 1545 | 1578 | 1508 | 1510 | 1569 | 1543 | 1503 | 1486 | 1462 | 1477 | 1515 | 1447 | 1442 | 1431 | 1426 | 1451 | 1444 | 1414 | 1377 | 1316 | 1348 | 1350 | 1362 | 1346 | 1310 | 1380 | 1343 | 1297 | 1327 | 1315 | 1278 | 1273 | 1281 | 1304 | 1278 | 1282 | 1249 | 1232 | 1233 | 1266 | 1239 | 1222 | 1238 | 1209 | 1220 | 1226 | 1192 | 1213 | 1218 | 1184 | 1200 | 1187 | 1161 | 1214 | 1219 | 1116 | 1122 | 1104 | 1133 | 1146 | 1144 | 1122 | 1101 | 1120 | 1081 | 1072 | 1053 | 1015 | 1021 | 1057 | 994 | 996 | 1014 | 987 | 934 | 956 | 974 | 962 | 948 | 940 | 925 | 908 | 876 | 859 | 862 | 886 | 887 | 852 | 841 | 832 | 825 | 826 | 837 | 841 | 828 | 818 | 826 | 840 | 840 | 812 | 791 | 791 | 791 | 780 | 766 | 757 | 742 | 738 | 747 | 747 | 735 | 751 | 769 | 759 | 747 | 743 | 753 | 756 | 741 | 761 | 739 | 689 | 705 | 729 | 740 | 701 | 689 | 725 | 713 | 683 | 690 | 685 | 699 | 695 | 666 | 643 | 660 | 687 | 686 | 664 | 637 | 650 | 674 | 670 | 659 | 653 | 653 | 656 | 660 | 664 | 673 | 678 | 667 | 677 | 683 | 651 | 665 | 666 | 664 | 650 | 648 | 656 | 653 | 679 | 669 | 643 | 652 | 650 | 664 | 624 | 674 | 663 | 649 | 672 | 717 | 734 | 653 | 682 | 678 | 720 | 706 | 702 | 728 | 712 | 777 | 765 | 770 | 871 | 890 | 917 | 955 | 1025 | 1060 | 1081 | 1172 | 1245 | 1276 | 1382 | 1523 | 1666 | 1780 | 1914 | 2109 | 2256 | 2447 | 2703 | 2968 | 3155 | 3359 | 3773 | 4113 | 4441 | 4727 | 5164 | 5586 | 6005 | 6484 | 6809 | 7219 | 7815 | 8405 | 8929 | 9455 | 9926 | 10419 | 10945 | 11416 | 11778 | 12260 | 12851 | 13332 | 13787 | 13961 | 14136 | 14570 | 14811 | 15042 | 15240 | 15326 | 15511 | 15522 | 15457 | 15309 | 15267 | 14978 | 14510 | 14192 | 13771 | 13249 | 12858 | 12483 | 11934 | 11343 | 10870 | 10387 | 9782 | 9152 | 8579 | 8002 | 7466 | 6978 | 6423 | 5863 | 5385 | 4933 | 4534 | 4192 | 3786 | 3350 | 3046 | 2757 | 2482 | 2340 | 2161 | 1964 | 1810 | 1659 | 1545 | 1435 | 1311 | 1203 | 1149 | 1096 | 1024 | 987 | 971 | 960 | 954 | 920 | 886 | 881 | 858 | 861 | 863 | 855 | 845 | 846 | 854 | 850 | 835 | 812 | 796 | 797 | 806 | 806 | 811 | 846 | 849 | 810 | 777 | 803 | 824 | 833 | 841 | 844 | 792 | 783 | 811 | 782 | 802 | 801 | 846 | 797 | 774 | 768 | 788 | 790 | 771 | 747 | 748 | 748 | 755 | 755 | 773 | 757 | 753 | 763 | 749 | 764 | 776 | 749 | 728 | 731 | 754 | 723 | 719 | 715 | 715 | 735 | 726 | 716 | 724 | 728 | 716 | 704 | 709 | 722 | 711 | 702 | 707 | 687 | 689 | 687 | 670 | 692 | 725 | 715 | 690 | 683 | 718 | 721 | 727 | 693 | 710 | 707 | 728 | 735 | 762 | 740 | 717 | 722 | 698 | 698 | 680 | 694 | 749 | 733 | 737 | 757 | 738 | 717 | 716 | 724 | 749 | 753 | 755 | 737 | 741 | 794 | 796 | 771 | 783 | 799 | 824 | 839 | 856 | 876 | 861 | 852 | 860 | 881 | 884 | 878 | 890 | 880 | 878 | 910 | 947 | 948 | 937 | 919 | 911 | 907 | 923 | 919 | 916 | 932 | 955 | 973 | 942 | 911 | 909 | 909 | 919 | 922 | 885 | 878 | 901 | 931 | 954 | 905 | 890 | 920 | 921 | 886 | 851 | 826 | 801 | 819 | 844 | 813 | 778 | 791 | 792 | 766 | 737 | 732 | 757 | 786 | 767 | 730 | 728 | 738 | 716 | 701 | 695 | 696 | 693 | 682 | 664 | 635 | 622 | 619 | 612 | 634 | 636 | 613 | 635 | 632 | 591 | 595 | 586 | 561 | 574 | 560 | 542 | 520 | 510 | 532 | 545 | 525 | 502 | 493 | 486 | 484 | 486 | 484 | 528 | 511 | 462 | 488 | 471 | 454 | 429 | 431 | 454 | 469 | 453 | 434 | 440 | 442 | 420 | 427 | 458 | 431 | 424 | 430 | 456 | 457 | 423 | 350 | 374 | 383 | 343 | 359 | 382 | 391 | 383 | 377 | 385 | 384 | 364 | 366 | 367 | 346 | 360 | 366 | 361 | 339 | 354 | 360 | 359 | 367 | 326 | 332 | 335 | 326 | 328 | 321 | 317 | 323 | 317 | 316 | 300 | 302 | 307 | 300 | 308 | 292 | 303 | 293 | 294 | 292 | 285 | 281 | 291 | 274 | 262 | 258 | 252 | 252 | 257 | 265 | 265 | 254 | 252 | 252 | 273 | 267 | 250 | 239 | 239 | 243 | 248 | 237 | 230 | 226 | 225 | 228 | 219 | 208 | 204 | 203 | 208 | 207 | 201 | 198 | 195 | 197 | 205 | 205 | 193 | 193 | 206 | 205 | 192 | 181 | 184 | 190 | 183 | 176 | 184 | 193 | 174 | 162 | 169 | 172 | 170 | 178 | 175 | 179 | 169 | 160 | 173 | 185 | 183 | 168 | 157 | 155 | 155 | 163 | 162 | 167 | 163 | 161 | 147 | 156 | 167 | 155 | 165 | 169 | 166 | 169 | 155 | 140 | 146 | 150 | 150 | 154 | 153 | 147 | 147 | 158 | 164 | 156 | 149 | 166 | 161 | 142 | 137 | 160 | 152 | 146 | 150 | 169 | 164 | 142 | 136 | 144 | 153 | 150 | 145 | 153 | 154 | 145 | 139 | 162 | 171 | 173 | 154 | 169 | 178 | 193 | 202 | 193 | 187 | 202 | 198 | 185 | 204 | 214 | 190 | 207 | 240 | 212 | 247 | 283 | 255 | 298 | 297 | 319 | 327 | 356 | 378 | 389 | 383 | 405 | 434 | 439 | 452 | 487 | 520 | 528 | 596 | 644 | 633 | 717 | 750 | 754 | 778 | 814 | 874 | 919 | 954 | 978 | 998 | 1030 | 1075 | 1111 | 1127 | 1156 | 1204 | 1267 | 1326 | 1364 | 1354 | 1463 | 1430 | 1420 | 1433 | 1450 | 1494 | 1552 | 1575 | 1578 | 1603 | 1623 | 1625 | 1602 | 1567 | 1545 | 1519 | 1535 | 1570 | 1548 | 1558 | 1571 | 1507 | 1459 | 1449 | 1398 | 1375 | 1355 | 1303 | 1246 | 1206 | 1189 | 1170 | 1107 | 1024 | 1001 | 983 | 939 | 923 | 896 | 851 | 767 | 714 | 714 | 683 | 631 | 574 | 525 | 469 | 425 | 423 | 417 | 387 | 363 | 335 | 295 | 269 | 246 | 230 | 203 | 187 | 183 | 167 | 155 | 140 | 124 | 114 | 110 | 106 | 100 | 93 | 87 | 82 | 72 | 56 | 58 | 55 | 45 | 55 | 56 | 50 | 45 | 52 | 45 | 37 | 35 | 27 | 29 | 35 | 37 | 35 | 29 | 27 | 29 | 31 | 34 | 29 | 25 | 32 | 29 | 31 | 29 | 27 | 26 | 29 | 27 | 30 | 28 | 25 | 25 | 23 | 24 | 28 | 25 | 20 | 29 | 31 | 31 | 25 | 25 | 42 | 29 | 31 | 24 | 22 | 27 | 28 | 28 | 26 | 25 | 25 | 25 | 23 | 28 | 33 | 26 | 23 | 28 | 22 | 28 | 21 | 19 | 28 | 33 | 24 | 26 | 27 | 22 | 19 | 20 | 22 | 24 | 23 | 21 | 20 | 23 | 28 | 23 | 21 | 24 | 20 | 27 | 25 | 23 | 21 | 18 | 25 | 24 | 21 | 21 | 19 | 20 | 20 | 20 | 21 | 21 | 23 | 28 | 26 | 23 | 24 | 20 | 19 | 22 | 21 | 20 | 22 | 22 | 23 | 24 | 23 | 20 | 19 | 21 | 22 | 23 | 21 | 20 | 24 | 23 | 22 | 21 | 22 | 22 | 25 | 26 | 20 | 16 | 17 | 20 | 22 | 23 | 21 | 18 | 19 | 19 | 18 | 17 | 19 | 22 | 21 | 20 | 17 | 16 | 20 | 21 | 19 | 16 | 15 | 17 | 21 | 21 | 18 | 20 | 21 | 20 | 19 | 19 | 23 | 17 | 13 | 15 | 22 | 21 | 18 | 18 | 17 | 16 | 15 | 16 | 15 | 14 | 15 | 14 | 16 | 18 | 16 | 17 | 19 | 13 | 17 | 21 | 12 | 15 | 11 | 14 | 20 | 18 | 19 | 22 | 15 | 13 | 10 | 9 | 18 | 18 | 17 | 13 | 9 | 14 | 13 | 7 | 17 | 12 | 12 | 12 | 15 | 13 | 14 | 14 | 11 | 9 | 7 | 11 | 15 | 20 | 10 | 21 | 15 | 12 | 9 | 13 | 15 | 8 | 7 | 12 | 15 | 10 | 11 | 14 | 11 | 11 | 11 | 7 | 11 | 14 | 10 | 10 | 13 | 12 | 11 | 11 | 13 | 10 | 10 | 12 | 13 | 9 | 7 | 9 | 11 | 11 | 12 | 10 | 7 | 8 | 13 | 10 | 13 | 11 | 9 | 8 | 8 | 9 | 12 | 12 | 12 | 9 | 7 | 8 | 11 | 11 | 12 | 11 | 10 | 10 | 10 | 7 | 10 | 13 | 9 | 8 | 9 | 9 | 10 | 10 | 10 | 12 | 11 | 10 | 11 | 12 | 11 | 9 | 7 | 12 | 10 | 8 | 9 | 7 | 10 | 9 | 9 | 10 | 11 | 8 | 8 | 10 | 9 | 8 | 8 | 9 | 11 | 12 | 12 | 11 | 7 | 9 | 8 | 9 | 10 | 8 | 10 | 9 | 6 | 5 | 6 | 5 | 6 | 8 | 7 | 4 | 5 | 8 | 12 | 9 | 9 | 5 | 6 | 7 | 12 | 11 | 6 | 10 | 12 | 9 | 7 | 6 | 6 | 9 | 9 | 12 | 14 | 12 | 13 | 12 | 9 | 10 | 7 | 6 | 11 | 13 | 8 | 12 | 6 | 5 | 9 | 6 | 13 | 9 | 10 | 13 | 13 | 10 | 6 | 7 | 9 | 8 | 10 | 9 | 6 | 16 | 12 | 5 | 10 | 10 | 9 | 7 | 8 | 13 | 12 | 10 | 11 | 14 | 9 | 12 | 5 | 12 | 13 | 14 | 10 | 13 | 12 | 14 | 15 | 13 | 11 | 17 | 16 | 13 | 15 | 15 | 13 | 15 | 9 | 12 | 13 | 14 | 19 | 17 | 13 | 15 | 14 | 13 | 11 | 13 | 9 | 10 | 18 | 14 | 15 | 16 | 10 | 10 | 12 | 13 | 9 | 8 | 10 | 15 | 17 | 17 | 16 | 14 | 10 | 10 | 9 | 11 | 11 | 10 | 12 | 9 | 11 | 14 | 10 | 7 | 11 | 12 | 10 | 7 | 5 | 5 | 7 | 7 | 6 | 7 | 3 | 3 | 5 | 6 | 3 | 2 | 4 | 2 | 3 | 4 | 5 | 4 | 4 | 3 | 3 | 5 | 3 | 3 | 2 | 1 | 3 | 5 | 4 | 4 | 4 | 3 | 4 | 3 | 3 | 2 | 2 | 2 | 2 | 1 | 1 | 3 | 2 | 2 | 2 | 1 | 3 | 2 | 3 | 5 | 4 | 3 | 2 | 3 | 2 | 0 | 1 | 2 | 4 | 2 | 1 | 2 | 3 | 2 | 2 | 1 | 3 | 3 | 2 | 1 | 3 | 5 | 2 | 4 | 2 | 2 | 2 | 2 | 4 | 3 | 3 | 1 | 5 | 1 | 0 | 2 | 2 | 2 | 3 | 2 | 3 | 2 | 1 | 3 | 3 | 4 | 2 | 2 | 2 | 2 | 3 | 3 | 0 | 4 | 2 | 1 | 1 | 3 | 1 | 4 | 2 | 2 | 2 | 2 | 1 | 3 | 2 | 1 | 1 | 0 | 2 | 1 | 3 | 2 | 2 | 0 | 1 | 2 | 2 | 1 | 1 | 2 | 2 | 2 | 1 | 2 | 2 | 0 | 2 | 3 | 3 | 5 | 2 | 2 | 0 | 1 | 2 | 2 | 2 | 2 | 4 | 1 | 1 | 2 | 1 | 2 | 3 | 2 | 1 | 1 | 3 | 3 | 2 | 2 | 3 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 2 | 3 | 1 | 0 | 1 | 1 | 1 | 2 | 2 | 1 | 1 | 2 | 1 | 2 | 2 | 3 | 2 | 1 | 1 | 2 | 1 | 2 | 1 | 1 | 1 | 2 | 2 | 2 | 1 | 1 | 2 | 2 | 2 | 1 | 2 | 2 | 1 | 2 | 1 | 0 | 0 | 1 | 0 | 2 | 1 | 0 | 2 | 1 | 0 | 1 | 1 | 0 | 1 | 2 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 2 | 1 | 2 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 2 | 1 | 0 | 2 | 1 | 3 | 0 | 0 | 2 | 0 | 1 | 1 | 1 | 2 | 0 | 2 | 1 | 0 | 1 | 1 | 1 | 2 | 0 | 0 | 1 | 1 | 2 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 3 | 1 | 1 | 2 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 2 | 3 | 1 | 2 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 2 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 2 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 2 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 2 | 0 | 1 | 2 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 3 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 3 | 490 | 38 | 49 | 53 | 55 | 65 | 67 | 75 | 72 | 82 | 108 | 128 | 150 | 167 | 179 | 203 | 233 | 272 | 326 | 384 | 441 | 531 | 640 | 711 | 788 | 898 | 1019 | 1125 | 1228 | 1361 | 1490 | 1603 | 1744 | 1810 | 1828 | 1870 | 1902 | 1934 | 1964 | 1988 | 1999 | 2003 | 2033 | 2055 | 2016 | 1983 | 2035 | 2109 | 2103 | 2060 | 2040 | 2023 | 2051 | 2150 | 2163 | 2077 | 2036 | 2014 | 1971 | 1972 | 2030 | 2051 | 2074 | 2045 | 2123 | 2091 | 2014 | 2045 | 2063 | 2026 | 2007 | 2038 | 2015 | 2058 | 2057 | 2214 | 2083 | 2015 | 2091 | 2136 | 2145 | 2140 | 2159 | 2148 | 2147 | 2132 | 2214 | 2100 | 2127 | 2106 | 2093 | 2098 | 2108 | 2035 | 2010 | 1977 | 2036 | 2005 | 1947 | 1947 | 1957 | 1931 | 1909 | 1894 | 1940 | 1963 | 1890 | 2053 | 1918 | 1915 | 1942 | 1927 | 1921 | 1906 | 1958 | 1969 | 1944 | 1999 | 1870 | 1782 | 1773 | 1707 | 1794 | 1760 | 1792 | 1850 | 1761 | 1748 | 1740 | 1773 | 1757 | 1771 | 1766 | 1775 | 1776 | 1825 | 1849 | 1824 | 1801 | 1783 | 1791 | 1813 | 1822 | 1820 | 1814 | 1806 | 1795 | 1817 | 1838 | 1852 | 1876 | 1863 | 1879 | 1924 | 1987 | 1914 | 1882 | 1895 | 1942 | 2018 | 2012 | 1993 | 2025 | 2052 | 2059 | 2059 | 2116 | 2215 | 2222 | 2263 | 2337 | 2410 | 2452 | 2516 | 2512 | 2602 | 2645 | 2680 | 2674 | 2701 | 2735 | 2742 | 2777 | 2762 | 2726 | 2746 | 2770 | 2766 | 2735 | 2704 | 2690 | 2692 | 2695 | 2675 | 2628 | 2595 | 2559 | 2517 | 2486 | 2484 | 2468 | 2408 | 2397 | 2383 | 2333 | 2286 | 2287 | 2280 | 2248 | 2240 | 2257 | 2219 | 2160 | 2174 | 2143 | 2166 | 2160 | 2145 | 2132 | 2002 | 2002 | 2063 | 2037 | 1968 | 1982 | 1977 | 1972 | 1967 | 1853 | 1813 | 1842 | 1846 | 1819 | 1823 | 1798 | 1787 | 1808 | 1805 | 1776 | 1750 | 1719 | 1672 | 1676 | 1702 | 1713 | 1688 | 1644 | 1629 | 1649 | 1635 | 1626 | 1658 | 1654 | 1619 | 1563 | 1557 | 1525 | 1547 | 1517 | 1520 | 1545 | 1482 | 1458 | 1515 | 1521 | 1476 | 1516 | 1454 | 1457 | 1436 | 1464 | 1395 | 1374 | 1352 | 1403 | 1408 | 1340 | 1379 | 1310 | 1357 | 1401 | 1305 | 1323 | 1296 | 1279 | 1318 | 1352 | 1330 | 1304 | 1292 | 1279 | 1286 | 1247 | 1241 | 1233 | 1193 | 1216 | 1263 | 1260 | 1232 | 1237 | 1238 | 1221 | 1228 | 1210 | 1198 | 1210 | 1205 | 1229 | 1228 | 1157 | 1115 | 1141 | 1166 | 1166 | 1142 | 1148 | 1183 | 1154 | 1129 | 1108 | 1096 | 1076 | 1063 | 1073 | 1043 | 1040 | 987 | 975 | 1002 | 995 | 995 | 965 | 932 | 928 | 898 | 880 | 863 | 878 | 906 | 894 | 873 | 868 | 881 | 859 | 849 | 843 | 837 | 807 | 795 | 810 | 807 | 801 | 797 | 787 | 799 | 807 | 803 | 797 | 777 | 769 | 756 | 745 | 754 | 759 | 770 | 768 | 749 | 755 | 732 | 710 | 704 | 689 | 719 | 740 | 742 | 726 | 720 | 721 | 713 | 706 | 709 | 720 | 715 | 694 | 688 | 703 | 718 | 700 | 700 | 705 | 703 | 708 | 673 | 646 | 646 | 651 | 658 | 651 | 631 | 626 | 635 | 654 | 662 | 646 | 640 | 641 | 642 | 640 | 629 | 642 | 644 | 666 | 663 | 657 | 646 | 658 | 617 | 629 | 653 | 636 | 669 | 669 | 665 | 663 | 643 | 615 | 634 | 680 | 640 | 711 | 686 | 667 | 627 | 686 | 694 | 690 | 673 | 661 | 697 | 693 | 715 | 749 | 792 | 803 | 847 | 851 | 888 | 933 | 978 | 979 | 1005 | 1103 | 1222 | 1308 | 1366 | 1437 | 1549 | 1643 | 1809 | 2061 | 2222 | 2409 | 2596 | 2773 | 3020 | 3207 | 3502 | 3796 | 4024 | 4448 | 4792 | 5261 | 5697 | 6093 | 6516 | 7009 | 7433 | 8091 | 8584 | 8952 | 9696 | 9948 | 10465 | 11118 | 11505 | 11869 | 12501 | 12835 | 13205 | 13703 | 14119 | 14357 | 14570 | 14898 | 15114 | 15289 | 15391 | 15442 | 15424 | 15202 | 14993 | 14859 | 14673 | 14390 | 14057 | 13698 | 13229 | 12697 | 12153 | 11588 | 11049 | 10498 | 9921 | 9321 | 8736 | 8219 | 7661 | 7149 | 6686 | 6187 | 5713 | 5194 | 4738 | 4346 | 3961 | 3639 | 3292 | 2946 | 2645 | 2380 | 2209 | 2009 | 1861 | 1723 | 1547 | 1431 | 1298 | 1196 | 1182 | 1129 | 1065 | 1024 | 984 | 967 | 954 | 925 | 894 | 874 | 851 | 837 | 845 | 832 | 815 | 823 | 806 | 810 | 851 | 842 | 822 | 836 | 831 | 814 | 818 | 808 | 784 | 788 | 826 | 840 | 818 | 795 | 777 | 803 | 815 | 793 | 789 | 775 | 780 | 779 | 780 | 789 | 809 | 790 | 809 | 749 | 787 | 760 | 798 | 768 | 782 | 728 | 725 | 756 | 778 | 747 | 733 | 717 | 719 | 731 | 734 | 722 | 712 | 717 | 710 | 697 | 717 | 725 | 741 | 751 | 720 | 691 | 683 | 691 | 713 | 720 | 709 | 699 | 695 | 699 | 683 | 671 | 679 | 691 | 682 | 690 | 658 | 676 | 661 | 701 | 725 | 708 | 730 | 723 | 703 | 737 | 687 | 685 | 645 | 691 | 692 | 707 | 719 | 711 | 697 | 686 | 703 | 706 | 725 | 747 | 741 | 730 | 722 | 745 | 758 | 756 | 731 | 770 | 839 | 766 | 793 | 799 | 788 | 826 | 848 | 821 | 795 | 812 | 822 | 822 | 844 | 875 | 894 | 892 | 885 | 916 | 922 | 950 | 959 | 915 | 970 | 985 | 922 | 911 | 947 | 963 | 956 | 931 | 919 | 927 | 947 | 942 | 905 | 901 | 919 | 905 | 892 | 918 | 905 | 895 | 931 | 913 | 865 | 876 | 857 | 827 | 814 | 819 | 847 | 852 | 815 | 783 | 762 | 729 | 735 | 761 | 733 | 735 | 750 | 731 | 730 | 732 | 693 | 669 | 668 | 675 | 670 | 652 | 658 | 658 | 635 | 639 | 647 | 619 | 629 | 632 | 599 | 595 | 593 | 563 | 529 | 570 | 596 | 577 | 553 | 541 | 537 | 562 | 565 | 523 | 513 | 515 | 512 | 512 | 520 | 506 | 500 | 480 | 467 | 474 | 462 | 478 | 452 | 468 | 467 | 452 | 458 | 456 | 436 | 421 | 424 | 436 | 419 | 448 | 426 | 431 | 435 | 402 | 420 | 413 | 387 | 380 | 390 | 406 | 388 | 377 | 365 | 351 | 351 | 361 | 359 | 340 | 350 | 378 | 344 | 351 | 358 | 339 | 353 | 330 | 336 | 351 | 337 | 305 | 331 | 346 | 316 | 297 | 317 | 338 | 322 | 314 | 315 | 301 | 306 | 296 | 306 | 281 | 272 | 288 | 278 | 280 | 290 | 286 | 290 | 280 | 273 | 277 | 266 | 252 | 260 | 252 | 237 | 240 | 229 | 232 | 253 | 239 | 236 | 248 | 252 | 256 | 230 | 219 | 221 | 222 | 215 | 216 | 212 | 204 | 217 | 219 | 203 | 193 | 198 | 219 | 195 | 195 | 196 | 186 | 187 | 198 | 194 | 198 | 209 | 205 | 184 | 180 | 188 | 195 | 202 | 184 | 164 | 183 | 185 | 169 | 177 | 173 | 165 | 175 | 179 | 165 | 163 | 165 | 165 | 170 | 169 | 163 | 156 | 155 | 155 | 155 | 150 | 154 | 165 | 165 | 178 | 164 | 152 | 148 | 146 | 154 | 163 | 152 | 147 | 160 | 167 | 164 | 162 | 153 | 152 | 155 | 165 | 164 | 164 | 148 | 147 | 144 | 137 | 136 | 139 | 155 | 162 | 156 | 142 | 137 | 140 | 145 | 151 | 155 | 167 | 163 | 161 | 155 | 151 | 163 | 164 | 162 | 156 | 160 | 174 | 182 | 163 | 180 | 196 | 195 | 195 | 199 | 203 | 203 | 199 | 214 | 228 | 244 | 230 | 270 | 237 | 273 | 258 | 289 | 317 | 329 | 330 | 354 | 372 | 391 | 413 | 439 | 463 | 458 | 488 | 514 | 550 | 559 | 611 | 683 | 667 | 705 | 733 | 791 | 831 | 851 | 904 | 914 | 924 | 951 | 1011 | 1081 | 1114 | 1146 | 1184 | 1223 | 1185 | 1273 | 1367 | 1450 | 1380 | 1407 | 1479 | 1492 | 1459 | 1513 | 1531 | 1546 | 1535 | 1572 | 1595 | 1584 | 1611 | 1615 | 1591 | 1589 | 1573 | 1557 | 1614 | 1619 | 1499 | 1482 | 1473 | 1444 | 1395 | 1360 | 1323 | 1278 | 1242 | 1239 | 1207 | 1121 | 1050 | 1026 | 1017 | 971 | 910 | 861 | 820 | 784 | 714 | 696 | 670 | 624 | 598 | 554 | 499 | 472 | 456 | 423 | 378 | 337 | 317 | 309 | 281 | 262 | 246 | 205 | 191 | 199 | 191 | 167 | 136 | 116 | 112 | 115 | 114 | 108 | 103 | 91 | 76 | 72 | 76 | 69 | 58 | 48 | 40 | 42 | 51 | 44 | 46 | 39 | 38 | 38 | 34 | 34 | 32 | 29 | 27 | 29 | 30 | 26 | 26 | 28 | 30 | 24 | 30 | 32 | 31 | 34 | 29 | 20 | 25 | 26 | 23 | 25 | 29 | 30 | 26 | 23 | 25 | 29 | 31 | 31 | 27 | 24 | 29 | 37 | 34 | 22 | 20 | 23 | 28 | 19 | 27 | 30 | 24 | 21 | 25 | 26 | 24 | 28 | 21 | 24 | 25 | 28 | 19 | 23 | 26 | 21 | 35 | 30 | 24 | 23 | 25 | 26 | 26 | 27 | 28 | 29 | 25 | 23 | 26 | 27 | 25 | 29 | 24 | 26 | 27 | 20 | 28 | 21 | 23 | 26 | 31 | 26 | 24 | 27 | 26 | 23 | 20 | 21 | 23 | 25 | 24 | 27 | 26 | 27 | 20 | 17 | 24 | 25 | 26 | 22 | 22 | 21 | 23 | 25 | 24 | 21 | 20 | 23 | 22 | 22 | 22 | 24 | 22 | 17 | 19 | 20 | 22 | 23 | 24 | 23 | 22 | 25 | 26 | 23 | 20 | 17 | 17 | 18 | 23 | 24 | 22 | 21 | 20 | 24 | 24 | 20 | 15 | 16 | 18 | 14 | 13 | 15 | 17 | 18 | 19 | 19 | 18 | 15 | 17 | 19 | 20 | 16 | 14 | 17 | 18 | 15 | 16 | 16 | 14 | 14 | 15 | 17 | 17 | 15 | 17 | 19 | 15 | 14 | 15 | 14 | 12 | 15 | 13 | 14 | 12 | 15 | 16 | 17 | 15 | 17 | 18 | 16 | 13 | 14 | 14 | 16 | 17 | 14 | 16 | 15 | 17 | 15 | 15 | 14 | 19 | 18 | 15 | 13 | 10 | 15 | 13 | 11 | 15 | 15 | 11 | 9 | 12 | 14 | 13 | 13 | 15 | 9 | 8 | 11 | 10 | 7 | 12 | 11 | 10 | 11 | 15 | 12 | 10 | 11 | 13 | 13 | 12 | 8 | 9 | 8 | 11 | 5 | 7 | 3 | 10 | 9 | 9 | 13 | 10 | 12 | 10 | 10 | 11 | 10 | 8 | 9 | 14 | 15 | 11 | 7 | 10 | 7 | 12 | 7 | 16 | 12 | 13 | 14 | 11 | 9 | 10 | 11 | 10 | 10 | 11 | 9 | 9 | 9 | 9 | 10 | 10 | 11 | 9 | 8 | 11 | 7 | 7 | 8 | 10 | 7 | 8 | 6 | 4 | 7 | 11 | 11 | 10 | 10 | 9 | 6 | 8 | 10 | 7 | 6 | 8 | 8 | 9 | 12 | 12 | 10 | 10 | 8 | 7 | 6 | 7 | 9 | 8 | 6 | 7 | 7 | 8 | 9 | 7 | 5 | 7 | 8 | 8 | 8 | 7 | 8 | 7 | 8 | 10 | 9 | 10 | 9 | 7 | 7 | 11 | 11 | 10 | 9 | 6 | 9 | 12 | 9 | 8 | 9 | 9 | 8 | 7 | 6 | 7 | 7 | 6 | 7 | 9 | 10 | 9 | 11 | 11 | 6 | 4 | 7 | 7 | 9 | 7 | 12 | 11 | 11 | 8 | 9 | 9 | 7 | 9 | 12 | 11 | 5 | 13 | 14 | 14 | 12 | 11 | 9 | 7 | 8 | 9 | 11 | 6 | 4 | 6 | 9 | 10 | 13 | 11 | 9 | 16 | 10 | 10 | 10 | 6 | 14 | 11 | 13 | 5 | 19 | 5 | 8 | 5 | 13 | 16 | 13 | 14 | 13 | 10 | 12 | 11 | 10 | 13 | 9 | 15 | 13 | 13 | 9 | 15 | 12 | 13 | 16 | 9 | 6 | 14 | 15 | 8 | 11 | 13 | 13 | 10 | 5 | 9 | 12 | 11 | 14 | 11 | 7 | 9 | 7 | 9 | 11 | 9 | 12 | 8 | 9 | 9 | 8 | 7 | 9 | 11 | 6 | 7 | 7 | 7 | 8 | 7 | 5 | 5 | 7 | 8 | 9 | 8 | 5 | 5 | 10 | 12 | 8 | 6 | 6 | 8 | 7 | 6 | 8 | 8 | 7 | 3 | 3 | 3 | 3 | 4 | 2 | 0 | 2 | 4 | 2 | 3 | 3 | 1 | 4 | 5 | 3 | 2 | 2 | 4 | 5 | 3 | 3 | 4 | 3 | 2 | 3 | 3 | 3 | 2 | 3 | 2 | 3 | 2 | 2 | 2 | 2 | 2 | 1 | 0 | 1 | 1 | 1 | 2 | 2 | 1 | 1 | 3 | 4 | 4 | 2 | 3 | 1 | 1 | 2 | 2 | 2 | 1 | 3 | 1 | 1 | 3 | 1 | 1 | 1 | 0 | 5 | 2 | 1 | 1 | 2 | 3 | 3 | 0 | 3 | 1 | 3 | 3 | 0 | 1 | 3 | 1 | 3 | 2 | 0 | 0 | 2 | 2 | 0 | 5 | 1 | 1 | 1 | 1 | 1 | 8 | 1 | 3 | 2 | 5 | 1 | 1 | 2 | 0 | 3 | 2 | 1 | 1 | 1 | 2 | 0 | 0 | 2 | 2 | 3 | 1 | 1 | 1 | 0 | 0 | 3 | 4 | 2 | 1 | 2 | 1 | 1 | 2 | 0 | 1 | 3 | 2 | 2 | 1 | 4 | 1 | 2 | 3 | 3 | 2 | 1 | 1 | 2 | 3 | 5 | 1 | 1 | 2 | 2 | 2 | 3 | 2 | 2 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 2 | 3 | 1 | 1 | 1 | 1 | 0 | 2 | 3 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 2 | 2 | 2 | 1 | 2 | 2 | 1 | 0 | 1 | 1 | 1 | 3 | 1 | 0 | 1 | 1 | 3 | 5 | 2 | 0 | 1 | 3 | 1 | 1 | 3 | 3 | 1 | 0 | 1 | 1 | 1 | 2 | 1 | 1 | 2 | 2 | 2 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 2 | 0 | 0 | 2 | 0 | 0 | 0 | 2 | 0 | 1 | 3 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 2 | 0 | 1 | 0 | 1 | 0 | 2 | 0 | 0 | 0 | 1 | 2 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 2 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 2 | 0 | 2 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 2 | 2 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 2 | 2 | 1 | 1 | 2 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 3 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 2 | 2 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 2 | 0 | 1 | 0 | 0 | 1 | 2 | 0 | 0 | 2 | 0 | 1 | 0 | 1 | 3 | 1 | 3 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 4 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 2 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 2 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 4 | 459 | 47 | 48 | 53 | 55 | 57 | 64 | 74 | 67 | 76 | 88 | 113 | 133 | 155 | 180 | 220 | 248 | 284 | 333 | 380 | 442 | 516 | 609 | 686 | 748 | 854 | 1000 | 1130 | 1215 | 1318 | 1443 | 1619 | 1769 | 1781 | 1790 | 1841 | 1852 | 1858 | 1886 | 1919 | 1925 | 1950 | 2043 | 2117 | 2018 | 1919 | 1930 | 1970 | 1938 | 1919 | 1957 | 1947 | 1958 | 2053 | 2027 | 1993 | 1994 | 1981 | 1981 | 2023 | 2019 | 2002 | 2034 | 2064 | 2163 | 2023 | 1934 | 1966 | 1992 | 1980 | 1983 | 2015 | 2037 | 2028 | 2016 | 2125 | 2111 | 2043 | 2034 | 2023 | 2031 | 2044 | 2087 | 2111 | 2089 | 2101 | 2125 | 2110 | 2089 | 2060 | 2026 | 2025 | 2026 | 1968 | 1969 | 1956 | 2042 | 2037 | 1952 | 1894 | 1891 | 1917 | 1919 | 1928 | 1942 | 1917 | 1950 | 2080 | 1960 | 1946 | 1885 | 1862 | 1922 | 1894 | 1929 | 1889 | 1864 | 1858 | 1849 | 1715 | 1732 | 1741 | 1712 | 1748 | 1728 | 1763 | 1703 | 1718 | 1764 | 1756 | 1804 | 1752 | 1734 | 1738 | 1718 | 1760 | 1695 | 1707 | 1763 | 1795 | 1790 | 1791 | 1806 | 1816 | 1800 | 1790 | 1782 | 1780 | 1792 | 1797 | 1799 | 1798 | 1832 | 1855 | 1866 | 1851 | 1879 | 1907 | 1937 | 1970 | 1962 | 1968 | 2011 | 2038 | 2063 | 2131 | 2175 | 2168 | 2215 | 2257 | 2342 | 2388 | 2393 | 2396 | 2437 | 2504 | 2548 | 2588 | 2656 | 2680 | 2680 | 2693 | 2659 | 2695 | 2749 | 2743 | 2774 | 2715 | 2667 | 2676 | 2665 | 2663 | 2650 | 2616 | 2594 | 2583 | 2555 | 2548 | 2534 | 2475 | 2416 | 2362 | 2325 | 2320 | 2330 | 2340 | 2313 | 2271 | 2251 | 2253 | 2254 | 2202 | 2155 | 2135 | 2086 | 2062 | 2038 | 1989 | 1982 | 1990 | 2019 | 2062 | 1964 | 1905 | 1938 | 1887 | 1855 | 1840 | 1905 | 1913 | 1894 | 1848 | 1818 | 1789 | 1799 | 1828 | 1800 | 1728 | 1685 | 1686 | 1756 | 1747 | 1727 | 1719 | 1636 | 1572 | 1587 | 1600 | 1579 | 1582 | 1584 | 1576 | 1586 | 1583 | 1598 | 1550 | 1520 | 1546 | 1503 | 1461 | 1463 | 1498 | 1485 | 1495 | 1484 | 1443 | 1467 | 1441 | 1423 | 1435 | 1423 | 1388 | 1375 | 1405 | 1411 | 1358 | 1363 | 1341 | 1313 | 1309 | 1246 | 1301 | 1303 | 1298 | 1312 | 1278 | 1310 | 1296 | 1258 | 1245 | 1268 | 1252 | 1234 | 1230 | 1229 | 1221 | 1244 | 1247 | 1248 | 1240 | 1211 | 1181 | 1182 | 1202 | 1209 | 1213 | 1180 | 1178 | 1192 | 1187 | 1173 | 1116 | 1140 | 1126 | 1139 | 1102 | 1112 | 1101 | 1018 | 1023 | 1076 | 1038 | 1012 | 999 | 1045 | 1038 | 1016 | 998 | 981 | 1025 | 999 | 986 | 968 | 934 | 936 | 931 | 910 | 923 | 932 | 915 | 912 | 901 | 887 | 865 | 849 | 842 | 832 | 799 | 776 | 788 | 802 | 794 | 788 | 790 | 788 | 785 | 788 | 794 | 791 | 773 | 778 | 794 | 768 | 738 | 720 | 734 | 733 | 723 | 732 | 712 | 716 | 723 | 720 | 723 | 714 | 721 | 715 | 725 | 725 | 714 | 742 | 717 | 689 | 688 | 690 | 695 | 697 | 700 | 690 | 680 | 664 | 673 | 675 | 648 | 663 | 679 | 673 | 659 | 660 | 655 | 638 | 626 | 634 | 650 | 661 | 660 | 648 | 660 | 649 | 623 | 625 | 639 | 646 | 625 | 605 | 614 | 621 | 614 | 642 | 645 | 639 | 619 | 624 | 627 | 613 | 598 | 602 | 621 | 626 | 596 | 642 | 656 | 657 | 656 | 647 | 667 | 642 | 657 | 673 | 705 | 706 | 701 | 724 | 753 | 809 | 800 | 824 | 867 | 895 | 952 | 968 | 1034 | 1123 | 1186 | 1256 | 1327 | 1428 | 1522 | 1652 | 1845 | 1953 | 2093 | 2294 | 2446 | 2671 | 2924 | 3176 | 3399 | 3727 | 4162 | 4500 | 4830 | 5191 | 5696 | 6165 | 6595 | 6907 | 7405 | 7890 | 8299 | 8863 | 9575 | 9840 | 10399 | 10777 | 11285 | 11888 | 12461 | 12905 | 13204 | 13616 | 13769 | 14282 | 14421 | 14698 | 14826 | 14968 | 15038 | 15141 | 15228 | 15042 | 14922 | 14622 | 14313 | 14005 | 13765 | 13460 | 13000 | 12451 | 11915 | 11463 | 10910 | 10272 | 9725 | 9216 | 8705 | 8157 | 7449 | 6850 | 6452 | 5960 | 5468 | 5001 | 4576 | 4236 | 3904 | 3506 | 3147 | 2830 | 2596 | 2360 | 2084 | 1953 | 1837 | 1669 | 1562 | 1411 | 1302 | 1240 | 1172 | 1134 | 1051 | 979 | 941 | 919 | 893 | 872 | 866 | 854 | 833 | 830 | 815 | 804 | 817 | 819 | 815 | 818 | 826 | 841 | 844 | 852 | 861 | 839 | 813 | 814 | 821 | 823 | 815 | 804 | 800 | 818 | 831 | 801 | 785 | 793 | 790 | 783 | 814 | 820 | 802 | 791 | 753 | 784 | 805 | 744 | 772 | 755 | 766 | 777 | 766 | 747 | 797 | 786 | 734 | 711 | 728 | 754 | 730 | 748 | 749 | 716 | 723 | 736 | 737 | 706 | 703 | 711 | 705 | 718 | 742 | 724 | 690 | 694 | 718 | 715 | 687 | 681 | 671 | 675 | 689 | 692 | 716 | 692 | 677 | 692 | 690 | 650 | 696 | 733 | 737 | 697 | 704 | 721 | 690 | 722 | 693 | 695 | 742 | 631 | 690 | 700 | 702 | 710 | 714 | 694 | 689 | 698 | 712 | 711 | 709 | 731 | 747 | 770 | 747 | 751 | 778 | 767 | 822 | 793 | 799 | 809 | 777 | 792 | 815 | 807 | 821 | 855 | 860 | 860 | 868 | 857 | 846 | 861 | 915 | 899 | 895 | 925 | 882 | 913 | 946 | 939 | 932 | 928 | 888 | 896 | 922 | 929 | 929 | 937 | 942 | 944 | 925 | 884 | 891 | 901 | 901 | 914 | 916 | 836 | 819 | 844 | 867 | 864 | 857 | 859 | 850 | 825 | 797 | 794 | 794 | 788 | 793 | 803 | 792 | 763 | 768 | 764 | 726 | 688 | 676 | 683 | 697 | 708 | 677 | 651 | 666 | 650 | 633 | 644 | 638 | 619 | 610 | 626 | 617 | 613 | 612 | 613 | 601 | 577 | 561 | 562 | 557 | 546 | 550 | 550 | 546 | 541 | 540 | 540 | 515 | 495 | 501 | 502 | 498 | 512 | 507 | 491 | 503 | 479 | 489 | 489 | 474 | 464 | 443 | 427 | 425 | 430 | 442 | 439 | 425 | 436 | 442 | 416 | 412 | 432 | 415 | 414 | 401 | 403 | 397 | 394 | 364 | 383 | 402 | 398 | 385 | 367 | 348 | 354 | 383 | 379 | 387 | 356 | 346 | 372 | 367 | 350 | 333 | 348 | 372 | 325 | 313 | 325 | 334 | 354 | 339 | 322 | 320 | 308 | 306 | 316 | 336 | 303 | 280 | 299 | 293 | 267 | 290 | 283 | 293 | 270 | 265 | 269 | 261 | 277 | 291 | 280 | 268 | 266 | 261 | 249 | 239 | 248 | 251 | 253 | 259 | 258 | 244 | 238 | 245 | 229 | 215 | 216 | 220 | 211 | 204 | 209 | 209 | 207 | 208 | 207 | 205 | 207 | 213 | 221 | 205 | 178 | 186 | 185 | 177 | 188 | 194 | 193 | 197 | 196 | 189 | 183 | 178 | 170 | 165 | 169 | 177 | 179 | 178 | 182 | 182 | 179 | 175 | 169 | 159 | 156 | 153 | 160 | 160 | 159 | 156 | 153 | 157 | 158 | 159 | 155 | 161 | 179 | 168 | 152 | 140 | 145 | 155 | 145 | 145 | 151 | 157 | 163 | 161 | 156 | 158 | 161 | 154 | 151 | 157 | 154 | 146 | 149 | 151 | 144 | 142 | 157 | 155 | 151 | 145 | 145 | 142 | 137 | 137 | 149 | 166 | 166 | 160 | 154 | 145 | 135 | 139 | 154 | 157 | 158 | 156 | 144 | 158 | 169 | 165 | 162 | 178 | 175 | 184 | 192 | 195 | 202 | 197 | 198 | 201 | 223 | 205 | 232 | 270 | 273 | 280 | 260 | 275 | 271 | 297 | 324 | 334 | 346 | 368 | 378 | 396 | 440 | 460 | 472 | 539 | 542 | 569 | 515 | 603 | 680 | 691 | 717 | 750 | 815 | 856 | 854 | 891 | 923 | 959 | 995 | 1040 | 1082 | 1108 | 1159 | 1166 | 1185 | 1265 | 1354 | 1336 | 1459 | 1406 | 1467 | 1450 | 1476 | 1465 | 1487 | 1522 | 1513 | 1544 | 1588 | 1573 | 1568 | 1602 | 1592 | 1556 | 1558 | 1571 | 1542 | 1531 | 1611 | 1520 | 1464 | 1465 | 1438 | 1370 | 1344 | 1328 | 1297 | 1226 | 1182 | 1154 | 1121 | 1105 | 1062 | 994 | 962 | 925 | 851 | 824 | 842 | 748 | 694 | 663 | 616 | 576 | 522 | 497 | 485 | 467 | 423 | 377 | 357 | 341 | 314 | 292 | 267 | 242 | 219 | 199 | 175 | 153 | 149 | 140 | 123 | 115 | 107 | 97 | 93 | 91 | 87 | 80 | 71 | 66 | 60 | 55 | 52 | 51 | 45 | 36 | 39 | 38 | 29 | 28 | 34 | 34 | 39 | 43 | 39 | 37 | 38 | 34 | 27 | 21 | 26 | 30 | 26 | 27 | 24 | 21 | 27 | 36 | 30 | 25 | 29 | 27 | 24 | 25 | 28 | 31 | 32 | 31 | 28 | 32 | 31 | 24 | 22 | 19 | 23 | 26 | 22 | 24 | 29 | 24 | 21 | 24 | 26 | 22 | 19 | 22 | 22 | 26 | 29 | 28 | 25 | 24 | 23 | 25 | 18 | 18 | 23 | 28 | 29 | 24 | 18 | 18 | 26 | 26 | 23 | 23 | 23 | 26 | 27 | 28 | 27 | 20 | 21 | 26 | 21 | 25 | 28 | 27 | 30 | 27 | 28 | 25 | 22 | 18 | 20 | 25 | 21 | 17 | 22 | 23 | 21 | 18 | 20 | 30 | 20 | 19 | 17 | 21 | 18 | 18 | 20 | 22 | 24 | 25 | 22 | 19 | 18 | 22 | 25 | 18 | 14 | 19 | 20 | 22 | 26 | 21 | 18 | 18 | 21 | 17 | 23 | 23 | 22 | 23 | 20 | 17 | 20 | 24 | 21 | 20 | 21 | 18 | 18 | 21 | 21 | 17 | 21 | 22 | 18 | 20 | 21 | 19 | 15 | 18 | 21 | 20 | 18 | 18 | 19 | 16 | 16 | 18 | 16 | 14 | 17 | 23 | 19 | 16 | 21 | 19 | 19 | 18 | 16 | 14 | 15 | 18 | 19 | 20 | 18 | 16 | 17 | 17 | 16 | 15 | 12 | 11 | 13 | 13 | 12 | 11 | 11 | 14 | 16 | 16 | 15 | 13 | 9 | 9 | 11 | 11 | 9 | 13 | 14 | 14 | 16 | 17 | 13 | 14 | 10 | 11 | 10 | 12 | 15 | 14 | 15 | 13 | 11 | 11 | 13 | 14 | 13 | 16 | 14 | 10 | 7 | 8 | 13 | 12 | 10 | 11 | 10 | 14 | 17 | 13 | 12 | 13 | 10 | 13 | 7 | 11 | 11 | 12 | 8 | 13 | 10 | 8 | 9 | 7 | 7 | 9 | 8 | 11 | 10 | 8 | 9 | 9 | 9 | 7 | 7 | 11 | 10 | 7 | 10 | 14 | 13 | 7 | 4 | 8 | 13 | 10 | 10 | 9 | 9 | 8 | 7 | 8 | 5 | 5 | 9 | 12 | 12 | 8 | 9 | 10 | 6 | 11 | 5 | 9 | 11 | 9 | 7 | 5 | 4 | 4 | 6 | 10 | 11 | 6 | 8 | 9 | 7 | 11 | 8 | 10 | 11 | 10 | 9 | 12 | 5 | 4 | 6 | 10 | 11 | 9 | 8 | 8 | 7 | 7 | 7 | 9 | 9 | 7 | 11 | 11 | 10 | 10 | 8 | 9 | 9 | 8 | 6 | 6 | 8 | 10 | 11 | 14 | 13 | 11 | 7 | 5 | 8 | 10 | 8 | 8 | 11 | 10 | 8 | 11 | 8 | 6 | 6 | 5 | 5 | 5 | 5 | 6 | 6 | 7 | 8 | 10 | 7 | 7 | 7 | 7 | 10 | 10 | 11 | 8 | 8 | 6 | 8 | 10 | 9 | 8 | 11 | 9 | 9 | 8 | 6 | 10 | 14 | 10 | 9 | 11 | 8 | 8 | 7 | 10 | 12 | 10 | 12 | 13 | 10 | 12 | 11 | 11 | 9 | 13 | 16 | 12 | 13 | 10 | 9 | 9 | 8 | 12 | 10 | 7 | 8 | 12 | 13 | 8 | 13 | 16 | 10 | 10 | 3 | 17 | 17 | 14 | 16 | 22 | 14 | 5 | 8 | 12 | 22 | 10 | 13 | 11 | 18 | 8 | 8 | 14 | 10 | 13 | 7 | 9 | 9 | 20 | 16 | 8 | 17 | 14 | 15 | 15 | 9 | 10 | 9 | 12 | 10 | 11 | 12 | 8 | 9 | 7 | 7 | 11 | 5 | 12 | 6 | 6 | 4 | 7 | 7 | 13 | 4 | 6 | 7 | 6 | 10 | 10 | 9 | 8 | 6 | 7 | 7 | 6 | 5 | 5 | 6 | 4 | 3 | 3 | 6 | 5 | 6 | 5 | 4 | 3 | 5 | 5 | 1 | 2 | 2 | 2 | 3 | 4 | 2 | 3 | 2 | 1 | 2 | 2 | 2 | 2 | 2 | 4 | 4 | 2 | 3 | 2 | 2 | 2 | 2 | 2 | 4 | 4 | 2 | 2 | 1 | 0 | 1 | 3 | 3 | 1 | 1 | 2 | 1 | 2 | 3 | 2 | 2 | 1 | 2 | 2 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 3 | 3 | 4 | 1 | 1 | 0 | 1 | 1 | 2 | 1 | 1 | 3 | 2 | 0 | 2 | 4 | 2 | 2 | 4 | 2 | 2 | 3 | 2 | 0 | 1 | 1 | 2 | 3 | 0 | 2 | 1 | 4 | 3 | 1 | 1 | 2 | 1 | 4 | 3 | 3 | 1 | 1 | 0 | 0 | 3 | 0 | 2 | 2 | 2 | 2 | 3 | 2 | 2 | 2 | 3 | 1 | 1 | 4 | 2 | 1 | 4 | 2 | 2 | 5 | 1 | 1 | 3 | 2 | 0 | 0 | 3 | 1 | 2 | 1 | 1 | 0 | 1 | 1 | 2 | 1 | 0 | 2 | 2 | 1 | 3 | 1 | 0 | 0 | 2 | 1 | 2 | 2 | 2 | 3 | 1 | 0 | 2 | 1 | 1 | 2 | 1 | 0 | 1 | 2 | 1 | 1 | 1 | 2 | 2 | 0 | 1 | 3 | 1 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 1 | 2 | 1 | 1 | 2 | 4 | 2 | 2 | 2 | 0 | 1 | 2 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 2 | 1 | 1 | 0 | 2 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 3 | 1 | 0 | 1 | 0 | 1 | 0 | 2 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 4 | 0 | 2 | 1 | 2 | 0 | 0 | 1 | 0 | 3 | 1 | 0 | 2 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 2 | 0 | 0 | 1 | 0 | 1 | 1 | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 2 | 2 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 2 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 2 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 2 | 0 | 1 | 0 | 1 | 1 | 1 | 2 | 0 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 5 | 481 | 48 | 46 | 51 | 59 | 60 | 68 | 71 | 62 | 77 | 96 | 123 | 145 | 164 | 186 | 210 | 244 | 289 | 343 | 394 | 461 | 559 | 653 | 722 | 795 | 909 | 1033 | 1163 | 1276 | 1381 | 1488 | 1619 | 1743 | 1768 | 1780 | 1833 | 1882 | 1920 | 1927 | 1935 | 1976 | 2029 | 2058 | 2017 | 1939 | 1963 | 1999 | 2010 | 1998 | 1993 | 2014 | 1977 | 1960 | 2092 | 2063 | 2009 | 1993 | 1999 | 2007 | 2033 | 2010 | 1970 | 1990 | 2005 | 2072 | 1995 | 1985 | 2001 | 2000 | 1972 | 1985 | 2007 | 2022 | 1984 | 2027 | 2100 | 2079 | 2047 | 2043 | 2110 | 2130 | 2106 | 2081 | 2047 | 2085 | 2105 | 2133 | 2063 | 2024 | 1988 | 1967 | 1993 | 2039 | 2040 | 1980 | 2022 | 1973 | 1972 | 1946 | 1958 | 1914 | 1925 | 1955 | 1939 | 1895 | 1830 | 1917 | 2073 | 1975 | 1922 | 1964 | 1946 | 1912 | 1866 | 1903 | 1909 | 1912 | 1885 | 1901 | 1714 | 1750 | 1661 | 1791 | 1723 | 1759 | 1708 | 1755 | 1755 | 1816 | 1829 | 1759 | 1745 | 1718 | 1722 | 1731 | 1808 | 1791 | 1795 | 1761 | 1807 | 1819 | 1780 | 1842 | 1856 | 1780 | 1740 | 1762 | 1800 | 1809 | 1805 | 1805 | 1827 | 1845 | 1823 | 1858 | 1903 | 1929 | 1869 | 1899 | 1944 | 2001 | 2023 | 2036 | 2038 | 2097 | 2158 | 2169 | 2217 | 2257 | 2278 | 2336 | 2429 | 2478 | 2459 | 2495 | 2595 | 2586 | 2604 | 2677 | 2692 | 2670 | 2664 | 2713 | 2756 | 2750 | 2696 | 2685 | 2673 | 2686 | 2693 | 2679 | 2664 | 2638 | 2603 | 2575 | 2543 | 2542 | 2508 | 2470 | 2486 | 2442 | 2367 | 2376 | 2374 | 2331 | 2313 | 2273 | 2242 | 2286 | 2285 | 2218 | 2198 | 2154 | 2114 | 2098 | 2103 | 2084 | 2026 | 2028 | 2036 | 2032 | 2050 | 1988 | 1923 | 1887 | 1896 | 1904 | 1840 | 1809 | 1838 | 1849 | 1821 | 1783 | 1760 | 1770 | 1764 | 1748 | 1717 | 1707 | 1695 | 1726 | 1709 | 1676 | 1650 | 1654 | 1665 | 1654 | 1650 | 1620 | 1615 | 1618 | 1584 | 1583 | 1571 | 1559 | 1539 | 1529 | 1534 | 1498 | 1460 | 1503 | 1523 | 1483 | 1428 | 1385 | 1385 | 1379 | 1377 | 1406 | 1442 | 1464 | 1388 | 1391 | 1376 | 1319 | 1298 | 1334 | 1366 | 1368 | 1366 | 1309 | 1273 | 1281 | 1297 | 1265 | 1276 | 1265 | 1269 | 1270 | 1271 | 1227 | 1231 | 1202 | 1230 | 1245 | 1255 | 1208 | 1201 | 1219 | 1192 | 1181 | 1182 | 1175 | 1155 | 1154 | 1157 | 1136 | 1133 | 1147 | 1174 | 1145 | 1128 | 1146 | 1109 | 1025 | 1057 | 1097 | 1099 | 1076 | 1084 | 1090 | 1075 | 1030 | 1014 | 968 | 1025 | 1058 | 997 | 1010 | 1006 | 954 | 915 | 913 | 938 | 923 | 884 | 907 | 922 | 933 | 874 | 855 | 832 | 811 | 828 | 824 | 828 | 815 | 823 | 826 | 807 | 791 | 794 | 801 | 793 | 784 | 785 | 767 | 755 | 779 | 773 | 772 | 788 | 773 | 746 | 747 | 759 | 761 | 754 | 758 | 750 | 722 | 739 | 734 | 732 | 742 | 726 | 717 | 743 | 729 | 698 | 682 | 689 | 721 | 712 | 698 | 694 | 673 | 658 | 672 | 678 | 675 | 689 | 710 | 706 | 677 | 663 | 670 | 673 | 670 | 664 | 653 | 651 | 654 | 652 | 657 | 658 | 652 | 663 | 661 | 644 | 640 | 624 | 625 | 645 | 668 | 662 | 617 | 634 | 639 | 639 | 625 | 628 | 625 | 636 | 645 | 633 | 653 | 642 | 638 | 682 | 691 | 685 | 670 | 615 | 629 | 681 | 635 | 664 | 667 | 682 | 694 | 749 | 750 | 792 | 806 | 834 | 886 | 907 | 918 | 943 | 990 | 1037 | 1150 | 1250 | 1278 | 1342 | 1428 | 1553 | 1717 | 1877 | 2081 | 2291 | 2432 | 2555 | 2831 | 2995 | 3305 | 3647 | 3859 | 4167 | 4516 | 4901 | 5263 | 5616 | 6205 | 6556 | 7070 | 7538 | 8053 | 8555 | 9116 | 9674 | 10385 | 10668 | 11094 | 11559 | 12009 | 12499 | 12903 | 13389 | 13689 | 13940 | 14227 | 14550 | 14739 | 14810 | 15140 | 15235 | 15165 | 15031 | 14930 | 14839 | 14512 | 14212 | 13970 | 13614 | 13032 | 12545 | 12201 | 11670 | 11134 | 10665 | 10041 | 9370 | 8837 | 8395 | 7881 | 7278 | 6680 | 6210 | 5663 | 5246 | 4797 | 4424 | 4048 | 3670 | 3311 | 2990 | 2750 | 2520 | 2279 | 2049 | 1865 | 1741 | 1590 | 1458 | 1367 | 1309 | 1207 | 1106 | 1068 | 987 | 971 | 976 | 923 | 872 | 869 | 868 | 859 | 869 | 877 | 896 | 876 | 818 | 805 | 818 | 821 | 812 | 809 | 802 | 810 | 840 | 836 | 820 | 803 | 778 | 766 | 775 | 800 | 804 | 788 | 778 | 772 | 768 | 783 | 795 | 791 | 784 | 777 | 796 | 800 | 774 | 755 | 778 | 773 | 821 | 781 | 783 | 766 | 742 | 731 | 725 | 745 | 767 | 743 | 696 | 719 | 721 | 685 | 699 | 712 | 691 | 693 | 709 | 719 | 723 | 725 | 719 | 692 | 689 | 671 | 661 | 693 | 715 | 707 | 672 | 679 | 702 | 679 | 692 | 685 | 717 | 695 | 682 | 737 | 712 | 702 | 726 | 677 | 651 | 701 | 700 | 632 | 692 | 672 | 711 | 739 | 678 | 715 | 675 | 697 | 732 | 715 | 691 | 704 | 727 | 736 | 703 | 701 | 727 | 710 | 702 | 705 | 753 | 738 | 752 | 749 | 809 | 734 | 772 | 780 | 792 | 796 | 812 | 843 | 839 | 832 | 862 | 874 | 856 | 867 | 898 | 900 | 859 | 893 | 899 | 902 | 939 | 898 | 925 | 920 | 897 | 909 | 922 | 916 | 925 | 937 | 925 | 891 | 880 | 901 | 921 | 945 | 925 | 901 | 876 | 882 | 921 | 874 | 899 | 855 | 840 | 835 | 824 | 807 | 788 | 789 | 801 | 787 | 752 | 750 | 790 | 793 | 749 | 716 | 724 | 766 | 749 | 709 | 675 | 668 | 685 | 688 | 685 | 662 | 638 | 625 | 613 | 610 | 620 | 643 | 638 | 606 | 593 | 611 | 589 | 582 | 597 | 554 | 562 | 572 | 540 | 512 | 504 | 516 | 529 | 537 | 529 | 516 | 522 | 500 | 474 | 476 | 475 | 479 | 497 | 480 | 474 | 446 | 438 | 453 | 467 | 459 | 439 | 431 | 429 | 431 | 434 | 430 | 423 | 427 | 449 | 439 | 396 | 399 | 404 | 408 | 398 | 422 | 403 | 404 | 399 | 419 | 397 | 373 | 355 | 361 | 383 | 375 | 366 | 362 | 362 | 333 | 317 | 362 | 323 | 347 | 366 | 354 | 327 | 334 | 339 | 331 | 336 | 318 | 304 | 304 | 320 | 331 | 310 | 309 | 298 | 285 | 295 | 267 | 340 | 302 | 283 | 286 | 281 | 279 | 286 | 260 | 244 | 246 | 255 | 265 | 260 | 244 | 241 | 247 | 261 | 258 | 233 | 243 | 230 | 256 | 238 | 242 | 235 | 224 | 204 | 206 | 203 | 210 | 215 | 220 | 218 | 203 | 203 | 211 | 217 | 204 | 202 | 185 | 198 | 203 | 201 | 184 | 205 | 200 | 188 | 181 | 182 | 187 | 180 | 175 | 182 | 180 | 176 | 181 | 172 | 173 | 170 | 166 | 172 | 183 | 179 | 164 | 173 | 168 | 163 | 167 | 173 | 177 | 180 | 170 | 155 | 157 | 161 | 150 | 144 | 151 | 152 | 156 | 161 | 151 | 149 | 160 | 160 | 147 | 142 | 152 | 158 | 157 | 154 | 150 | 143 | 145 | 143 | 142 | 146 | 140 | 145 | 153 | 162 | 166 | 149 | 141 | 141 | 137 | 149 | 157 | 144 | 137 | 145 | 153 | 158 | 158 | 153 | 148 | 146 | 157 | 165 | 156 | 147 | 163 | 164 | 171 | 160 | 164 | 172 | 183 | 181 | 171 | 177 | 187 | 194 | 202 | 203 | 208 | 213 | 222 | 244 | 278 | 282 | 285 | 286 | 314 | 330 | 311 | 348 | 370 | 384 | 402 | 424 | 455 | 469 | 489 | 513 | 533 | 553 | 596 | 613 | 703 | 691 | 700 | 759 | 812 | 842 | 875 | 910 | 946 | 957 | 990 | 1050 | 1077 | 1094 | 1166 | 1225 | 1242 | 1221 | 1291 | 1317 | 1382 | 1354 | 1436 | 1500 | 1506 | 1451 | 1522 | 1534 | 1535 | 1514 | 1540 | 1553 | 1530 | 1524 | 1506 | 1538 | 1496 | 1550 | 1546 | 1599 | 1584 | 1546 | 1427 | 1468 | 1408 | 1326 | 1312 | 1278 | 1250 | 1199 | 1164 | 1128 | 1055 | 1016 | 1015 | 971 | 917 | 873 | 814 | 763 | 792 | 692 | 634 | 612 | 576 | 538 | 508 | 478 | 460 | 418 | 395 | 373 | 326 | 304 | 298 | 270 | 244 | 225 | 205 | 191 | 181 | 170 | 164 | 140 | 120 | 113 | 104 | 100 | 91 | 81 | 76 | 70 | 65 | 66 | 62 | 56 | 47 | 41 | 44 | 44 | 46 | 44 | 37 | 37 | 39 | 37 | 36 | 30 | 24 | 26 | 29 | 30 | 33 | 31 | 31 | 31 | 31 | 25 | 28 | 33 | 29 | 28 | 32 | 29 | 26 | 26 | 30 | 30 | 27 | 27 | 24 | 22 | 24 | 23 | 24 | 25 | 29 | 23 | 24 | 29 | 38 | 31 | 24 | 24 | 25 | 24 | 21 | 22 | 21 | 22 | 28 | 29 | 22 | 22 | 20 | 22 | 26 | 25 | 30 | 29 | 33 | 27 | 21 | 23 | 31 | 24 | 20 | 22 | 24 | 28 | 28 | 25 | 24 | 20 | 18 | 23 | 21 | 20 | 21 | 25 | 23 | 30 | 30 | 27 | 21 | 23 | 26 | 27 | 23 | 19 | 16 | 16 | 19 | 25 | 17 | 26 | 24 | 15 | 19 | 23 | 25 | 20 | 23 | 21 | 20 | 23 | 19 | 17 | 13 | 13 | 18 | 22 | 23 | 16 | 14 | 16 | 22 | 22 | 19 | 17 | 22 | 19 | 18 | 20 | 19 | 19 | 17 | 18 | 22 | 20 | 21 | 21 | 17 | 19 | 20 | 22 | 24 | 20 | 22 | 26 | 17 | 18 | 19 | 21 | 18 | 22 | 21 | 19 | 19 | 18 | 21 | 20 | 15 | 16 | 16 | 17 | 19 | 22 | 19 | 20 | 20 | 16 | 14 | 14 | 14 | 14 | 14 | 14 | 15 | 12 | 14 | 17 | 15 | 15 | 15 | 13 | 13 | 13 | 14 | 18 | 20 | 16 | 12 | 15 | 16 | 15 | 17 | 17 | 14 | 14 | 16 | 15 | 15 | 14 | 13 | 11 | 13 | 13 | 13 | 12 | 10 | 12 | 12 | 12 | 12 | 11 | 12 | 11 | 12 | 11 | 12 | 13 | 13 | 14 | 11 | 10 | 7 | 8 | 14 | 15 | 14 | 13 | 12 | 10 | 12 | 11 | 9 | 12 | 14 | 12 | 10 | 10 | 12 | 11 | 11 | 12 | 12 | 16 | 13 | 12 | 15 | 8 | 9 | 8 | 12 | 14 | 9 | 7 | 11 | 15 | 12 | 10 | 8 | 6 | 13 | 8 | 10 | 12 | 6 | 10 | 11 | 8 | 9 | 12 | 9 | 10 | 12 | 10 | 9 | 9 | 11 | 13 | 9 | 8 | 13 | 8 | 8 | 7 | 7 | 9 | 13 | 8 | 8 | 13 | 12 | 8 | 5 | 5 | 7 | 6 | 8 | 10 | 8 | 7 | 12 | 6 | 12 | 12 | 7 | 6 | 9 | 12 | 7 | 7 | 6 | 8 | 11 | 10 | 7 | 6 | 7 | 6 | 3 | 7 | 12 | 9 | 8 | 8 | 6 | 8 | 12 | 10 | 10 | 11 | 11 | 9 | 9 | 8 | 7 | 9 | 12 | 11 | 6 | 5 | 8 | 9 | 6 | 8 | 6 | 7 | 9 | 9 | 8 | 8 | 9 | 8 | 6 | 6 | 7 | 8 | 9 | 8 | 8 | 10 | 8 | 10 | 9 | 7 | 7 | 10 | 9 | 5 | 8 | 5 | 9 | 9 | 7 | 6 | 6 | 12 | 9 | 5 | 10 | 8 | 7 | 10 | 11 | 9 | 11 | 12 | 11 | 11 | 8 | 7 | 11 | 14 | 13 | 13 | 17 | 14 | 14 | 11 | 8 | 14 | 9 | 11 | 10 | 11 | 15 | 10 | 9 | 12 | 15 | 10 | 9 | 8 | 12 | 14 | 11 | 15 | 9 | 9 | 11 | 13 | 16 | 10 | 12 | 15 | 12 | 10 | 14 | 10 | 11 | 14 | 16 | 12 | 15 | 11 | 11 | 11 | 12 | 9 | 9 | 12 | 10 | 10 | 5 | 7 | 10 | 10 | 10 | 5 | 9 | 8 | 14 | 15 | 5 | 8 | 5 | 12 | 9 | 9 | 9 | 6 | 9 | 5 | 7 | 11 | 10 | 9 | 8 | 6 | 3 | 4 | 6 | 6 | 7 | 4 | 7 | 6 | 4 | 2 | 0 | 1 | 3 | 4 | 3 | 1 | 2 | 2 | 4 | 4 | 3 | 2 | 2 | 3 | 5 | 3 | 2 | 2 | 3 | 3 | 4 | 4 | 2 | 5 | 5 | 1 | 1 | 2 | 1 | 1 | 3 | 4 | 3 | 3 | 4 | 3 | 2 | 2 | 1 | 1 | 0 | 1 | 3 | 3 | 1 | 2 | 3 | 2 | 2 | 2 | 4 | 2 | 2 | 2 | 1 | 1 | 3 | 2 | 2 | 1 | 1 | 1 | 2 | 1 | 2 | 4 | 2 | 2 | 1 | 1 | 5 | 3 | 2 | 3 | 2 | 0 | 1 | 1 | 1 | 3 | 5 | 3 | 1 | 1 | 1 | 2 | 2 | 3 | 2 | 2 | 1 | 4 | 3 | 1 | 0 | 1 | 1 | 1 | 1 | 2 | 3 | 1 | 2 | 4 | 2 | 0 | 3 | 0 | 1 | 3 | 3 | 2 | 0 | 3 | 3 | 1 | 2 | 4 | 3 | 0 | 0 | 2 | 2 | 2 | 2 | 0 | 3 | 0 | 1 | 0 | 2 | 3 | 3 | 0 | 1 | 4 | 0 | 5 | 0 | 2 | 2 | 4 | 1 | 3 | 2 | 1 | 2 | 0 | 1 | 1 | 1 | 4 | 0 | 4 | 2 | 2 | 1 | 1 | 0 | 2 | 0 | 2 | 0 | 1 | 3 | 2 | 1 | 2 | 0 | 0 | 1 | 0 | 2 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 2 | 1 | 1 | 1 | 2 | 2 | 1 | 1 | 0 | 1 | 2 | 1 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 2 | 0 | 2 | 2 | 2 | 2 | 0 | 1 | 2 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 2 | 1 | 2 | 2 | 1 | 2 | 2 | 1 | 1 | 0 | 1 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 2 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 2 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 2 | 2 | 2 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 1 | 0 | 2 | 1 | 0 | 1 | 1 | 2 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 2 | 2 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 2 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 2 | 2 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 1 | 2 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 2 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 6 | 91 | 7 | 5 | 8 | 9 | 8 | 8 | 7 | 11 | 14 | 15 | 21 | 27 | 36 | 47 | 59 | 65 | 77 | 104 | 137 | 165 | 207 | 271 | 336 | 401 | 453 | 499 | 561 | 637 | 672 | 672 | 686 | 732 | 704 | 679 | 697 | 697 | 708 | 702 | 708 | 727 | 739 | 737 | 740 | 772 | 761 | 716 | 726 | 761 | 772 | 754 | 725 | 723 | 736 | 768 | 757 | 721 | 732 | 739 | 747 | 746 | 739 | 762 | 775 | 738 | 780 | 772 | 735 | 755 | 765 | 749 | 735 | 753 | 776 | 774 | 780 | 815 | 782 | 770 | 781 | 775 | 753 | 739 | 780 | 794 | 774 | 782 | 777 | 759 | 768 | 755 | 724 | 710 | 711 | 726 | 717 | 701 | 728 | 737 | 712 | 680 | 666 | 695 | 699 | 696 | 734 | 701 | 714 | 734 | 726 | 709 | 705 | 722 | 751 | 732 | 739 | 729 | 732 | 709 | 734 | 661 | 620 | 671 | 674 | 677 | 647 | 657 | 655 | 634 | 645 | 648 | 625 | 611 | 612 | 645 | 647 | 638 | 624 | 640 | 653 | 663 | 668 | 649 | 646 | 660 | 675 | 668 | 648 | 650 | 671 | 679 | 661 | 673 | 668 | 669 | 674 | 677 | 724 | 746 | 713 | 726 | 739 | 715 | 732 | 772 | 798 | 791 | 821 | 834 | 886 | 916 | 933 | 931 | 960 | 976 | 1015 | 986 | 994 | 998 | 1012 | 980 | 976 | 992 | 982 | 972 | 974 | 979 | 974 | 958 | 932 | 949 | 942 | 909 | 931 | 952 | 925 | 898 | 906 | 907 | 876 | 849 | 818 | 811 | 844 | 825 | 793 | 795 | 805 | 809 | 773 | 764 | 780 | 768 | 745 | 747 | 713 | 742 | 732 | 688 | 671 | 688 | 714 | 712 | 664 | 672 | 621 | 653 | 652 | 677 | 654 | 648 | 616 | 648 | 617 | 650 | 631 | 633 | 608 | 599 | 594 | 598 | 597 | 595 | 566 | 573 | 593 | 576 | 570 | 559 | 544 | 538 | 536 | 531 | 536 | 551 | 537 | 519 | 509 | 527 | 544 | 521 | 486 | 477 | 475 | 501 | 499 | 498 | 495 | 470 | 510 | 457 | 481 | 452 | 442 | 472 | 442 | 480 | 422 | 469 | 456 | 436 | 413 | 438 | 435 | 424 | 431 | 460 | 443 | 422 | 424 | 426 | 420 | 416 | 421 | 419 | 424 | 394 | 388 | 402 | 405 | 407 | 400 | 391 | 392 | 396 | 401 | 395 | 395 | 395 | 388 | 396 | 399 | 384 | 366 | 381 | 351 | 360 | 355 | 340 | 343 | 352 | 360 | 348 | 308 | 325 | 359 | 357 | 361 | 323 | 308 | 316 | 309 | 332 | 311 | 316 | 308 | 308 | 300 | 296 | 315 | 299 | 297 | 284 | 295 | 283 | 274 | 284 | 280 | 285 | 280 | 267 | 268 | 272 | 276 | 269 | 275 | 292 | 292 | 294 | 297 | 292 | 294 | 297 | 279 | 268 | 265 | 255 | 273 | 251 | 243 | 258 | 255 | 257 | 248 | 238 | 241 | 224 | 248 | 272 | 246 | 272 | 259 | 229 | 242 | 241 | 252 | 250 | 246 | 241 | 229 | 232 | 239 | 234 | 239 | 243 | 237 | 208 | 206 | 229 | 230 | 223 | 223 | 233 | 226 | 208 | 211 | 216 | 215 | 221 | 220 | 218 | 219 | 212 | 215 | 215 | 210 | 203 | 200 | 209 | 220 | 224 | 212 | 214 | 215 | 219 | 213 | 202 | 208 | 213 | 212 | 233 | 228 | 217 | 223 | 231 | 217 | 221 | 237 | 231 | 219 | 219 | 235 | 239 | 255 | 263 | 272 | 280 | 276 | 279 | 313 | 345 | 360 | 377 | 417 | 461 | 491 | 540 | 601 | 656 | 711 | 785 | 858 | 933 | 1058 | 1198 | 1284 | 1376 | 1531 | 1673 | 1826 | 2007 | 2176 | 2376 | 2559 | 2703 | 2890 | 3183 | 3427 | 3616 | 3901 | 4160 | 4284 | 4457 | 4780 | 5029 | 5199 | 5390 | 5581 | 5688 | 5736 | 5881 | 5943 | 5999 | 6099 | 6113 | 6097 | 6025 | 5978 | 5925 | 5827 | 5761 | 5724 | 5593 | 5362 | 5110 | 4873 | 4659 | 4459 | 4219 | 3958 | 3718 | 3470 | 3218 | 2997 | 2787 | 2563 | 2367 | 2186 | 1975 | 1800 | 1667 | 1516 | 1353 | 1241 | 1148 | 1048 | 974 | 884 | 828 | 769 | 670 | 634 | 592 | 525 | 511 | 469 | 447 | 440 | 439 | 440 | 409 | 409 | 403 | 369 | 364 | 379 | 379 | 363 | 359 | 364 | 364 | 354 | 340 | 339 | 342 | 360 | 374 | 372 | 368 | 365 | 359 | 347 | 349 | 363 | 374 | 381 | 371 | 356 | 350 | 363 | 376 | 374 | 362 | 348 | 350 | 351 | 358 | 357 | 349 | 347 | 350 | 352 | 342 | 329 | 340 | 349 | 343 | 358 | 352 | 350 | 338 | 335 | 351 | 357 | 344 | 334 | 317 | 324 | 320 | 330 | 326 | 332 | 335 | 304 | 316 | 323 | 320 | 311 | 299 | 313 | 322 | 326 | 331 | 319 | 305 | 308 | 305 | 296 | 300 | 299 | 319 | 336 | 313 | 291 | 287 | 318 | 307 | 318 | 304 | 306 | 296 | 306 | 318 | 303 | 302 | 305 | 287 | 289 | 308 | 326 | 338 | 309 | 290 | 287 | 324 | 310 | 314 | 316 | 321 | 342 | 344 | 340 | 357 | 350 | 329 | 331 | 341 | 358 | 357 | 318 | 379 | 335 | 363 | 366 | 369 | 379 | 396 | 395 | 385 | 383 | 410 | 432 | 422 | 415 | 410 | 413 | 404 | 419 | 392 | 437 | 465 | 424 | 392 | 415 | 407 | 415 | 418 | 417 | 402 | 411 | 424 | 400 | 377 | 387 | 391 | 384 | 378 | 366 | 376 | 354 | 396 | 397 | 361 | 340 | 314 | 354 | 330 | 336 | 340 | 333 | 320 | 308 | 314 | 322 | 319 | 315 | 300 | 287 | 288 | 290 | 302 | 296 | 276 | 284 | 248 | 258 | 281 | 274 | 262 | 268 | 280 | 270 | 261 | 265 | 251 | 253 | 251 | 241 | 217 | 226 | 230 | 231 | 224 | 211 | 227 | 224 | 241 | 219 | 205 | 213 | 219 | 212 | 209 | 212 | 197 | 202 | 203 | 207 | 211 | 199 | 202 | 173 | 182 | 177 | 176 | 213 | 182 | 177 | 195 | 189 | 175 | 179 | 197 | 193 | 158 | 172 | 172 | 176 | 165 | 185 | 158 | 160 | 143 | 158 | 160 | 154 | 147 | 149 | 155 | 163 | 154 | 138 | 142 | 151 | 148 | 154 | 156 | 132 | 150 | 134 | 137 | 134 | 135 | 131 | 136 | 123 | 127 | 121 | 121 | 118 | 113 | 108 | 118 | 133 | 119 | 121 | 116 | 108 | 112 | 112 | 114 | 114 | 125 | 106 | 105 | 114 | 103 | 106 | 104 | 99 | 92 | 86 | 90 | 105 | 106 | 90 | 99 | 92 | 86 | 101 | 88 | 88 | 82 | 78 | 82 | 91 | 89 | 82 | 77 | 79 | 74 | 67 | 78 | 87 | 74 | 66 | 71 | 84 | 82 | 70 | 87 | 78 | 79 | 71 | 69 | 69 | 73 | 73 | 77 | 78 | 72 | 67 | 64 | 65 | 66 | 53 | 60 | 69 | 65 | 73 | 73 | 62 | 63 | 69 | 65 | 62 | 58 | 59 | 60 | 60 | 62 | 65 | 67 | 67 | 62 | 56 | 58 | 60 | 56 | 51 | 53 | 51 | 54 | 59 | 58 | 60 | 60 | 59 | 56 | 51 | 49 | 55 | 58 | 61 | 60 | 55 | 55 | 52 | 57 | 52 | 59 | 58 | 61 | 58 | 57 | 57 | 61 | 63 | 62 | 60 | 60 | 60 | 57 | 56 | 55 | 55 | 62 | 69 | 78 | 71 | 66 | 69 | 71 | 72 | 70 | 72 | 78 | 83 | 84 | 91 | 95 | 102 | 114 | 116 | 114 | 114 | 122 | 142 | 165 | 166 | 161 | 173 | 185 | 202 | 211 | 226 | 253 | 268 | 284 | 314 | 331 | 332 | 338 | 358 | 384 | 396 | 393 | 415 | 493 | 496 | 499 | 541 | 549 | 575 | 587 | 619 | 640 | 630 | 636 | 663 | 691 | 696 | 700 | 728 | 771 | 797 | 815 | 829 | 868 | 880 | 855 | 848 | 836 | 812 | 835 | 858 | 830 | 811 | 806 | 788 | 781 | 781 | 778 | 773 | 754 | 718 | 716 | 745 | 714 | 647 | 602 | 625 | 610 | 552 | 532 | 505 | 496 | 494 | 462 | 430 | 409 | 385 | 349 | 331 | 330 | 317 | 291 | 261 | 233 | 234 | 207 | 180 | 171 | 152 | 141 | 134 | 122 | 109 | 102 | 95 | 86 | 80 | 76 | 63 | 51 | 43 | 44 | 45 | 37 | 38 | 31 | 23 | 26 | 21 | 19 | 19 | 16 | 14 | 15 | 15 | 13 | 10 | 9 | 10 | 12 | 9 | 8 | 9 | 11 | 9 | 8 | 7 | 8 | 6 | 4 | 4 | 7 | 9 | 8 | 6 | 6 | 5 | 6 | 5 | 6 | 5 | 3 | 5 | 7 | 6 | 4 | 4 | 5 | 7 | 7 | 10 | 9 | 6 | 7 | 6 | 4 | 6 | 9 | 12 | 8 | 8 | 6 | 6 | 5 | 2 | 3 | 2 | 3 | 4 | 5 | 6 | 6 | 8 | 6 | 5 | 6 | 6 | 4 | 3 | 5 | 3 | 4 | 3 | 8 | 4 | 10 | 7 | 4 | 4 | 7 | 8 | 7 | 8 | 9 | 8 | 8 | 6 | 3 | 3 | 4 | 5 | 4 | 3 | 3 | 5 | 3 | 4 | 8 | 5 | 6 | 4 | 4 | 6 | 5 | 4 | 3 | 5 | 10 | 6 | 5 | 3 | 5 | 3 | 4 | 5 | 8 | 5 | 6 | 3 | 4 | 5 | 4 | 4 | 5 | 4 | 2 | 3 | 6 | 7 | 2 | 3 | 5 | 6 | 5 | 5 | 6 | 3 | 3 | 4 | 3 | 2 | 4 | 5 | 3 | 4 | 3 | 3 | 4 | 6 | 2 | 2 | 3 | 5 | 1 | 6 | 1 | 5 | 4 | 4 | 4 | 4 | 4 | 4 | 5 | 5 | 4 | 4 | 6 | 4 | 3 | 3 | 1 | 5 | 6 | 6 | 4 | 5 | 3 | 2 | 6 | 6 | 5 | 4 | 3 | 2 | 3 | 6 | 2 | 3 | 2 | 3 | 4 | 9 | 2 | 3 | 3 | 4 | 3 | 1 | 1 | 1 | 3 | 3 | 6 | 5 | 3 | 3 | 2 | 2 | 2 | 3 | 4 | 4 | 2 | 5 | 3 | 0 | 2 | 1 | 3 | 4 | 3 | 2 | 2 | 2 | 3 | 3 | 1 | 3 | 3 | 2 | 1 | 3 | 3 | 2 | 0 | 2 | 1 | 2 | 3 | 2 | 2 | 4 | 4 | 3 | 4 | 4 | 2 | 3 | 2 | 0 | 2 | 2 | 2 | 3 | 2 | 4 | 2 | 0 | 2 | 1 | 2 | 2 | 1 | 3 | 3 | 3 | 3 | 4 | 2 | 3 | 3 | 4 | 5 | 3 | 1 | 0 | 0 | 1 | 2 | 2 | 2 | 1 | 1 | 2 | 1 | 2 | 3 | 4 | 2 | 2 | 2 | 2 | 2 | 1 | 1 | 2 | 1 | 0 | 0 | 2 | 3 | 2 | 1 | 2 | 3 | 2 | 2 | 2 | 0 | 1 | 2 | 2 | 3 | 3 | 4 | 3 | 2 | 2 | 1 | 1 | 2 | 1 | 2 | 2 | 1 | 3 | 4 | 3 | 3 | 2 | 2 | 1 | 2 | 2 | 1 | 1 | 0 | 0 | 2 | 3 | 2 | 2 | 3 | 2 | 1 | 2 | 2 | 1 | 1 | 2 | 2 | 2 | 1 | 1 | 2 | 2 | 2 | 2 | 3 | 3 | 2 | 2 | 2 | 2 | 2 | 2 | 0 | 2 | 3 | 1 | 0 | 3 | 4 | 1 | 2 | 1 | 1 | 1 | 2 | 1 | 2 | 3 | 4 | 2 | 1 | 1 | 2 | 1 | 2 | 3 | 3 | 5 | 3 | 1 | 1 | 1 | 3 | 4 | 1 | 1 | 0 | 1 | 1 | 1 | 3 | 3 | 4 | 3 | 2 | 2 | 1 | 0 | 1 | 1 | 3 | 3 | 2 | 3 | 3 | 4 | 3 | 3 | 1 | 0 | 2 | 2 | 1 | 1 | 2 | 2 | 0 | 2 | 4 | 2 | 2 | 1 | 2 | 3 | 3 | 2 | 5 | 5 | 2 | 2 | 1 | 2 | 2 | 1 | 1 | 0 | 0 | 2 | 2 | 0 | 1 | 4 | 2 | 2 | 2 | 0 | 1 | 1 | 1 | 0 | 1 | 2 | 1 | 0 | 1 | 1 | 1 | 2 | 2 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 2 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 2 | 0 | 1 | 0 | 0 | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 4 | 1 | 3 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 2 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 2 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 7 | 95 | 6 | 7 | 9 | 7 | 7 | 7 | 9 | 12 | 13 | 15 | 19 | 26 | 34 | 35 | 49 | 72 | 83 | 96 | 134 | 163 | 200 | 263 | 324 | 384 | 448 | 496 | 542 | 589 | 605 | 631 | 672 | 703 | 687 | 669 | 684 | 692 | 693 | 689 | 692 | 713 | 732 | 734 | 727 | 719 | 690 | 679 | 694 | 713 | 725 | 729 | 727 | 720 | 702 | 741 | 747 | 741 | 730 | 700 | 708 | 734 | 733 | 710 | 685 | 698 | 751 | 745 | 710 | 705 | 701 | 701 | 705 | 730 | 724 | 719 | 734 | 778 | 745 | 722 | 745 | 752 | 764 | 746 | 717 | 708 | 709 | 726 | 768 | 743 | 712 | 701 | 695 | 704 | 703 | 714 | 709 | 732 | 730 | 738 | 701 | 705 | 690 | 661 | 651 | 660 | 676 | 683 | 668 | 703 | 706 | 711 | 716 | 691 | 675 | 649 | 654 | 676 | 694 | 683 | 646 | 624 | 618 | 619 | 630 | 629 | 621 | 626 | 597 | 606 | 600 | 620 | 624 | 609 | 620 | 607 | 623 | 629 | 630 | 650 | 640 | 612 | 623 | 630 | 623 | 637 | 634 | 610 | 622 | 630 | 625 | 641 | 650 | 657 | 642 | 658 | 677 | 695 | 691 | 701 | 696 | 728 | 702 | 677 | 727 | 757 | 762 | 823 | 834 | 838 | 822 | 811 | 900 | 914 | 899 | 931 | 963 | 964 | 961 | 938 | 966 | 976 | 993 | 963 | 971 | 985 | 978 | 976 | 955 | 915 | 894 | 889 | 877 | 879 | 897 | 881 | 856 | 856 | 841 | 828 | 826 | 812 | 809 | 808 | 798 | 791 | 746 | 719 | 716 | 742 | 758 | 757 | 700 | 694 | 718 | 697 | 667 | 696 | 676 | 671 | 660 | 722 | 681 | 655 | 634 | 647 | 648 | 655 | 644 | 622 | 643 | 613 | 583 | 572 | 581 | 561 | 624 | 591 | 571 | 582 | 584 | 560 | 546 | 543 | 538 | 536 | 516 | 518 | 517 | 531 | 555 | 553 | 539 | 531 | 543 | 543 | 519 | 500 | 503 | 513 | 475 | 463 | 497 | 490 | 512 | 504 | 471 | 468 | 466 | 489 | 470 | 465 | 495 | 425 | 459 | 456 | 459 | 441 | 455 | 439 | 423 | 396 | 423 | 406 | 398 | 414 | 427 | 447 | 421 | 416 | 416 | 418 | 426 | 433 | 411 | 397 | 395 | 383 | 372 | 381 | 374 | 379 | 379 | 374 | 376 | 385 | 390 | 384 | 388 | 375 | 370 | 374 | 355 | 389 | 376 | 371 | 370 | 352 | 375 | 356 | 323 | 316 | 325 | 326 | 329 | 341 | 320 | 336 | 354 | 323 | 307 | 308 | 298 | 289 | 280 | 274 | 292 | 305 | 271 | 300 | 299 | 288 | 273 | 269 | 282 | 278 | 266 | 276 | 283 | 278 | 266 | 262 | 263 | 274 | 275 | 276 | 279 | 264 | 253 | 252 | 256 | 258 | 245 | 256 | 263 | 251 | 253 | 245 | 240 | 243 | 241 | 239 | 235 | 238 | 249 | 243 | 237 | 232 | 228 | 241 | 246 | 237 | 250 | 257 | 244 | 222 | 202 | 213 | 237 | 247 | 213 | 221 | 215 | 224 | 235 | 222 | 229 | 215 | 223 | 226 | 225 | 217 | 210 | 207 | 208 | 222 | 227 | 216 | 210 | 209 | 207 | 208 | 210 | 212 | 208 | 212 | 220 | 214 | 207 | 205 | 207 | 194 | 191 | 207 | 207 | 206 | 220 | 216 | 209 | 208 | 203 | 202 | 194 | 176 | 191 | 190 | 192 | 209 | 226 | 236 | 225 | 226 | 247 | 254 | 242 | 256 | 271 | 260 | 267 | 275 | 284 | 306 | 343 | 391 | 423 | 438 | 459 | 496 | 546 | 609 | 688 | 778 | 865 | 949 | 1028 | 1125 | 1258 | 1374 | 1480 | 1631 | 1804 | 1991 | 2171 | 2297 | 2524 | 2695 | 2826 | 2999 | 3177 | 3361 | 3620 | 3930 | 4230 | 4419 | 4618 | 4779 | 4929 | 5118 | 5229 | 5347 | 5537 | 5688 | 5735 | 5874 | 5936 | 5940 | 5900 | 5885 | 5873 | 5779 | 5697 | 5585 | 5442 | 5239 | 4988 | 4769 | 4587 | 4457 | 4302 | 4072 | 3790 | 3563 | 3371 | 3104 | 2865 | 2686 | 2478 | 2254 | 2084 | 1940 | 1744 | 1571 | 1440 | 1306 | 1188 | 1088 | 1009 | 944 | 841 | 728 | 687 | 648 | 600 | 559 | 504 | 477 | 478 | 466 | 436 | 417 | 418 | 420 | 403 | 378 | 376 | 377 | 354 | 338 | 349 | 349 | 352 | 371 | 374 | 364 | 361 | 364 | 367 | 350 | 348 | 359 | 343 | 328 | 337 | 349 | 353 | 349 | 345 | 349 | 351 | 348 | 342 | 340 | 338 | 337 | 347 | 335 | 330 | 339 | 341 | 327 | 325 | 317 | 319 | 329 | 329 | 304 | 300 | 312 | 331 | 335 | 331 | 316 | 336 | 332 | 295 | 321 | 318 | 306 | 330 | 310 | 312 | 313 | 315 | 337 | 323 | 301 | 309 | 299 | 302 | 325 | 318 | 305 | 302 | 293 | 288 | 291 | 299 | 304 | 300 | 300 | 296 | 294 | 296 | 300 | 300 | 304 | 303 | 297 | 285 | 280 | 302 | 294 | 290 | 289 | 294 | 283 | 283 | 289 | 295 | 289 | 302 | 306 | 302 | 321 | 306 | 309 | 299 | 300 | 289 | 301 | 322 | 320 | 313 | 314 | 317 | 336 | 333 | 334 | 333 | 322 | 359 | 367 | 368 | 374 | 371 | 361 | 363 | 367 | 363 | 374 | 377 | 373 | 380 | 393 | 409 | 414 | 407 | 388 | 379 | 373 | 373 | 413 | 388 | 389 | 389 | 382 | 391 | 400 | 408 | 406 | 401 | 396 | 388 | 381 | 376 | 372 | 366 | 356 | 362 | 366 | 374 | 382 | 380 | 364 | 354 | 347 | 347 | 335 | 327 | 317 | 314 | 326 | 326 | 315 | 311 | 288 | 295 | 279 | 290 | 292 | 284 | 298 | 296 | 262 | 268 | 259 | 279 | 273 | 268 | 273 | 280 | 272 | 259 | 256 | 253 | 245 | 245 | 236 | 229 | 212 | 222 | 211 | 238 | 242 | 209 | 201 | 210 | 208 | 215 | 216 | 210 | 192 | 189 | 199 | 198 | 204 | 189 | 200 | 190 | 196 | 191 | 191 | 189 | 190 | 185 | 176 | 179 | 166 | 170 | 169 | 161 | 172 | 177 | 160 | 161 | 174 | 166 | 159 | 163 | 158 | 157 | 171 | 163 | 171 | 143 | 157 | 147 | 145 | 150 | 164 | 171 | 158 | 143 | 136 | 136 | 142 | 139 | 145 | 153 | 140 | 127 | 125 | 126 | 117 | 115 | 115 | 126 | 130 | 121 | 134 | 122 | 110 | 115 | 111 | 112 | 109 | 113 | 114 | 114 | 119 | 129 | 120 | 111 | 112 | 105 | 117 | 101 | 102 | 103 | 96 | 94 | 98 | 102 | 101 | 103 | 99 | 94 | 85 | 79 | 84 | 87 | 78 | 86 | 82 | 92 | 90 | 82 | 89 | 84 | 81 | 85 | 77 | 65 | 67 | 75 | 81 | 81 | 81 | 73 | 80 | 69 | 77 | 80 | 70 | 70 | 62 | 67 | 68 | 64 | 66 | 67 | 63 | 59 | 64 | 69 | 67 | 69 | 72 | 59 | 50 | 53 | 69 | 63 | 70 | 59 | 53 | 56 | 63 | 71 | 69 | 65 | 62 | 58 | 61 | 65 | 59 | 54 | 56 | 56 | 50 | 58 | 51 | 51 | 48 | 46 | 43 | 47 | 48 | 50 | 51 | 47 | 45 | 48 | 49 | 52 | 53 | 51 | 54 | 52 | 51 | 53 | 53 | 58 | 52 | 51 | 49 | 47 | 55 | 57 | 50 | 50 | 52 | 48 | 47 | 58 | 70 | 70 | 62 | 62 | 64 | 60 | 70 | 76 | 76 | 77 | 74 | 79 | 79 | 77 | 83 | 90 | 93 | 104 | 115 | 118 | 120 | 114 | 117 | 120 | 138 | 152 | 162 | 167 | 179 | 192 | 204 | 227 | 236 | 234 | 246 | 270 | 280 | 289 | 315 | 340 | 364 | 391 | 395 | 409 | 459 | 468 | 456 | 486 | 521 | 549 | 576 | 613 | 633 | 619 | 641 | 670 | 688 | 701 | 705 | 717 | 763 | 802 | 809 | 791 | 780 | 803 | 811 | 820 | 828 | 838 | 845 | 849 | 850 | 837 | 821 | 825 | 824 | 819 | 797 | 761 | 742 | 722 | 709 | 688 | 679 | 657 | 611 | 592 | 572 | 560 | 546 | 536 | 523 | 483 | 437 | 410 | 391 | 364 | 330 | 313 | 314 | 291 | 269 | 253 | 247 | 216 | 186 | 187 | 166 | 153 | 144 | 136 | 128 | 109 | 97 | 90 | 81 | 71 | 65 | 58 | 50 | 46 | 47 | 51 | 41 | 34 | 34 | 25 | 22 | 21 | 20 | 16 | 14 | 16 | 15 | 13 | 13 | 11 | 9 | 9 | 12 | 9 | 11 | 12 | 10 | 7 | 5 | 7 | 9 | 7 | 6 | 5 | 6 | 7 | 7 | 9 | 11 | 11 | 7 | 6 | 7 | 6 | 5 | 6 | 3 | 7 | 6 | 6 | 5 | 5 | 4 | 6 | 8 | 8 | 7 | 5 | 5 | 5 | 4 | 3 | 4 | 4 | 7 | 9 | 5 | 4 | 2 | 2 | 5 | 6 | 5 | 6 | 5 | 3 | 4 | 6 | 7 | 6 | 4 | 2 | 4 | 7 | 8 | 5 | 3 | 4 | 4 | 6 | 6 | 5 | 4 | 4 | 4 | 5 | 4 | 4 | 3 | 5 | 5 | 3 | 5 | 4 | 9 | 3 | 7 | 3 | 4 | 3 | 4 | 6 | 6 | 6 | 5 | 4 | 5 | 6 | 3 | 3 | 3 | 5 | 4 | 7 | 3 | 4 | 4 | 3 | 4 | 2 | 2 | 4 | 4 | 4 | 4 | 5 | 4 | 4 | 3 | 5 | 7 | 6 | 4 | 4 | 7 | 6 | 4 | 4 | 4 | 7 | 4 | 5 | 7 | 4 | 4 | 6 | 5 | 5 | 4 | 3 | 3 | 4 | 2 | 2 | 6 | 2 | 2 | 5 | 3 | 1 | 2 | 4 | 4 | 2 | 3 | 4 | 1 | 2 | 2 | 1 | 1 | 3 | 2 | 3 | 3 | 3 | 2 | 2 | 3 | 3 | 2 | 2 | 5 | 5 | 4 | 5 | 5 | 4 | 4 | 3 | 1 | 2 | 2 | 1 | 1 | 4 | 2 | 3 | 3 | 3 | 3 | 2 | 4 | 2 | 3 | 5 | 7 | 4 | 3 | 4 | 4 | 0 | 1 | 1 | 2 | 4 | 1 | 3 | 3 | 3 | 2 | 1 | 4 | 2 | 2 | 2 | 2 | 4 | 4 | 3 | 2 | 2 | 2 | 2 | 4 | 2 | 2 | 2 | 2 | 2 | 1 | 3 | 2 | 0 | 1 | 0 | 2 | 1 | 0 | 1 | 3 | 2 | 2 | 2 | 3 | 4 | 3 | 0 | 3 | 2 | 0 | 1 | 2 | 2 | 3 | 2 | 2 | 1 | 2 | 1 | 0 | 1 | 0 | 2 | 0 | 3 | 1 | 3 | 3 | 2 | 1 | 1 | 0 | 1 | 2 | 1 | 1 | 2 | 1 | 0 | 0 | 2 | 2 | 1 | 0 | 3 | 1 | 1 | 2 | 2 | 3 | 1 | 1 | 2 | 2 | 1 | 1 | 1 | 0 | 2 | 1 | 2 | 2 | 3 | 1 | 0 | 2 | 3 | 3 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 2 | 2 | 2 | 2 | 0 | 0 | 0 | 2 | 3 | 1 | 1 | 2 | 1 | 2 | 2 | 2 | 2 | 1 | 2 | 2 | 1 | 1 | 2 | 2 | 0 | 1 | 1 | 1 | 1 | 1 | 2 | 0 | 2 | 3 | 1 | 0 | 1 | 2 | 2 | 1 | 0 | 2 | 3 | 3 | 2 | 1 | 1 | 3 | 3 | 3 | 2 | 2 | 3 | 1 | 0 | 1 | 2 | 2 | 2 | 2 | 3 | 2 | 1 | 2 | 2 | 4 | 2 | 2 | 3 | 0 | 1 | 3 | 4 | 1 | 1 | 2 | 3 | 3 | 1 | 1 | 1 | 2 | 4 | 2 | 0 | 0 | 1 | 1 | 1 | 2 | 3 | 1 | 0 | 1 | 1 | 1 | 2 | 2 | 2 | 2 | 4 | 4 | 3 | 2 | 1 | 1 | 2 | 1 | 2 | 2 | 1 | 4 | 5 | 3 | 2 | 2 | 1 | 1 | 2 | 1 | 2 | 2 | 2 | 3 | 3 | 1 | 2 | 2 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 2 | 1 | 2 | 2 | 3 | 2 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 3 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 2 | 2 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 8 | 103 | 4 | 4 | 7 | 8 | 9 | 11 | 9 | 11 | 13 | 18 | 20 | 19 | 28 | 40 | 53 | 67 | 77 | 95 | 121 | 153 | 207 | 285 | 367 | 406 | 437 | 500 | 554 | 610 | 647 | 660 | 655 | 681 | 718 | 707 | 684 | 665 | 661 | 675 | 687 | 671 | 684 | 709 | 733 | 727 | 700 | 697 | 707 | 716 | 691 | 703 | 731 | 704 | 708 | 728 | 707 | 696 | 696 | 701 | 713 | 716 | 694 | 701 | 729 | 707 | 739 | 717 | 673 | 673 | 682 | 711 | 718 | 711 | 712 | 716 | 723 | 752 | 725 | 710 | 731 | 722 | 715 | 714 | 707 | 716 | 739 | 750 | 686 | 679 | 682 | 668 | 679 | 678 | 688 | 691 | 697 | 672 | 677 | 705 | 691 | 672 | 672 | 686 | 676 | 639 | 632 | 620 | 628 | 706 | 731 | 700 | 686 | 671 | 655 | 677 | 696 | 673 | 644 | 650 | 676 | 605 | 603 | 610 | 614 | 605 | 576 | 586 | 569 | 583 | 577 | 583 | 591 | 614 | 613 | 616 | 616 | 606 | 597 | 595 | 620 | 626 | 620 | 622 | 622 | 632 | 616 | 602 | 634 | 636 | 620 | 628 | 623 | 622 | 643 | 648 | 650 | 672 | 672 | 709 | 695 | 703 | 714 | 683 | 689 | 744 | 770 | 780 | 767 | 804 | 796 | 862 | 855 | 856 | 894 | 911 | 944 | 966 | 909 | 913 | 906 | 960 | 951 | 933 | 985 | 934 | 946 | 929 | 907 | 901 | 902 | 892 | 872 | 865 | 852 | 840 | 830 | 795 | 785 | 795 | 818 | 823 | 792 | 767 | 729 | 749 | 753 | 723 | 703 | 719 | 734 | 747 | 746 | 745 | 719 | 706 | 673 | 674 | 683 | 669 | 650 | 638 | 674 | 635 | 642 | 619 | 580 | 592 | 592 | 631 | 610 | 581 | 583 | 593 | 602 | 587 | 581 | 584 | 551 | 564 | 554 | 533 | 564 | 545 | 547 | 549 | 528 | 522 | 556 | 544 | 500 | 518 | 521 | 484 | 482 | 501 | 498 | 470 | 481 | 484 | 474 | 500 | 459 | 472 | 469 | 459 | 476 | 468 | 473 | 443 | 430 | 438 | 491 | 458 | 435 | 447 | 416 | 432 | 436 | 419 | 422 | 408 | 430 | 402 | 414 | 457 | 427 | 421 | 406 | 403 | 411 | 419 | 402 | 389 | 383 | 392 | 386 | 378 | 404 | 395 | 394 | 399 | 375 | 378 | 385 | 372 | 371 | 375 | 373 | 366 | 373 | 381 | 366 | 346 | 353 | 352 | 341 | 312 | 333 | 331 | 340 | 327 | 324 | 302 | 320 | 319 | 315 | 292 | 311 | 300 | 319 | 307 | 315 | 286 | 282 | 293 | 292 | 303 | 304 | 284 | 290 | 285 | 264 | 266 | 289 | 294 | 280 | 264 | 272 | 282 | 276 | 269 | 266 | 273 | 271 | 259 | 262 | 267 | 263 | 266 | 268 | 249 | 245 | 257 | 265 | 255 | 242 | 232 | 243 | 233 | 240 | 235 | 256 | 256 | 246 | 226 | 235 | 232 | 227 | 217 | 233 | 218 | 240 | 221 | 243 | 219 | 225 | 199 | 205 | 223 | 239 | 228 | 216 | 213 | 221 | 225 | 230 | 216 | 209 | 232 | 223 | 212 | 213 | 215 | 218 | 216 | 213 | 209 | 198 | 192 | 201 | 203 | 204 | 207 | 202 | 196 | 198 | 207 | 219 | 218 | 212 | 217 | 196 | 181 | 195 | 218 | 205 | 206 | 220 | 209 | 209 | 194 | 204 | 222 | 194 | 186 | 190 | 217 | 215 | 209 | 212 | 208 | 221 | 240 | 247 | 255 | 254 | 253 | 256 | 269 | 287 | 304 | 320 | 343 | 373 | 403 | 436 | 482 | 531 | 586 | 656 | 737 | 810 | 886 | 992 | 1119 | 1212 | 1283 | 1434 | 1598 | 1727 | 1906 | 2089 | 2232 | 2395 | 2600 | 2832 | 3002 | 3190 | 3383 | 3541 | 3733 | 4040 | 4268 | 4578 | 4785 | 4846 | 5033 | 5204 | 5306 | 5538 | 5641 | 5620 | 5641 | 5735 | 5769 | 5805 | 5740 | 5680 | 5687 | 5641 | 5509 | 5350 | 5245 | 5109 | 4898 | 4665 | 4411 | 4172 | 3970 | 3729 | 3470 | 3271 | 3088 | 2875 | 2671 | 2479 | 2267 | 2047 | 1855 | 1712 | 1563 | 1400 | 1276 | 1155 | 1050 | 972 | 902 | 826 | 724 | 647 | 618 | 593 | 545 | 502 | 482 | 490 | 460 | 441 | 426 | 410 | 413 | 381 | 374 | 394 | 367 | 343 | 357 | 358 | 356 | 354 | 346 | 354 | 356 | 360 | 376 | 367 | 346 | 346 | 344 | 341 | 342 | 345 | 355 | 353 | 339 | 346 | 363 | 352 | 343 | 349 | 353 | 344 | 342 | 342 | 334 | 341 | 359 | 363 | 339 | 356 | 353 | 325 | 340 | 351 | 358 | 341 | 326 | 326 | 338 | 327 | 328 | 330 | 322 | 328 | 304 | 296 | 309 | 319 | 321 | 326 | 324 | 312 | 299 | 287 | 288 | 299 | 306 | 304 | 305 | 302 | 304 | 305 | 308 | 309 | 296 | 291 | 292 | 289 | 295 | 300 | 293 | 295 | 302 | 298 | 301 | 303 | 284 | 271 | 278 | 291 | 288 | 299 | 337 | 304 | 305 | 321 | 328 | 307 | 293 | 283 | 301 | 302 | 297 | 289 | 304 | 304 | 306 | 303 | 306 | 314 | 316 | 305 | 303 | 322 | 334 | 324 | 326 | 321 | 332 | 368 | 353 | 360 | 360 | 368 | 360 | 380 | 404 | 381 | 357 | 368 | 373 | 367 | 384 | 405 | 400 | 391 | 365 | 368 | 380 | 390 | 421 | 393 | 396 | 385 | 366 | 359 | 367 | 362 | 374 | 390 | 380 | 368 | 374 | 375 | 361 | 361 | 370 | 359 | 352 | 338 | 369 | 360 | 360 | 333 | 325 | 316 | 318 | 337 | 343 | 336 | 322 | 299 | 288 | 294 | 297 | 306 | 282 | 289 | 266 | 270 | 303 | 284 | 265 | 255 | 261 | 253 | 270 | 262 | 246 | 243 | 245 | 245 | 240 | 240 | 261 | 238 | 236 | 233 | 228 | 221 | 235 | 224 | 200 | 217 | 216 | 221 | 220 | 210 | 202 | 207 | 214 | 211 | 193 | 174 | 183 | 181 | 186 | 185 | 218 | 194 | 190 | 198 | 206 | 183 | 189 | 178 | 173 | 185 | 160 | 140 | 155 | 173 | 169 | 157 | 162 | 147 | 161 | 147 | 153 | 153 | 155 | 178 | 155 | 145 | 160 | 170 | 157 | 140 | 142 | 151 | 143 | 134 | 145 | 150 | 142 | 123 | 123 | 129 | 137 | 123 | 126 | 127 | 124 | 129 | 121 | 128 | 128 | 122 | 121 | 121 | 116 | 114 | 115 | 109 | 113 | 121 | 113 | 107 | 101 | 93 | 113 | 90 | 102 | 103 | 103 | 102 | 111 | 107 | 103 | 107 | 106 | 99 | 99 | 101 | 86 | 80 | 86 | 79 | 81 | 86 | 97 | 80 | 87 | 82 | 88 | 87 | 82 | 86 | 86 | 78 | 81 | 84 | 76 | 71 | 74 | 81 | 80 | 81 | 68 | 68 | 76 | 64 | 78 | 74 | 69 | 64 | 61 | 60 | 65 | 64 | 60 | 59 | 60 | 63 | 68 | 60 | 61 | 61 | 74 | 76 | 56 | 64 | 60 | 61 | 60 | 67 | 65 | 66 | 66 | 67 | 62 | 50 | 46 | 52 | 54 | 49 | 53 | 52 | 44 | 58 | 59 | 58 | 61 | 54 | 56 | 58 | 62 | 58 | 54 | 49 | 52 | 56 | 55 | 60 | 60 | 56 | 54 | 49 | 49 | 43 | 55 | 47 | 50 | 59 | 48 | 45 | 53 | 49 | 48 | 51 | 50 | 51 | 53 | 54 | 55 | 57 | 58 | 62 | 65 | 66 | 62 | 62 | 75 | 71 | 63 | 67 | 77 | 81 | 90 | 89 | 89 | 99 | 103 | 100 | 106 | 123 | 127 | 132 | 139 | 153 | 170 | 183 | 203 | 203 | 200 | 214 | 228 | 240 | 260 | 279 | 293 | 313 | 337 | 345 | 347 | 374 | 408 | 428 | 451 | 465 | 474 | 508 | 531 | 533 | 551 | 606 | 616 | 618 | 642 | 664 | 674 | 687 | 690 | 709 | 752 | 772 | 806 | 805 | 789 | 797 | 807 | 825 | 846 | 839 | 839 | 846 | 832 | 810 | 791 | 785 | 801 | 801 | 770 | 743 | 742 | 740 | 709 | 675 | 678 | 698 | 673 | 627 | 600 | 570 | 542 | 533 | 512 | 487 | 449 | 405 | 379 | 356 | 331 | 319 | 317 | 282 | 251 | 245 | 237 | 228 | 204 | 191 | 176 | 151 | 130 | 121 | 116 | 108 | 100 | 88 | 76 | 65 | 54 | 54 | 51 | 45 | 42 | 41 | 43 | 34 | 27 | 27 | 23 | 20 | 21 | 18 | 14 | 14 | 14 | 12 | 10 | 10 | 11 | 13 | 11 | 9 | 11 | 13 | 11 | 6 | 6 | 7 | 5 | 7 | 8 | 6 | 7 | 6 | 7 | 8 | 6 | 4 | 7 | 7 | 6 | 5 | 3 | 4 | 6 | 7 | 5 | 4 | 5 | 5 | 6 | 5 | 4 | 7 | 7 | 4 | 6 | 7 | 8 | 6 | 6 | 6 | 4 | 5 | 4 | 4 | 7 | 6 | 4 | 5 | 5 | 5 | 4 | 5 | 6 | 7 | 7 | 6 | 5 | 4 | 6 | 7 | 5 | 4 | 5 | 6 | 5 | 4 | 4 | 4 | 3 | 3 | 3 | 4 | 5 | 5 | 6 | 9 | 8 | 5 | 3 | 3 | 4 | 5 | 7 | 6 | 4 | 3 | 3 | 3 | 2 | 4 | 5 | 4 | 3 | 3 | 3 | 3 | 3 | 5 | 4 | 4 | 6 | 5 | 5 | 7 | 3 | 2 | 4 | 7 | 7 | 5 | 4 | 4 | 5 | 4 | 5 | 3 | 3 | 4 | 3 | 4 | 5 | 5 | 2 | 3 | 4 | 7 | 5 | 4 | 5 | 4 | 3 | 2 | 1 | 3 | 3 | 2 | 2 | 5 | 2 | 3 | 2 | 5 | 4 | 2 | 3 | 7 | 8 | 6 | 3 | 3 | 3 | 0 | 3 | 5 | 4 | 4 | 1 | 3 | 6 | 4 | 3 | 1 | 4 | 2 | 3 | 2 | 4 | 2 | 3 | 1 | 1 | 3 | 3 | 3 | 6 | 5 | 3 | 2 | 2 | 3 | 4 | 2 | 4 | 6 | 0 | 2 | 3 | 2 | 2 | 2 | 4 | 4 | 2 | 3 | 5 | 3 | 1 | 2 | 2 | 2 | 1 | 3 | 3 | 0 | 3 | 2 | 1 | 2 | 2 | 2 | 2 | 2 | 2 | 1 | 1 | 2 | 3 | 3 | 2 | 1 | 1 | 3 | 3 | 3 | 0 | 2 | 2 | 2 | 1 | 2 | 1 | 0 | 1 | 3 | 3 | 1 | 1 | 3 | 4 | 3 | 2 | 0 | 3 | 1 | 2 | 1 | 2 | 0 | 3 | 2 | 2 | 2 | 1 | 3 | 2 | 1 | 1 | 4 | 2 | 3 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 2 | 1 | 1 | 1 | 2 | 2 | 2 | 2 | 3 | 3 | 1 | 4 | 3 | 1 | 1 | 2 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 2 | 2 | 2 | 2 | 3 | 0 | 2 | 2 | 1 | 1 | 2 | 1 | 2 | 4 | 4 | 2 | 1 | 2 | 2 | 3 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 2 | 0 | 0 | 1 | 1 | 1 | 2 | 2 | 2 | 0 | 1 | 1 | 2 | 3 | 1 | 2 | 1 | 0 | 1 | 3 | 1 | 2 | 2 | 1 | 1 | 1 | 1 | 2 | 2 | 2 | 1 | 1 | 0 | 0 | 1 | 2 | 3 | 2 | 1 | 2 | 2 | 1 | 0 | 1 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 2 | 2 | 1 | 1 | 2 | 4 | 2 | 1 | 2 | 3 | 2 | 3 | 2 | 3 | 3 | 2 | 2 | 3 | 3 | 3 | 4 | 4 | 5 | 5 | 4 | 2 | 3 | 1 | 2 | 3 | 3 | 2 | 2 | 1 | 0 | 1 | 2 | 3 | 1 | 2 | 1 | 1 | 3 | 5 | 2 | 3 | 3 | 4 | 3 | 1 | 3 | 4 | 2 | 2 | 3 | 2 | 3 | 2 | 1 | 2 | 0 | 1 | 1 | 1 | 2 | 3 | 2 | 1 | 2 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 2 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 2 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 2 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 2 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 2 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 9 | 111 | 5 | 6 | 6 | 6 | 6 | 6 | 7 | 9 | 16 | 16 | 11 | 23 | 37 | 43 | 46 | 58 | 72 | 93 | 128 | 163 | 211 | 269 | 317 | 383 | 464 | 514 | 537 | 570 | 610 | 624 | 642 | 689 | 687 | 683 | 703 | 676 | 671 | 696 | 692 | 678 | 696 | 701 | 709 | 738 | 720 | 704 | 706 | 689 | 671 | 665 | 687 | 730 | 735 | 714 | 710 | 687 | 687 | 693 | 692 | 691 | 693 | 686 | 698 | 723 | 750 | 737 | 708 | 711 | 707 | 706 | 706 | 688 | 673 | 698 | 699 | 716 | 727 | 711 | 689 | 706 | 742 | 744 | 731 | 718 | 736 | 739 | 729 | 681 | 667 | 699 | 705 | 673 | 685 | 689 | 676 | 685 | 702 | 732 | 678 | 638 | 643 | 659 | 662 | 630 | 654 | 686 | 678 | 713 | 695 | 680 | 674 | 677 | 655 | 625 | 624 | 628 | 646 | 660 | 685 | 617 | 594 | 606 | 618 | 629 | 624 | 607 | 583 | 593 | 617 | 618 | 617 | 605 | 605 | 633 | 657 | 659 | 635 | 600 | 622 | 627 | 622 | 626 | 618 | 612 | 618 | 620 | 612 | 623 | 627 | 616 | 626 | 638 | 643 | 646 | 652 | 671 | 681 | 664 | 661 | 688 | 703 | 714 | 736 | 757 | 755 | 809 | 814 | 804 | 826 | 872 | 881 | 886 | 875 | 898 | 937 | 970 | 954 | 960 | 955 | 935 | 936 | 946 | 941 | 954 | 951 | 928 | 904 | 890 | 897 | 889 | 899 | 885 | 847 | 845 | 865 | 856 | 820 | 795 | 798 | 794 | 773 | 759 | 752 | 741 | 734 | 752 | 736 | 729 | 720 | 702 | 694 | 717 | 703 | 719 | 705 | 669 | 700 | 705 | 653 | 644 | 635 | 638 | 641 | 677 | 625 | 592 | 564 | 589 | 583 | 592 | 575 | 546 | 565 | 575 | 599 | 594 | 577 | 559 | 551 | 525 | 534 | 530 | 539 | 516 | 514 | 541 | 554 | 530 | 501 | 510 | 514 | 503 | 508 | 503 | 482 | 487 | 496 | 474 | 486 | 463 | 468 | 484 | 493 | 473 | 457 | 451 | 503 | 454 | 457 | 420 | 464 | 419 | 412 | 410 | 450 | 444 | 423 | 437 | 416 | 384 | 409 | 437 | 409 | 421 | 384 | 410 | 399 | 383 | 379 | 393 | 398 | 369 | 363 | 378 | 374 | 387 | 372 | 381 | 392 | 382 | 373 | 375 | 382 | 379 | 379 | 371 | 353 | 351 | 347 | 352 | 342 | 338 | 348 | 335 | 334 | 348 | 322 | 321 | 369 | 357 | 351 | 318 | 342 | 326 | 338 | 294 | 283 | 301 | 309 | 313 | 303 | 329 | 287 | 295 | 307 | 294 | 281 | 289 | 274 | 283 | 286 | 269 | 275 | 264 | 269 | 275 | 265 | 259 | 262 | 263 | 262 | 264 | 262 | 256 | 255 | 241 | 230 | 238 | 238 | 239 | 252 | 256 | 246 | 236 | 242 | 242 | 242 | 253 | 250 | 250 | 251 | 235 | 246 | 247 | 247 | 233 | 240 | 229 | 225 | 229 | 234 | 249 | 235 | 262 | 261 | 236 | 222 | 214 | 231 | 229 | 231 | 220 | 219 | 213 | 206 | 213 | 225 | 225 | 209 | 190 | 205 | 219 | 209 | 211 | 215 | 216 | 217 | 215 | 213 | 215 | 216 | 206 | 205 | 206 | 203 | 208 | 206 | 211 | 207 | 187 | 194 | 207 | 208 | 199 | 199 | 217 | 197 | 176 | 206 | 217 | 198 | 198 | 202 | 212 | 213 | 212 | 217 | 221 | 208 | 209 | 209 | 229 | 227 | 240 | 240 | 258 | 258 | 256 | 274 | 314 | 328 | 345 | 396 | 447 | 480 | 510 | 550 | 599 | 655 | 720 | 785 | 866 | 967 | 1083 | 1189 | 1297 | 1456 | 1586 | 1752 | 1920 | 2056 | 2200 | 2409 | 2676 | 2846 | 3015 | 3287 | 3436 | 3544 | 3718 | 4014 | 4239 | 4548 | 4722 | 4821 | 5059 | 5238 | 5317 | 5451 | 5609 | 5652 | 5682 | 5815 | 5816 | 5785 | 5745 | 5728 | 5625 | 5547 | 5487 | 5362 | 5197 | 4969 | 4746 | 4561 | 4343 | 4143 | 3945 | 3695 | 3426 | 3209 | 3017 | 2810 | 2596 | 2382 | 2214 | 2066 | 1877 | 1690 | 1510 | 1328 | 1198 | 1103 | 1024 | 951 | 868 | 786 | 704 | 648 | 602 | 570 | 545 | 504 | 484 | 460 | 436 | 413 | 379 | 373 | 373 | 364 | 371 | 385 | 370 | 362 | 340 | 324 | 348 | 363 | 357 | 344 | 341 | 343 | 358 | 364 | 353 | 344 | 345 | 346 | 341 | 343 | 355 | 358 | 349 | 347 | 345 | 342 | 344 | 342 | 344 | 341 | 338 | 346 | 341 | 332 | 325 | 305 | 308 | 316 | 320 | 339 | 332 | 328 | 343 | 319 | 317 | 324 | 322 | 294 | 314 | 330 | 318 | 324 | 325 | 325 | 312 | 298 | 305 | 309 | 319 | 312 | 315 | 327 | 316 | 295 | 287 | 281 | 297 | 309 | 294 | 281 | 286 | 293 | 305 | 321 | 311 | 296 | 297 | 301 | 300 | 295 | 302 | 307 | 301 | 298 | 292 | 297 | 296 | 286 | 276 | 277 | 291 | 296 | 295 | 295 | 281 | 272 | 281 | 292 | 298 | 280 | 284 | 312 | 309 | 276 | 280 | 294 | 292 | 290 | 306 | 323 | 313 | 296 | 299 | 314 | 329 | 343 | 342 | 326 | 367 | 350 | 342 | 329 | 334 | 336 | 327 | 344 | 365 | 382 | 389 | 390 | 376 | 343 | 342 | 366 | 364 | 376 | 376 | 393 | 414 | 397 | 379 | 379 | 382 | 372 | 361 | 381 | 400 | 381 | 364 | 361 | 355 | 353 | 356 | 373 | 364 | 358 | 350 | 340 | 346 | 345 | 338 | 354 | 333 | 331 | 323 | 332 | 330 | 307 | 299 | 304 | 299 | 298 | 305 | 322 | 305 | 280 | 291 | 275 | 277 | 266 | 268 | 269 | 278 | 272 | 254 | 246 | 243 | 230 | 231 | 227 | 216 | 234 | 232 | 229 | 228 | 213 | 219 | 227 | 217 | 202 | 214 | 217 | 221 | 226 | 212 | 201 | 196 | 184 | 190 | 204 | 204 | 204 | 192 | 191 | 197 | 175 | 196 | 188 | 197 | 185 | 158 | 174 | 170 | 180 | 183 | 176 | 184 | 175 | 157 | 152 | 152 | 162 | 157 | 147 | 145 | 169 | 166 | 141 | 165 | 154 | 164 | 142 | 149 | 142 | 135 | 146 | 140 | 133 | 137 | 131 | 123 | 128 | 126 | 134 | 129 | 123 | 125 | 129 | 154 | 143 | 134 | 128 | 119 | 110 | 109 | 121 | 115 | 111 | 114 | 110 | 101 | 107 | 118 | 106 | 113 | 91 | 95 | 100 | 102 | 93 | 103 | 110 | 110 | 90 | 95 | 94 | 88 | 86 | 97 | 96 | 83 | 88 | 86 | 85 | 87 | 83 | 84 | 96 | 86 | 85 | 81 | 77 | 67 | 78 | 71 | 76 | 80 | 78 | 78 | 81 | 77 | 66 | 68 | 70 | 70 | 70 | 69 | 73 | 78 | 72 | 67 | 70 | 76 | 76 | 67 | 65 | 63 | 66 | 70 | 66 | 64 | 72 | 69 | 61 | 60 | 63 | 64 | 58 | 67 | 72 | 63 | 53 | 44 | 53 | 62 | 64 | 56 | 60 | 68 | 63 | 56 | 58 | 57 | 48 | 60 | 62 | 51 | 49 | 62 | 56 | 59 | 61 | 55 | 52 | 50 | 49 | 49 | 55 | 57 | 51 | 49 | 50 | 50 | 49 | 51 | 49 | 49 | 54 | 43 | 46 | 46 | 52 | 51 | 52 | 47 | 44 | 50 | 57 | 56 | 53 | 54 | 52 | 50 | 52 | 59 | 61 | 66 | 67 | 66 | 78 | 77 | 70 | 68 | 75 | 79 | 79 | 79 | 84 | 90 | 95 | 102 | 105 | 110 | 125 | 142 | 136 | 128 | 134 | 135 | 163 | 164 | 165 | 201 | 226 | 232 | 255 | 281 | 294 | 292 | 299 | 326 | 342 | 349 | 371 | 377 | 406 | 426 | 438 | 483 | 512 | 528 | 536 | 572 | 571 | 586 | 618 | 651 | 677 | 696 | 704 | 713 | 735 | 742 | 739 | 758 | 784 | 782 | 768 | 805 | 806 | 812 | 811 | 807 | 783 | 785 | 789 | 781 | 767 | 769 | 786 | 773 | 740 | 732 | 725 | 692 | 676 | 688 | 658 | 608 | 588 | 559 | 547 | 528 | 489 | 467 | 443 | 404 | 373 | 363 | 364 | 350 | 306 | 273 | 260 | 248 | 237 | 223 | 209 | 192 | 176 | 157 | 144 | 134 | 132 | 120 | 104 | 94 | 82 | 76 | 68 | 55 | 51 | 47 | 48 | 53 | 40 | 35 | 42 | 35 | 28 | 27 | 21 | 15 | 14 | 15 | 16 | 16 | 16 | 14 | 15 | 15 | 12 | 10 | 7 | 7 | 8 | 8 | 8 | 7 | 10 | 8 | 6 | 6 | 7 | 9 | 10 | 11 | 9 | 6 | 6 | 6 | 5 | 6 | 9 | 8 | 5 | 2 | 4 | 5 | 7 | 8 | 7 | 5 | 6 | 7 | 6 | 4 | 3 | 4 | 5 | 7 | 5 | 2 | 2 | 4 | 7 | 5 | 6 | 6 | 6 | 6 | 3 | 5 | 5 | 8 | 9 | 7 | 5 | 4 | 4 | 6 | 5 | 4 | 3 | 4 | 6 | 8 | 6 | 5 | 4 | 5 | 5 | 4 | 4 | 6 | 6 | 5 | 5 | 4 | 5 | 6 | 6 | 7 | 5 | 5 | 4 | 5 | 7 | 4 | 3 | 5 | 5 | 5 | 5 | 5 | 3 | 3 | 4 | 3 | 3 | 3 | 3 | 4 | 5 | 4 | 3 | 4 | 3 | 2 | 3 | 3 | 4 | 4 | 3 | 3 | 4 | 4 | 3 | 3 | 3 | 2 | 3 | 6 | 8 | 3 | 3 | 3 | 5 | 5 | 7 | 4 | 3 | 4 | 5 | 7 | 5 | 1 | 2 | 4 | 4 | 5 | 3 | 5 | 3 | 4 | 3 | 2 | 4 | 2 | 4 | 5 | 3 | 2 | 2 | 3 | 2 | 3 | 4 | 4 | 4 | 4 | 5 | 5 | 3 | 5 | 3 | 3 | 3 | 3 | 3 | 3 | 2 | 2 | 2 | 1 | 2 | 3 | 2 | 1 | 1 | 0 | 2 | 2 | 2 | 3 | 3 | 3 | 6 | 2 | 2 | 2 | 4 | 4 | 2 | 2 | 3 | 3 | 1 | 1 | 3 | 1 | 2 | 3 | 2 | 1 | 2 | 4 | 0 | 0 | 2 | 5 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 6 | 4 | 0 | 2 | 3 | 6 | 2 | 1 | 1 | 1 | 3 | 0 | 1 | 1 | 1 | 0 | 1 | 2 | 2 | 4 | 4 | 2 | 2 | 2 | 2 | 1 | 1 | 1 | 2 | 0 | 0 | 2 | 1 | 1 | 2 | 3 | 1 | 1 | 1 | 1 | 2 | 2 | 1 | 0 | 1 | 4 | 2 | 3 | 1 | 3 | 2 | 2 | 1 | 1 | 1 | 1 | 3 | 2 | 2 | 1 | 1 | 3 | 4 | 0 | 0 | 1 | 4 | 1 | 1 | 3 | 1 | 2 | 0 | 1 | 0 | 1 | 2 | 1 | 1 | 2 | 1 | 1 | 2 | 2 | 3 | 3 | 0 | 2 | 2 | 2 | 3 | 0 | 2 | 2 | 1 | 2 | 3 | 3 | 2 | 2 | 1 | 2 | 2 | 1 | 4 | 4 | 3 | 1 | 3 | 0 | 3 | 2 | 2 | 1 | 2 | 1 | 1 | 0 | 2 | 3 | 3 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 2 | 3 | 1 | 2 | 2 | 2 | 1 | 1 | 1 | 1 | 2 | 2 | 3 | 2 | 3 | 2 | 2 | 4 | 3 | 2 | 0 | 1 | 1 | 2 | 2 | 3 | 1 | 1 | 2 | 1 | 1 | 1 | 0 | 1 | 1 | 2 | 2 | 1 | 3 | 2 | 0 | 1 | 1 | 2 | 3 | 3 | 2 | 2 | 2 | 1 | 3 | 4 | 2 | 2 | 4 | 2 | 1 | 1 | 1 | 1 | 2 | 2 | 1 | 2 | 1 | 2 | 2 | 2 | 1 | 2 | 1 | 1 | 1 | 2 | 2 | 3 | 2 | 1 | 1 | 1 | 3 | 3 | 1 | 2 | 2 | 2 | 2 | 1 | 1 | 4 | 3 | 2 | 1 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 2 | 3 | 1 | 1 | 2 | 2 | 2 | 0 | 0 | 1 | 1 | 2 | 2 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 2 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 3 | 1 | 1 | 2 | 1 | 1 | 1 | 0 | 2 | 2 | 1 | 0 | 2 | 0 | 0 | 1 | 0 | 2 | 0 | 1 | 0 | 0 | 2 | 2 | 2 | 1 | 3 | 2 | 3 | 1 | 3 | 1 | 2 | 4 | 3 | 1 | 2 | 0 | 0 | 2 | 1 | 2 | 1 | 1 | 2 | 2 | 1 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 2 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 2 | 2 | 2 | 0 | 1 | 2 | 0 | 0 | 3 | 0 | 1 | 1 | 0 | 0 | 1 | 2 | 1 | 1 | 1 | 1 | 0 | 2 | 1 | 1 | 2 | 0 | 0 | 0 | 1 | 2 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 2 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 2 | 1 | 1 | 2 | 1 | 0 | 1 | 2 | 0 | 2 | 0 | 1 | 3 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 2 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 2 | 2 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 2 | 2 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 2 | 1 | 0 | 0 | 1 | 1 | 1 | 3 | 1 | 0 | 1 | 1 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 2 | 1 | 1 | 2 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 2 | 0 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0