## Results
Hidex Automatic Gamma Counter
Counting 18F
Run
Template | Counting 18F
Run status | Finished
Run started | 2021-09-01 21:18
Run type | Count only, single rack
Software version | 1.8.0.0
AMG serial number | 2210227
Firmware versions | 0: 50, 1: 50, F: 100
Counting time | 60
Result normalization | 2021-09-01 11:48
Comment
Counting windows
Window start | Window end | Additional windows | Isotope | Counting efficiency | Half-life
400 | 650 |  | ¹⁸F | 1 | 1.82951 h
Racks
Rack | Time | Barcode | Slot count | Latest QC check | QC status
1 | 2021-09-01 21:18 | 006 | 10 | 2021-09-01 10:39 | Calibration suggested
Results
Rack | Vial | Time | Counted time (s) | Dead time factor | ¹⁸F (counts) | ¹⁸F (CPM) | Normalized ¹⁸F (CPM) | Normalized ¹⁸F (Bq) | Tags
1 | 1 | 2021-09-01 21:19 | 60 | 1.000 | 53 | 53 | 1951 | 33
1 | 2 | 2021-09-01 21:21 | 60 | 1.140 | 623243 | 710524 | 26398892 | 439982
1 | 3 | 2021-09-01 21:22 | 60 | 1.140 | 617378 | 703878 | 26397352 | 439956
1 | 4 | 2021-09-01 21:23 | 60 | 1.137 | 609242 | 692956 | 26231132 | 437186
1 | 5 | 2021-09-01 21:25 | 60 | 1.137 | 606961 | 690096 | 26367384 | 439456
1 | 6 | 2021-09-01 21:26 | 60 | 1.048 | 232071 | 243273 | 9382496 | 156375
1 | 7 | 2021-09-01 21:28 | 60 | 1.046 | 222378 | 232498 | 9051186 | 150853
1 | 8 | 2021-09-01 21:29 | 60 | 1.045 | 219148 | 229060 | 9001224 | 150020
1 | 9 | 2021-09-01 21:31 | 60 | 1.044 | 217218 | 226825 | 8997329 | 149955
1 | 10 | 2021-09-01 21:32 | 60 | 1.000 | 53 | 53 | 2122 | 35

## Spectra
Hidex Automatic Gamma Counter
Counting 18F
Run
Template | Counting 18F
Run status | Finished
Run started | 2021-09-01 21:18
Run type | Count only, single rack
Software version | 1.8.0.0
AMG serial number | 2210227
Firmware versions | 0: 50, 1: 50, F: 100
Counting time | 60
Result normalization | 2021-09-01 11:48
Comment
Spectra
Channel
Rack | Vial | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66 | 67 | 68 | 69 | 70 | 71 | 72 | 73 | 74 | 75 | 76 | 77 | 78 | 79 | 80 | 81 | 82 | 83 | 84 | 85 | 86 | 87 | 88 | 89 | 90 | 91 | 92 | 93 | 94 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 117 | 118 | 119 | 120 | 121 | 122 | 123 | 124 | 125 | 126 | 127 | 128 | 129 | 130 | 131 | 132 | 133 | 134 | 135 | 136 | 137 | 138 | 139 | 140 | 141 | 142 | 143 | 144 | 145 | 146 | 147 | 148 | 149 | 150 | 151 | 152 | 153 | 154 | 155 | 156 | 157 | 158 | 159 | 160 | 161 | 162 | 163 | 164 | 165 | 166 | 167 | 168 | 169 | 170 | 171 | 172 | 173 | 174 | 175 | 176 | 177 | 178 | 179 | 180 | 181 | 182 | 183 | 184 | 185 | 186 | 187 | 188 | 189 | 190 | 191 | 192 | 193 | 194 | 195 | 196 | 197 | 198 | 199 | 200 | 201 | 202 | 203 | 204 | 205 | 206 | 207 | 208 | 209 | 210 | 211 | 212 | 213 | 214 | 215 | 216 | 217 | 218 | 219 | 220 | 221 | 222 | 223 | 224 | 225 | 226 | 227 | 228 | 229 | 230 | 231 | 232 | 233 | 234 | 235 | 236 | 237 | 238 | 239 | 240 | 241 | 242 | 243 | 244 | 245 | 246 | 247 | 248 | 249 | 250 | 251 | 252 | 253 | 254 | 255 | 256 | 257 | 258 | 259 | 260 | 261 | 262 | 263 | 264 | 265 | 266 | 267 | 268 | 269 | 270 | 271 | 272 | 273 | 274 | 275 | 276 | 277 | 278 | 279 | 280 | 281 | 282 | 283 | 284 | 285 | 286 | 287 | 288 | 289 | 290 | 291 | 292 | 293 | 294 | 295 | 296 | 297 | 298 | 299 | 300 | 301 | 302 | 303 | 304 | 305 | 306 | 307 | 308 | 309 | 310 | 311 | 312 | 313 | 314 | 315 | 316 | 317 | 318 | 319 | 320 | 321 | 322 | 323 | 324 | 325 | 326 | 327 | 328 | 329 | 330 | 331 | 332 | 333 | 334 | 335 | 336 | 337 | 338 | 339 | 340 | 341 | 342 | 343 | 344 | 345 | 346 | 347 | 348 | 349 | 350 | 351 | 352 | 353 | 354 | 355 | 356 | 357 | 358 | 359 | 360 | 361 | 362 | 363 | 364 | 365 | 366 | 367 | 368 | 369 | 370 | 371 | 372 | 373 | 374 | 375 | 376 | 377 | 378 | 379 | 380 | 381 | 382 | 383 | 384 | 385 | 386 | 387 | 388 | 389 | 390 | 391 | 392 | 393 | 394 | 395 | 396 | 397 | 398 | 399 | 400 | 401 | 402 | 403 | 404 | 405 | 406 | 407 | 408 | 409 | 410 | 411 | 412 | 413 | 414 | 415 | 416 | 417 | 418 | 419 | 420 | 421 | 422 | 423 | 424 | 425 | 426 | 427 | 428 | 429 | 430 | 431 | 432 | 433 | 434 | 435 | 436 | 437 | 438 | 439 | 440 | 441 | 442 | 443 | 444 | 445 | 446 | 447 | 448 | 449 | 450 | 451 | 452 | 453 | 454 | 455 | 456 | 457 | 458 | 459 | 460 | 461 | 462 | 463 | 464 | 465 | 466 | 467 | 468 | 469 | 470 | 471 | 472 | 473 | 474 | 475 | 476 | 477 | 478 | 479 | 480 | 481 | 482 | 483 | 484 | 485 | 486 | 487 | 488 | 489 | 490 | 491 | 492 | 493 | 494 | 495 | 496 | 497 | 498 | 499 | 500 | 501 | 502 | 503 | 504 | 505 | 506 | 507 | 508 | 509 | 510 | 511 | 512 | 513 | 514 | 515 | 516 | 517 | 518 | 519 | 520 | 521 | 522 | 523 | 524 | 525 | 526 | 527 | 528 | 529 | 530 | 531 | 532 | 533 | 534 | 535 | 536 | 537 | 538 | 539 | 540 | 541 | 542 | 543 | 544 | 545 | 546 | 547 | 548 | 549 | 550 | 551 | 552 | 553 | 554 | 555 | 556 | 557 | 558 | 559 | 560 | 561 | 562 | 563 | 564 | 565 | 566 | 567 | 568 | 569 | 570 | 571 | 572 | 573 | 574 | 575 | 576 | 577 | 578 | 579 | 580 | 581 | 582 | 583 | 584 | 585 | 586 | 587 | 588 | 589 | 590 | 591 | 592 | 593 | 594 | 595 | 596 | 597 | 598 | 599 | 600 | 601 | 602 | 603 | 604 | 605 | 606 | 607 | 608 | 609 | 610 | 611 | 612 | 613 | 614 | 615 | 616 | 617 | 618 | 619 | 620 | 621 | 622 | 623 | 624 | 625 | 626 | 627 | 628 | 629 | 630 | 631 | 632 | 633 | 634 | 635 | 636 | 637 | 638 | 639 | 640 | 641 | 642 | 643 | 644 | 645 | 646 | 647 | 648 | 649 | 650 | 651 | 652 | 653 | 654 | 655 | 656 | 657 | 658 | 659 | 660 | 661 | 662 | 663 | 664 | 665 | 666 | 667 | 668 | 669 | 670 | 671 | 672 | 673 | 674 | 675 | 676 | 677 | 678 | 679 | 680 | 681 | 682 | 683 | 684 | 685 | 686 | 687 | 688 | 689 | 690 | 691 | 692 | 693 | 694 | 695 | 696 | 697 | 698 | 699 | 700 | 701 | 702 | 703 | 704 | 705 | 706 | 707 | 708 | 709 | 710 | 711 | 712 | 713 | 714 | 715 | 716 | 717 | 718 | 719 | 720 | 721 | 722 | 723 | 724 | 725 | 726 | 727 | 728 | 729 | 730 | 731 | 732 | 733 | 734 | 735 | 736 | 737 | 738 | 739 | 740 | 741 | 742 | 743 | 744 | 745 | 746 | 747 | 748 | 749 | 750 | 751 | 752 | 753 | 754 | 755 | 756 | 757 | 758 | 759 | 760 | 761 | 762 | 763 | 764 | 765 | 766 | 767 | 768 | 769 | 770 | 771 | 772 | 773 | 774 | 775 | 776 | 777 | 778 | 779 | 780 | 781 | 782 | 783 | 784 | 785 | 786 | 787 | 788 | 789 | 790 | 791 | 792 | 793 | 794 | 795 | 796 | 797 | 798 | 799 | 800 | 801 | 802 | 803 | 804 | 805 | 806 | 807 | 808 | 809 | 810 | 811 | 812 | 813 | 814 | 815 | 816 | 817 | 818 | 819 | 820 | 821 | 822 | 823 | 824 | 825 | 826 | 827 | 828 | 829 | 830 | 831 | 832 | 833 | 834 | 835 | 836 | 837 | 838 | 839 | 840 | 841 | 842 | 843 | 844 | 845 | 846 | 847 | 848 | 849 | 850 | 851 | 852 | 853 | 854 | 855 | 856 | 857 | 858 | 859 | 860 | 861 | 862 | 863 | 864 | 865 | 866 | 867 | 868 | 869 | 870 | 871 | 872 | 873 | 874 | 875 | 876 | 877 | 878 | 879 | 880 | 881 | 882 | 883 | 884 | 885 | 886 | 887 | 888 | 889 | 890 | 891 | 892 | 893 | 894 | 895 | 896 | 897 | 898 | 899 | 900 | 901 | 902 | 903 | 904 | 905 | 906 | 907 | 908 | 909 | 910 | 911 | 912 | 913 | 914 | 915 | 916 | 917 | 918 | 919 | 920 | 921 | 922 | 923 | 924 | 925 | 926 | 927 | 928 | 929 | 930 | 931 | 932 | 933 | 934 | 935 | 936 | 937 | 938 | 939 | 940 | 941 | 942 | 943 | 944 | 945 | 946 | 947 | 948 | 949 | 950 | 951 | 952 | 953 | 954 | 955 | 956 | 957 | 958 | 959 | 960 | 961 | 962 | 963 | 964 | 965 | 966 | 967 | 968 | 969 | 970 | 971 | 972 | 973 | 974 | 975 | 976 | 977 | 978 | 979 | 980 | 981 | 982 | 983 | 984 | 985 | 986 | 987 | 988 | 989 | 990 | 991 | 992 | 993 | 994 | 995 | 996 | 997 | 998 | 999 | 1000 | 1001 | 1002 | 1003 | 1004 | 1005 | 1006 | 1007 | 1008 | 1009 | 1010 | 1011 | 1012 | 1013 | 1014 | 1015 | 1016 | 1017 | 1018 | 1019 | 1020 | 1021 | 1022 | 1023 | 1024 | 1025 | 1026 | 1027 | 1028 | 1029 | 1030 | 1031 | 1032 | 1033 | 1034 | 1035 | 1036 | 1037 | 1038 | 1039 | 1040 | 1041 | 1042 | 1043 | 1044 | 1045 | 1046 | 1047 | 1048 | 1049 | 1050 | 1051 | 1052 | 1053 | 1054 | 1055 | 1056 | 1057 | 1058 | 1059 | 1060 | 1061 | 1062 | 1063 | 1064 | 1065 | 1066 | 1067 | 1068 | 1069 | 1070 | 1071 | 1072 | 1073 | 1074 | 1075 | 1076 | 1077 | 1078 | 1079 | 1080 | 1081 | 1082 | 1083 | 1084 | 1085 | 1086 | 1087 | 1088 | 1089 | 1090 | 1091 | 1092 | 1093 | 1094 | 1095 | 1096 | 1097 | 1098 | 1099 | 1100 | 1101 | 1102 | 1103 | 1104 | 1105 | 1106 | 1107 | 1108 | 1109 | 1110 | 1111 | 1112 | 1113 | 1114 | 1115 | 1116 | 1117 | 1118 | 1119 | 1120 | 1121 | 1122 | 1123 | 1124 | 1125 | 1126 | 1127 | 1128 | 1129 | 1130 | 1131 | 1132 | 1133 | 1134 | 1135 | 1136 | 1137 | 1138 | 1139 | 1140 | 1141 | 1142 | 1143 | 1144 | 1145 | 1146 | 1147 | 1148 | 1149 | 1150 | 1151 | 1152 | 1153 | 1154 | 1155 | 1156 | 1157 | 1158 | 1159 | 1160 | 1161 | 1162 | 1163 | 1164 | 1165 | 1166 | 1167 | 1168 | 1169 | 1170 | 1171 | 1172 | 1173 | 1174 | 1175 | 1176 | 1177 | 1178 | 1179 | 1180 | 1181 | 1182 | 1183 | 1184 | 1185 | 1186 | 1187 | 1188 | 1189 | 1190 | 1191 | 1192 | 1193 | 1194 | 1195 | 1196 | 1197 | 1198 | 1199 | 1200 | 1201 | 1202 | 1203 | 1204 | 1205 | 1206 | 1207 | 1208 | 1209 | 1210 | 1211 | 1212 | 1213 | 1214 | 1215 | 1216 | 1217 | 1218 | 1219 | 1220 | 1221 | 1222 | 1223 | 1224 | 1225 | 1226 | 1227 | 1228 | 1229 | 1230 | 1231 | 1232 | 1233 | 1234 | 1235 | 1236 | 1237 | 1238 | 1239 | 1240 | 1241 | 1242 | 1243 | 1244 | 1245 | 1246 | 1247 | 1248 | 1249 | 1250 | 1251 | 1252 | 1253 | 1254 | 1255 | 1256 | 1257 | 1258 | 1259 | 1260 | 1261 | 1262 | 1263 | 1264 | 1265 | 1266 | 1267 | 1268 | 1269 | 1270 | 1271 | 1272 | 1273 | 1274 | 1275 | 1276 | 1277 | 1278 | 1279 | 1280 | 1281 | 1282 | 1283 | 1284 | 1285 | 1286 | 1287 | 1288 | 1289 | 1290 | 1291 | 1292 | 1293 | 1294 | 1295 | 1296 | 1297 | 1298 | 1299 | 1300 | 1301 | 1302 | 1303 | 1304 | 1305 | 1306 | 1307 | 1308 | 1309 | 1310 | 1311 | 1312 | 1313 | 1314 | 1315 | 1316 | 1317 | 1318 | 1319 | 1320 | 1321 | 1322 | 1323 | 1324 | 1325 | 1326 | 1327 | 1328 | 1329 | 1330 | 1331 | 1332 | 1333 | 1334 | 1335 | 1336 | 1337 | 1338 | 1339 | 1340 | 1341 | 1342 | 1343 | 1344 | 1345 | 1346 | 1347 | 1348 | 1349 | 1350 | 1351 | 1352 | 1353 | 1354 | 1355 | 1356 | 1357 | 1358 | 1359 | 1360 | 1361 | 1362 | 1363 | 1364 | 1365 | 1366 | 1367 | 1368 | 1369 | 1370 | 1371 | 1372 | 1373 | 1374 | 1375 | 1376 | 1377 | 1378 | 1379 | 1380 | 1381 | 1382 | 1383 | 1384 | 1385 | 1386 | 1387 | 1388 | 1389 | 1390 | 1391 | 1392 | 1393 | 1394 | 1395 | 1396 | 1397 | 1398 | 1399 | 1400 | 1401 | 1402 | 1403 | 1404 | 1405 | 1406 | 1407 | 1408 | 1409 | 1410 | 1411 | 1412 | 1413 | 1414 | 1415 | 1416 | 1417 | 1418 | 1419 | 1420 | 1421 | 1422 | 1423 | 1424 | 1425 | 1426 | 1427 | 1428 | 1429 | 1430 | 1431 | 1432 | 1433 | 1434 | 1435 | 1436 | 1437 | 1438 | 1439 | 1440 | 1441 | 1442 | 1443 | 1444 | 1445 | 1446 | 1447 | 1448 | 1449 | 1450 | 1451 | 1452 | 1453 | 1454 | 1455 | 1456 | 1457 | 1458 | 1459 | 1460 | 1461 | 1462 | 1463 | 1464 | 1465 | 1466 | 1467 | 1468 | 1469 | 1470 | 1471 | 1472 | 1473 | 1474 | 1475 | 1476 | 1477 | 1478 | 1479 | 1480 | 1481 | 1482 | 1483 | 1484 | 1485 | 1486 | 1487 | 1488 | 1489 | 1490 | 1491 | 1492 | 1493 | 1494 | 1495 | 1496 | 1497 | 1498 | 1499 | 1500 | 1501 | 1502 | 1503 | 1504 | 1505 | 1506 | 1507 | 1508 | 1509 | 1510 | 1511 | 1512 | 1513 | 1514 | 1515 | 1516 | 1517 | 1518 | 1519 | 1520 | 1521 | 1522 | 1523 | 1524 | 1525 | 1526 | 1527 | 1528 | 1529 | 1530 | 1531 | 1532 | 1533 | 1534 | 1535 | 1536 | 1537 | 1538 | 1539 | 1540 | 1541 | 1542 | 1543 | 1544 | 1545 | 1546 | 1547 | 1548 | 1549 | 1550 | 1551 | 1552 | 1553 | 1554 | 1555 | 1556 | 1557 | 1558 | 1559 | 1560 | 1561 | 1562 | 1563 | 1564 | 1565 | 1566 | 1567 | 1568 | 1569 | 1570 | 1571 | 1572 | 1573 | 1574 | 1575 | 1576 | 1577 | 1578 | 1579 | 1580 | 1581 | 1582 | 1583 | 1584 | 1585 | 1586 | 1587 | 1588 | 1589 | 1590 | 1591 | 1592 | 1593 | 1594 | 1595 | 1596 | 1597 | 1598 | 1599 | 1600 | 1601 | 1602 | 1603 | 1604 | 1605 | 1606 | 1607 | 1608 | 1609 | 1610 | 1611 | 1612 | 1613 | 1614 | 1615 | 1616 | 1617 | 1618 | 1619 | 1620 | 1621 | 1622 | 1623 | 1624 | 1625 | 1626 | 1627 | 1628 | 1629 | 1630 | 1631 | 1632 | 1633 | 1634 | 1635 | 1636 | 1637 | 1638 | 1639 | 1640 | 1641 | 1642 | 1643 | 1644 | 1645 | 1646 | 1647 | 1648 | 1649 | 1650 | 1651 | 1652 | 1653 | 1654 | 1655 | 1656 | 1657 | 1658 | 1659 | 1660 | 1661 | 1662 | 1663 | 1664 | 1665 | 1666 | 1667 | 1668 | 1669 | 1670 | 1671 | 1672 | 1673 | 1674 | 1675 | 1676 | 1677 | 1678 | 1679 | 1680 | 1681 | 1682 | 1683 | 1684 | 1685 | 1686 | 1687 | 1688 | 1689 | 1690 | 1691 | 1692 | 1693 | 1694 | 1695 | 1696 | 1697 | 1698 | 1699 | 1700 | 1701 | 1702 | 1703 | 1704 | 1705 | 1706 | 1707 | 1708 | 1709 | 1710 | 1711 | 1712 | 1713 | 1714 | 1715 | 1716 | 1717 | 1718 | 1719 | 1720 | 1721 | 1722 | 1723 | 1724 | 1725 | 1726 | 1727 | 1728 | 1729 | 1730 | 1731 | 1732 | 1733 | 1734 | 1735 | 1736 | 1737 | 1738 | 1739 | 1740 | 1741 | 1742 | 1743 | 1744 | 1745 | 1746 | 1747 | 1748 | 1749 | 1750 | 1751 | 1752 | 1753 | 1754 | 1755 | 1756 | 1757 | 1758 | 1759 | 1760 | 1761 | 1762 | 1763 | 1764 | 1765 | 1766 | 1767 | 1768 | 1769 | 1770 | 1771 | 1772 | 1773 | 1774 | 1775 | 1776 | 1777 | 1778 | 1779 | 1780 | 1781 | 1782 | 1783 | 1784 | 1785 | 1786 | 1787 | 1788 | 1789 | 1790 | 1791 | 1792 | 1793 | 1794 | 1795 | 1796 | 1797 | 1798 | 1799 | 1800 | 1801 | 1802 | 1803 | 1804 | 1805 | 1806 | 1807 | 1808 | 1809 | 1810 | 1811 | 1812 | 1813 | 1814 | 1815 | 1816 | 1817 | 1818 | 1819 | 1820 | 1821 | 1822 | 1823 | 1824 | 1825 | 1826 | 1827 | 1828 | 1829 | 1830 | 1831 | 1832 | 1833 | 1834 | 1835 | 1836 | 1837 | 1838 | 1839 | 1840 | 1841 | 1842 | 1843 | 1844 | 1845 | 1846 | 1847 | 1848 | 1849 | 1850 | 1851 | 1852 | 1853 | 1854 | 1855 | 1856 | 1857 | 1858 | 1859 | 1860 | 1861 | 1862 | 1863 | 1864 | 1865 | 1866 | 1867 | 1868 | 1869 | 1870 | 1871 | 1872 | 1873 | 1874 | 1875 | 1876 | 1877 | 1878 | 1879 | 1880 | 1881 | 1882 | 1883 | 1884 | 1885 | 1886 | 1887 | 1888 | 1889 | 1890 | 1891 | 1892 | 1893 | 1894 | 1895 | 1896 | 1897 | 1898 | 1899 | 1900 | 1901 | 1902 | 1903 | 1904 | 1905 | 1906 | 1907 | 1908 | 1909 | 1910 | 1911 | 1912 | 1913 | 1914 | 1915 | 1916 | 1917 | 1918 | 1919 | 1920 | 1921 | 1922 | 1923 | 1924 | 1925 | 1926 | 1927 | 1928 | 1929 | 1930 | 1931 | 1932 | 1933 | 1934 | 1935 | 1936 | 1937 | 1938 | 1939 | 1940 | 1941 | 1942 | 1943 | 1944 | 1945 | 1946 | 1947 | 1948 | 1949 | 1950 | 1951 | 1952 | 1953 | 1954 | 1955 | 1956 | 1957 | 1958 | 1959 | 1960 | 1961 | 1962 | 1963 | 1964 | 1965 | 1966 | 1967 | 1968 | 1969 | 1970 | 1971 | 1972 | 1973 | 1974 | 1975 | 1976 | 1977 | 1978 | 1979 | 1980 | 1981 | 1982 | 1983 | 1984 | 1985 | 1986 | 1987 | 1988 | 1989 | 1990 | 1991 | 1992 | 1993 | 1994 | 1995 | 1996 | 1997 | 1998 | 1999 | 2000 | 2001 | 2002 | 2003 | 2004 | 2005 | 2006 | 2007 | 2008 | 2009 | 2010 | 2011 | 2012 | 2013 | 2014 | 2015 | 2016 | 2017 | 2018 | 2019 | 2020 | 2021 | 2022 | 2023 | 2024 | 2025 | 2026 | 2027 | 2028 | 2029 | 2030 | 2031 | 2032 | 2033 | 2034 | 2035 | 2036 | 2037 | 2038 | 2039 | 2040 | 2041 | 2042 | 2043 | 2044 | 2045 | 2046 | 2047
1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 1 | 0 | 2 | 0 | 3 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 2 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 2 | 1 | 1 | 0 | 1 | 2 | 1 | 2 | 1 | 1 | 0 | 2 | 0 | 2 | 1 | 1 | 1 | 1 | 0 | 3 | 2 | 1 | 3 | 1 | 1 | 1 | 2 | 1 | 1 | 2 | 2 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 2 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 2 | 0 | 2 | 2 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 0 | 1 | 0 | 2 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 2 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 2 | 0 | 1 | 2 | 2 | 0 | 2 | 3 | 2 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 2 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 2 | 1 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 2 | 3 | 1 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 2 | 2 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 2 | 1 | 0 | 1 | 1 | 1 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 2 | 1 | 0 | 0 | 3 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 2 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 2 | 352 | 44 | 27 | 35 | 38 | 36 | 44 | 51 | 61 | 76 | 71 | 92 | 106 | 124 | 145 | 170 | 185 | 222 | 268 | 320 | 384 | 466 | 553 | 616 | 728 | 839 | 915 | 1032 | 1166 | 1280 | 1369 | 1401 | 1525 | 1540 | 1507 | 1535 | 1576 | 1588 | 1629 | 1649 | 1644 | 1653 | 1687 | 1733 | 1704 | 1639 | 1603 | 1624 | 1630 | 1679 | 1714 | 1672 | 1620 | 1690 | 1717 | 1624 | 1566 | 1576 | 1610 | 1637 | 1634 | 1601 | 1618 | 1629 | 1698 | 1785 | 1644 | 1603 | 1591 | 1624 | 1618 | 1604 | 1609 | 1673 | 1690 | 1707 | 1711 | 1651 | 1707 | 1726 | 1743 | 1741 | 1747 | 1730 | 1714 | 1733 | 1775 | 1714 | 1671 | 1698 | 1695 | 1653 | 1634 | 1623 | 1634 | 1662 | 1686 | 1675 | 1637 | 1604 | 1576 | 1554 | 1555 | 1566 | 1526 | 1533 | 1534 | 1532 | 1579 | 1565 | 1552 | 1547 | 1528 | 1516 | 1538 | 1578 | 1581 | 1568 | 1538 | 1477 | 1432 | 1402 | 1434 | 1421 | 1384 | 1381 | 1399 | 1413 | 1426 | 1440 | 1444 | 1452 | 1430 | 1427 | 1416 | 1408 | 1439 | 1449 | 1427 | 1438 | 1440 | 1425 | 1455 | 1477 | 1474 | 1455 | 1439 | 1473 | 1505 | 1488 | 1468 | 1488 | 1480 | 1464 | 1511 | 1560 | 1581 | 1590 | 1601 | 1592 | 1564 | 1598 | 1648 | 1671 | 1698 | 1736 | 1790 | 1840 | 1830 | 1868 | 1991 | 2019 | 2001 | 1993 | 2056 | 2085 | 2145 | 2161 | 2167 | 2162 | 2229 | 2236 | 2304 | 2309 | 2233 | 2213 | 2228 | 2214 | 2194 | 2180 | 2164 | 2163 | 2117 | 2080 | 2094 | 2060 | 2041 | 2026 | 2001 | 2025 | 2001 | 1961 | 1969 | 1919 | 1839 | 1870 | 1863 | 1847 | 1808 | 1779 | 1733 | 1768 | 1735 | 1756 | 1758 | 1669 | 1679 | 1617 | 1591 | 1638 | 1584 | 1629 | 1673 | 1513 | 1519 | 1549 | 1551 | 1538 | 1486 | 1503 | 1523 | 1447 | 1502 | 1459 | 1456 | 1421 | 1376 | 1366 | 1368 | 1434 | 1422 | 1369 | 1359 | 1364 | 1348 | 1328 | 1304 | 1305 | 1322 | 1308 | 1313 | 1311 | 1280 | 1284 | 1288 | 1265 | 1258 | 1261 | 1260 | 1211 | 1195 | 1193 | 1178 | 1172 | 1202 | 1190 | 1188 | 1169 | 1160 | 1138 | 1121 | 1118 | 1095 | 1105 | 1077 | 1087 | 1130 | 1096 | 1083 | 1066 | 1068 | 1036 | 1047 | 1077 | 1073 | 1051 | 1012 | 1058 | 1073 | 1015 | 969 | 968 | 993 | 982 | 972 | 1000 | 1023 | 1023 | 985 | 952 | 957 | 974 | 994 | 988 | 972 | 962 | 943 | 945 | 946 | 923 | 895 | 884 | 904 | 931 | 931 | 889 | 858 | 847 | 823 | 851 | 891 | 886 | 884 | 856 | 814 | 855 | 876 | 830 | 776 | 768 | 781 | 788 | 779 | 771 | 756 | 727 | 728 | 733 | 731 | 730 | 717 | 718 | 727 | 699 | 668 | 660 | 659 | 655 | 650 | 646 | 650 | 649 | 642 | 642 | 644 | 645 | 633 | 627 | 632 | 624 | 600 | 581 | 580 | 592 | 610 | 626 | 622 | 618 | 600 | 603 | 630 | 615 | 572 | 563 | 576 | 575 | 590 | 602 | 592 | 559 | 527 | 544 | 567 | 562 | 567 | 573 | 583 | 557 | 563 | 547 | 561 | 542 | 526 | 549 | 563 | 559 | 533 | 518 | 520 | 525 | 529 | 527 | 513 | 508 | 518 | 537 | 539 | 532 | 532 | 528 | 526 | 519 | 532 | 518 | 524 | 532 | 518 | 523 | 537 | 528 | 505 | 496 | 506 | 486 | 510 | 522 | 532 | 542 | 538 | 516 | 525 | 509 | 487 | 527 | 549 | 527 | 531 | 544 | 578 | 572 | 566 | 563 | 555 | 590 | 627 | 663 | 666 | 681 | 729 | 732 | 779 | 832 | 876 | 930 | 1015 | 1113 | 1181 | 1276 | 1388 | 1488 | 1601 | 1746 | 1907 | 2073 | 2318 | 2518 | 2708 | 3002 | 3277 | 3502 | 3831 | 4122 | 4468 | 4887 | 5202 | 5587 | 5851 | 6309 | 6895 | 7281 | 7633 | 8205 | 8722 | 8873 | 9396 | 9806 | 10089 | 10458 | 10823 | 11268 | 11612 | 12014 | 12107 | 12245 | 12397 | 12488 | 12551 | 12560 | 12483 | 12474 | 12338 | 12128 | 11933 | 11731 | 11348 | 10943 | 10563 | 10207 | 9835 | 9290 | 8810 | 8497 | 8100 | 7618 | 7109 | 6573 | 6082 | 5620 | 5163 | 4784 | 4437 | 4044 | 3688 | 3366 | 3031 | 2778 | 2523 | 2283 | 2036 | 1844 | 1668 | 1543 | 1438 | 1322 | 1211 | 1089 | 1058 | 1020 | 935 | 897 | 862 | 801 | 752 | 754 | 733 | 702 | 668 | 670 | 698 | 695 | 664 | 659 | 662 | 659 | 667 | 674 | 653 | 629 | 633 | 645 | 646 | 642 | 648 | 647 | 636 | 627 | 629 | 652 | 634 | 604 | 621 | 626 | 618 | 633 | 662 | 666 | 642 | 638 | 636 | 630 | 613 | 598 | 602 | 613 | 583 | 609 | 628 | 648 | 612 | 578 | 575 | 586 | 595 | 600 | 593 | 584 | 574 | 572 | 595 | 568 | 565 | 571 | 568 | 555 | 557 | 551 | 556 | 554 | 533 | 527 | 533 | 535 | 543 | 540 | 535 | 537 | 525 | 521 | 535 | 566 | 567 | 569 | 554 | 534 | 554 | 583 | 547 | 549 | 547 | 578 | 578 | 541 | 550 | 561 | 547 | 540 | 567 | 619 | 597 | 542 | 510 | 521 | 545 | 585 | 587 | 579 | 558 | 541 | 563 | 572 | 564 | 588 | 591 | 594 | 622 | 632 | 580 | 582 | 659 | 617 | 608 | 617 | 652 | 670 | 627 | 627 | 663 | 685 | 666 | 644 | 663 | 694 | 695 | 709 | 733 | 766 | 732 | 794 | 827 | 754 | 759 | 764 | 703 | 741 | 769 | 760 | 731 | 726 | 739 | 750 | 746 | 729 | 727 | 723 | 715 | 708 | 708 | 703 | 734 | 692 | 687 | 685 | 685 | 673 | 666 | 644 | 627 | 652 | 671 | 653 | 640 | 644 | 625 | 593 | 587 | 612 | 617 | 569 | 566 | 544 | 559 | 552 | 521 | 543 | 537 | 519 | 525 | 537 | 524 | 511 | 512 | 501 | 471 | 462 | 476 | 480 | 476 | 489 | 474 | 432 | 437 | 439 | 423 | 444 | 460 | 434 | 410 | 402 | 397 | 395 | 410 | 427 | 422 | 402 | 401 | 402 | 352 | 355 | 362 | 354 | 362 | 353 | 347 | 347 | 347 | 347 | 351 | 349 | 346 | 343 | 338 | 336 | 332 | 326 | 336 | 334 | 322 | 313 | 324 | 319 | 302 | 313 | 308 | 295 | 297 | 296 | 291 | 292 | 297 | 293 | 278 | 273 | 271 | 272 | 281 | 288 | 293 | 299 | 281 | 276 | 262 | 254 | 260 | 256 | 257 | 265 | 261 | 254 | 249 | 247 | 242 | 227 | 221 | 232 | 245 | 228 | 240 | 258 | 222 | 219 | 223 | 226 | 213 | 211 | 214 | 205 | 195 | 197 | 210 | 220 | 217 | 205 | 193 | 207 | 184 | 183 | 182 | 193 | 229 | 161 | 162 | 175 | 175 | 178 | 183 | 166 | 165 | 166 | 157 | 169 | 170 | 152 | 144 | 161 | 167 | 167 | 176 | 177 | 138 | 163 | 139 | 138 | 140 | 131 | 139 | 148 | 142 | 147 | 153 | 135 | 132 | 146 | 145 | 119 | 150 | 134 | 128 | 130 | 142 | 129 | 111 | 109 | 121 | 130 | 126 | 132 | 130 | 123 | 124 | 123 | 117 | 118 | 110 | 124 | 111 | 103 | 99 | 105 | 122 | 110 | 112 | 118 | 118 | 107 | 103 | 106 | 106 | 105 | 109 | 114 | 120 | 125 | 110 | 115 | 109 | 120 | 133 | 107 | 113 | 98 | 121 | 108 | 111 | 118 | 127 | 122 | 110 | 101 | 102 | 107 | 106 | 103 | 115 | 117 | 110 | 111 | 118 | 129 | 130 | 127 | 132 | 142 | 136 | 137 | 131 | 133 | 148 | 152 | 154 | 161 | 151 | 153 | 181 | 196 | 196 | 201 | 222 | 223 | 251 | 258 | 261 | 276 | 272 | 282 | 299 | 316 | 343 | 371 | 379 | 388 | 396 | 417 | 476 | 507 | 525 | 578 | 611 | 639 | 652 | 666 | 707 | 732 | 748 | 780 | 809 | 844 | 901 | 937 | 955 | 989 | 1000 | 985 | 1034 | 1085 | 1105 | 1144 | 1167 | 1163 | 1163 | 1221 | 1262 | 1251 | 1241 | 1261 | 1287 | 1276 | 1261 | 1287 | 1321 | 1294 | 1256 | 1248 | 1269 | 1228 | 1201 | 1243 | 1162 | 1159 | 1129 | 1074 | 1080 | 1042 | 978 | 957 | 945 | 905 | 859 | 833 | 803 | 734 | 692 | 676 | 647 | 613 | 565 | 558 | 529 | 493 | 472 | 443 | 405 | 382 | 337 | 310 | 290 | 265 | 247 | 230 | 223 | 213 | 187 | 165 | 160 | 145 | 131 | 119 | 97 | 97 | 91 | 88 | 77 | 78 | 70 | 63 | 60 | 54 | 48 | 41 | 33 | 41 | 38 | 32 | 31 | 34 | 22 | 30 | 27 | 27 | 25 | 23 | 25 | 25 | 22 | 19 | 24 | 21 | 16 | 17 | 22 | 20 | 18 | 12 | 16 | 22 | 17 | 25 | 18 | 23 | 21 | 13 | 18 | 16 | 19 | 16 | 16 | 16 | 16 | 20 | 23 | 20 | 12 | 15 | 15 | 15 | 20 | 18 | 15 | 20 | 17 | 12 | 20 | 16 | 17 | 14 | 15 | 15 | 13 | 15 | 13 | 20 | 17 | 14 | 17 | 17 | 19 | 26 | 18 | 15 | 19 | 20 | 12 | 13 | 15 | 17 | 16 | 13 | 11 | 13 | 15 | 14 | 14 | 16 | 18 | 16 | 16 | 16 | 15 | 16 | 12 | 14 | 17 | 16 | 16 | 15 | 14 | 13 | 15 | 16 | 16 | 15 | 11 | 12 | 14 | 16 | 17 | 18 | 11 | 16 | 18 | 14 | 11 | 12 | 12 | 14 | 16 | 18 | 16 | 13 | 14 | 16 | 14 | 13 | 17 | 16 | 11 | 14 | 17 | 13 | 14 | 14 | 16 | 17 | 15 | 12 | 9 | 9 | 14 | 17 | 13 | 12 | 15 | 15 | 12 | 10 | 10 | 12 | 13 | 9 | 6 | 9 | 12 | 13 | 12 | 11 | 11 | 11 | 12 | 11 | 9 | 11 | 12 | 11 | 13 | 12 | 8 | 9 | 11 | 15 | 13 | 11 | 11 | 11 | 11 | 12 | 14 | 14 | 11 | 10 | 9 | 8 | 8 | 5 | 7 | 7 | 8 | 7 | 10 | 8 | 8 | 9 | 8 | 9 | 10 | 9 | 11 | 13 | 9 | 7 | 5 | 5 | 5 | 7 | 6 | 7 | 12 | 10 | 5 | 8 | 7 | 7 | 11 | 11 | 9 | 8 | 8 | 10 | 11 | 11 | 10 | 7 | 6 | 5 | 7 | 10 | 7 | 5 | 8 | 10 | 12 | 6 | 5 | 7 | 9 | 7 | 4 | 4 | 6 | 6 | 9 | 10 | 8 | 4 | 3 | 9 | 7 | 9 | 6 | 7 | 7 | 9 | 7 | 7 | 6 | 5 | 6 | 8 | 7 | 4 | 6 | 7 | 8 | 4 | 5 | 10 | 6 | 5 | 6 | 7 | 6 | 9 | 6 | 5 | 5 | 6 | 7 | 4 | 5 | 7 | 3 | 4 | 3 | 7 | 6 | 8 | 10 | 3 | 1 | 6 | 11 | 10 | 7 | 5 | 6 | 5 | 5 | 6 | 4 | 2 | 5 | 7 | 6 | 5 | 7 | 7 | 8 | 4 | 3 | 3 | 4 | 4 | 4 | 5 | 6 | 6 | 5 | 7 | 7 | 6 | 4 | 4 | 6 | 8 | 7 | 8 | 9 | 6 | 3 | 2 | 5 | 5 | 3 | 3 | 5 | 6 | 6 | 5 | 6 | 6 | 5 | 5 | 6 | 5 | 8 | 8 | 7 | 6 | 6 | 7 | 6 | 4 | 6 | 7 | 4 | 3 | 4 | 5 | 5 | 4 | 4 | 6 | 7 | 8 | 8 | 7 | 5 | 5 | 3 | 3 | 5 | 4 | 4 | 4 | 5 | 4 | 8 | 9 | 8 | 7 | 6 | 5 | 8 | 10 | 9 | 7 | 7 | 6 | 5 | 4 | 6 | 10 | 9 | 8 | 9 | 7 | 6 | 10 | 10 | 5 | 6 | 5 | 6 | 5 | 3 | 2 | 5 | 8 | 8 | 7 | 10 | 7 | 4 | 6 | 8 | 8 | 7 | 6 | 9 | 8 | 8 | 8 | 7 | 7 | 11 | 14 | 10 | 8 | 7 | 6 | 3 | 6 | 10 | 7 | 6 | 8 | 10 | 9 | 10 | 7 | 6 | 4 | 7 | 9 | 8 | 6 | 3 | 4 | 8 | 4 | 3 | 8 | 3 | 8 | 6 | 5 | 9 | 7 | 5 | 8 | 4 | 5 | 5 | 4 | 7 | 2 | 3 | 9 | 6 | 4 | 2 | 2 | 5 | 4 | 2 | 5 | 2 | 0 | 4 | 4 | 2 | 7 | 1 | 1 | 2 | 1 | 4 | 1 | 3 | 3 | 1 | 2 | 0 | 1 | 1 | 1 | 2 | 2 | 2 | 2 | 1 | 1 | 2 | 2 | 2 | 2 | 0 | 2 | 0 | 3 | 3 | 0 | 1 | 3 | 1 | 1 | 1 | 0 | 1 | 3 | 3 | 1 | 0 | 1 | 3 | 1 | 0 | 1 | 2 | 2 | 2 | 2 | 2 | 4 | 2 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 2 | 2 | 0 | 0 | 1 | 2 | 1 | 1 | 2 | 1 | 3 | 1 | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 3 | 2 | 2 | 1 | 0 | 1 | 1 | 2 | 1 | 1 | 0 | 0 | 1 | 1 | 2 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 2 | 1 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 2 | 2 | 2 | 1 | 1 | 1 | 0 | 0 | 3 | 2 | 1 | 2 | 3 | 1 | 1 | 2 | 1 | 1 | 1 | 0 | 2 | 1 | 0 | 0 | 2 | 2 | 0 | 1 | 0 | 1 | 2 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 2 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 2 | 0 | 0 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 2 | 1 | 1 | 1 | 2 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 2 | 2 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 3 | 340 | 18 | 25 | 36 | 36 | 39 | 36 | 52 | 53 | 51 | 58 | 96 | 109 | 118 | 143 | 157 | 194 | 221 | 253 | 317 | 362 | 460 | 537 | 629 | 759 | 821 | 915 | 1062 | 1165 | 1243 | 1358 | 1453 | 1537 | 1561 | 1504 | 1540 | 1563 | 1549 | 1528 | 1584 | 1566 | 1595 | 1632 | 1681 | 1646 | 1626 | 1607 | 1594 | 1624 | 1640 | 1635 | 1649 | 1648 | 1656 | 1722 | 1647 | 1604 | 1638 | 1666 | 1664 | 1604 | 1613 | 1616 | 1583 | 1616 | 1680 | 1576 | 1615 | 1594 | 1583 | 1576 | 1607 | 1647 | 1655 | 1690 | 1706 | 1674 | 1602 | 1658 | 1694 | 1706 | 1723 | 1730 | 1719 | 1703 | 1741 | 1814 | 1775 | 1667 | 1622 | 1630 | 1653 | 1623 | 1623 | 1605 | 1609 | 1708 | 1714 | 1597 | 1551 | 1533 | 1519 | 1507 | 1486 | 1506 | 1548 | 1570 | 1565 | 1572 | 1544 | 1532 | 1571 | 1576 | 1529 | 1519 | 1550 | 1554 | 1559 | 1571 | 1517 | 1429 | 1426 | 1420 | 1408 | 1425 | 1442 | 1415 | 1422 | 1427 | 1434 | 1463 | 1437 | 1414 | 1415 | 1404 | 1406 | 1405 | 1406 | 1402 | 1415 | 1449 | 1473 | 1488 | 1465 | 1418 | 1425 | 1464 | 1470 | 1457 | 1449 | 1460 | 1484 | 1496 | 1509 | 1499 | 1487 | 1554 | 1607 | 1585 | 1542 | 1575 | 1592 | 1606 | 1653 | 1693 | 1728 | 1756 | 1816 | 1841 | 1840 | 1970 | 1971 | 1968 | 2007 | 2082 | 2165 | 2135 | 2128 | 2206 | 2202 | 2182 | 2271 | 2226 | 2189 | 2228 | 2240 | 2213 | 2186 | 2133 | 2121 | 2152 | 2164 | 2133 | 2103 | 2096 | 2089 | 2056 | 2033 | 2000 | 1931 | 1915 | 1922 | 1863 | 1826 | 1869 | 1840 | 1757 | 1800 | 1809 | 1794 | 1719 | 1718 | 1731 | 1743 | 1696 | 1684 | 1708 | 1704 | 1650 | 1683 | 1583 | 1636 | 1531 | 1630 | 1543 | 1506 | 1470 | 1554 | 1538 | 1557 | 1496 | 1438 | 1427 | 1415 | 1418 | 1423 | 1392 | 1347 | 1338 | 1393 | 1387 | 1360 | 1361 | 1368 | 1346 | 1335 | 1347 | 1359 | 1312 | 1275 | 1291 | 1294 | 1275 | 1258 | 1229 | 1209 | 1244 | 1237 | 1222 | 1204 | 1228 | 1213 | 1177 | 1149 | 1114 | 1173 | 1181 | 1137 | 1137 | 1134 | 1143 | 1127 | 1134 | 1145 | 1116 | 1086 | 1051 | 1064 | 1081 | 1010 | 1037 | 1035 | 1043 | 1026 | 1023 | 1037 | 1031 | 1023 | 1028 | 1019 | 1012 | 1012 | 1006 | 1009 | 1001 | 985 | 994 | 979 | 986 | 1004 | 984 | 967 | 959 | 950 | 956 | 955 | 955 | 953 | 937 | 923 | 923 | 933 | 922 | 913 | 918 | 929 | 890 | 860 | 859 | 843 | 844 | 834 | 837 | 842 | 841 | 828 | 810 | 827 | 828 | 829 | 773 | 744 | 773 | 785 | 763 | 736 | 738 | 736 | 716 | 703 | 695 | 684 | 680 | 687 | 693 | 683 | 668 | 646 | 636 | 645 | 646 | 643 | 644 | 633 | 613 | 620 | 627 | 618 | 612 | 605 | 613 | 606 | 589 | 595 | 589 | 591 | 622 | 628 | 597 | 597 | 597 | 613 | 588 | 599 | 581 | 592 | 571 | 575 | 581 | 613 | 572 | 597 | 579 | 566 | 563 | 551 | 532 | 559 | 563 | 563 | 538 | 538 | 528 | 529 | 524 | 490 | 503 | 539 | 548 | 552 | 549 | 516 | 506 | 522 | 526 | 537 | 530 | 517 | 538 | 539 | 518 | 495 | 481 | 512 | 529 | 512 | 519 | 505 | 514 | 483 | 496 | 500 | 539 | 511 | 516 | 547 | 525 | 514 | 494 | 509 | 520 | 536 | 508 | 536 | 544 | 535 | 550 | 520 | 558 | 556 | 567 | 584 | 566 | 593 | 589 | 612 | 660 | 666 | 703 | 726 | 817 | 842 | 865 | 929 | 1015 | 1083 | 1189 | 1270 | 1336 | 1468 | 1610 | 1760 | 1891 | 2079 | 2267 | 2458 | 2712 | 2990 | 3221 | 3519 | 3823 | 4191 | 4526 | 4858 | 5117 | 5569 | 5920 | 6410 | 6780 | 7248 | 7502 | 8189 | 8684 | 9030 | 9353 | 9743 | 10137 | 10514 | 10859 | 11157 | 11494 | 11784 | 11933 | 12146 | 12313 | 12350 | 12439 | 12584 | 12502 | 12337 | 12362 | 12217 | 11935 | 11654 | 11418 | 11061 | 10605 | 10177 | 9733 | 9277 | 8801 | 8316 | 7812 | 7297 | 6835 | 6374 | 5939 | 5505 | 5083 | 4686 | 4220 | 3877 | 3575 | 3187 | 2857 | 2675 | 2429 | 2200 | 2053 | 1860 | 1659 | 1533 | 1379 | 1252 | 1162 | 1049 | 977 | 966 | 895 | 862 | 814 | 770 | 764 | 754 | 736 | 692 | 688 | 688 | 651 | 653 | 652 | 648 | 655 | 632 | 613 | 634 | 645 | 634 | 628 | 634 | 657 | 654 | 622 | 601 | 600 | 612 | 612 | 619 | 637 | 622 | 601 | 602 | 595 | 622 | 644 | 626 | 629 | 645 | 611 | 614 | 625 | 609 | 588 | 593 | 586 | 585 | 589 | 599 | 582 | 585 | 565 | 567 | 568 | 523 | 543 | 592 | 605 | 557 | 556 | 585 | 595 | 590 | 555 | 563 | 586 | 561 | 518 | 519 | 546 | 551 | 546 | 542 | 533 | 524 | 524 | 545 | 564 | 562 | 557 | 551 | 553 | 560 | 533 | 550 | 551 | 531 | 547 | 552 | 547 | 535 | 534 | 529 | 539 | 535 | 533 | 524 | 556 | 544 | 579 | 529 | 560 | 523 | 578 | 559 | 579 | 552 | 549 | 577 | 577 | 556 | 568 | 581 | 558 | 599 | 604 | 605 | 630 | 638 | 631 | 631 | 613 | 609 | 647 | 631 | 664 | 657 | 655 | 672 | 695 | 691 | 690 | 725 | 735 | 724 | 685 | 718 | 777 | 753 | 765 | 770 | 726 | 739 | 740 | 762 | 733 | 707 | 687 | 710 | 725 | 726 | 736 | 715 | 703 | 725 | 706 | 711 | 725 | 713 | 736 | 693 | 692 | 701 | 664 | 664 | 673 | 648 | 651 | 651 | 640 | 640 | 618 | 580 | 592 | 609 | 588 | 606 | 583 | 573 | 560 | 555 | 562 | 547 | 527 | 531 | 516 | 498 | 495 | 487 | 488 | 496 | 489 | 485 | 487 | 476 | 445 | 436 | 451 | 443 | 454 | 422 | 409 | 425 | 415 | 422 | 428 | 429 | 419 | 401 | 399 | 399 | 401 | 405 | 396 | 393 | 388 | 382 | 385 | 376 | 352 | 366 | 369 | 355 | 367 | 380 | 372 | 354 | 345 | 338 | 334 | 339 | 344 | 337 | 332 | 343 | 340 | 316 | 335 | 355 | 340 | 312 | 285 | 288 | 295 | 298 | 287 | 280 | 282 | 289 | 286 | 274 | 273 | 290 | 287 | 264 | 256 | 267 | 275 | 290 | 269 | 268 | 286 | 279 | 268 | 259 | 257 | 261 | 263 | 259 | 247 | 230 | 220 | 220 | 235 | 240 | 238 | 239 | 239 | 231 | 216 | 219 | 236 | 220 | 214 | 216 | 221 | 211 | 197 | 209 | 219 | 210 | 210 | 213 | 191 | 184 | 179 | 176 | 181 | 197 | 171 | 185 | 170 | 169 | 174 | 181 | 180 | 164 | 152 | 158 | 169 | 165 | 155 | 170 | 163 | 154 | 170 | 151 | 155 | 165 | 155 | 136 | 123 | 144 | 136 | 137 | 145 | 134 | 138 | 146 | 133 | 128 | 128 | 133 | 131 | 140 | 114 | 113 | 139 | 129 | 133 | 118 | 122 | 128 | 124 | 126 | 123 | 120 | 119 | 119 | 115 | 114 | 122 | 125 | 116 | 111 | 135 | 132 | 116 | 102 | 117 | 111 | 126 | 135 | 118 | 100 | 106 | 119 | 113 | 109 | 114 | 118 | 125 | 123 | 104 | 107 | 109 | 119 | 107 | 113 | 107 | 116 | 103 | 111 | 115 | 110 | 104 | 110 | 117 | 120 | 118 | 117 | 120 | 119 | 122 | 114 | 105 | 117 | 140 | 111 | 119 | 132 | 119 | 120 | 113 | 125 | 132 | 137 | 140 | 142 | 152 | 155 | 158 | 176 | 184 | 190 | 189 | 198 | 203 | 225 | 237 | 239 | 244 | 264 | 289 | 309 | 307 | 313 | 349 | 380 | 392 | 401 | 418 | 456 | 498 | 535 | 531 | 547 | 596 | 620 | 646 | 647 | 668 | 715 | 744 | 781 | 817 | 846 | 886 | 924 | 944 | 958 | 1000 | 1046 | 1047 | 1099 | 1162 | 1179 | 1175 | 1170 | 1230 | 1233 | 1231 | 1237 | 1233 | 1241 | 1246 | 1259 | 1268 | 1272 | 1258 | 1245 | 1233 | 1238 | 1254 | 1190 | 1160 | 1187 | 1183 | 1145 | 1109 | 1087 | 1058 | 1021 | 973 | 941 | 924 | 895 | 856 | 814 | 782 | 766 | 741 | 705 | 631 | 603 | 587 | 538 | 501 | 480 | 456 | 433 | 421 | 395 | 355 | 332 | 310 | 276 | 249 | 225 | 202 | 187 | 192 | 179 | 165 | 146 | 130 | 127 | 117 | 107 | 104 | 85 | 78 | 74 | 59 | 52 | 50 | 51 | 48 | 37 | 35 | 37 | 31 | 22 | 31 | 29 | 27 | 32 | 28 | 26 | 22 | 28 | 24 | 24 | 22 | 19 | 21 | 19 | 17 | 21 | 18 | 15 | 19 | 20 | 19 | 23 | 14 | 18 | 24 | 20 | 20 | 23 | 16 | 13 | 17 | 16 | 16 | 18 | 17 | 18 | 19 | 17 | 15 | 18 | 16 | 22 | 17 | 9 | 16 | 20 | 11 | 25 | 21 | 18 | 18 | 18 | 14 | 16 | 17 | 15 | 15 | 16 | 18 | 13 | 22 | 18 | 14 | 19 | 13 | 17 | 15 | 17 | 12 | 14 | 15 | 16 | 17 | 19 | 18 | 14 | 11 | 8 | 10 | 15 | 13 | 17 | 12 | 15 | 15 | 11 | 9 | 13 | 13 | 11 | 13 | 12 | 15 | 18 | 19 | 19 | 14 | 13 | 13 | 16 | 12 | 14 | 18 | 15 | 18 | 17 | 24 | 17 | 14 | 13 | 15 | 16 | 14 | 13 | 14 | 13 | 13 | 12 | 12 | 13 | 11 | 14 | 16 | 14 | 15 | 11 | 10 | 10 | 9 | 11 | 12 | 12 | 13 | 14 | 12 | 9 | 11 | 15 | 14 | 12 | 9 | 12 | 15 | 15 | 12 | 9 | 8 | 12 | 15 | 12 | 9 | 8 | 8 | 9 | 11 | 11 | 12 | 12 | 12 | 13 | 8 | 10 | 13 | 10 | 8 | 10 | 12 | 15 | 14 | 9 | 11 | 10 | 8 | 8 | 8 | 10 | 12 | 11 | 9 | 9 | 9 | 10 | 12 | 8 | 5 | 6 | 8 | 12 | 12 | 12 | 10 | 8 | 7 | 8 | 10 | 8 | 8 | 7 | 5 | 7 | 10 | 11 | 8 | 6 | 9 | 7 | 5 | 3 | 5 | 8 | 9 | 8 | 6 | 5 | 4 | 5 | 6 | 6 | 8 | 6 | 7 | 8 | 9 | 7 | 5 | 7 | 9 | 5 | 5 | 7 | 7 | 7 | 6 | 6 | 8 | 6 | 5 | 5 | 8 | 7 | 4 | 3 | 9 | 5 | 7 | 6 | 6 | 3 | 7 | 9 | 9 | 7 | 5 | 5 | 6 | 4 | 6 | 7 | 3 | 4 | 7 | 2 | 6 | 5 | 4 | 3 | 3 | 7 | 5 | 5 | 6 | 7 | 7 | 4 | 5 | 4 | 4 | 8 | 6 | 6 | 2 | 7 | 5 | 7 | 7 | 7 | 3 | 3 | 8 | 6 | 6 | 4 | 5 | 5 | 6 | 5 | 7 | 7 | 4 | 3 | 6 | 3 | 5 | 6 | 7 | 6 | 7 | 4 | 4 | 3 | 4 | 4 | 6 | 7 | 5 | 4 | 6 | 4 | 2 | 5 | 7 | 5 | 6 | 5 | 6 | 5 | 6 | 5 | 5 | 5 | 6 | 5 | 4 | 4 | 7 | 7 | 4 | 3 | 4 | 5 | 8 | 11 | 8 | 6 | 6 | 5 | 4 | 4 | 4 | 3 | 3 | 4 | 6 | 6 | 5 | 6 | 7 | 6 | 5 | 2 | 5 | 9 | 6 | 5 | 6 | 4 | 5 | 7 | 5 | 4 | 5 | 7 | 8 | 6 | 5 | 7 | 8 | 5 | 4 | 5 | 8 | 11 | 11 | 9 | 9 | 8 | 7 | 6 | 8 | 7 | 5 | 6 | 10 | 12 | 7 | 7 | 7 | 6 | 5 | 6 | 6 | 5 | 6 | 6 | 7 | 10 | 9 | 6 | 8 | 9 | 12 | 10 | 6 | 6 | 9 | 7 | 6 | 7 | 6 | 6 | 5 | 4 | 5 | 9 | 8 | 5 | 9 | 7 | 5 | 6 | 8 | 9 | 7 | 8 | 8 | 7 | 5 | 6 | 8 | 4 | 3 | 3 | 3 | 2 | 2 | 3 | 5 | 8 | 5 | 6 | 9 | 4 | 3 | 10 | 5 | 7 | 6 | 5 | 3 | 1 | 6 | 4 | 1 | 5 | 4 | 2 | 5 | 4 | 1 | 2 | 0 | 2 | 1 | 4 | 1 | 1 | 1 | 3 | 2 | 1 | 3 | 3 | 5 | 2 | 0 | 3 | 5 | 1 | 4 | 2 | 1 | 0 | 3 | 1 | 0 | 2 | 1 | 2 | 3 | 1 | 0 | 0 | 2 | 1 | 2 | 0 | 0 | 3 | 2 | 0 | 3 | 1 | 0 | 2 | 0 | 1 | 3 | 1 | 2 | 1 | 0 | 1 | 0 | 2 | 0 | 2 | 0 | 2 | 2 | 1 | 1 | 4 | 0 | 1 | 2 | 2 | 0 | 2 | 1 | 0 | 2 | 1 | 3 | 1 | 1 | 1 | 1 | 1 | 2 | 2 | 2 | 1 | 1 | 3 | 5 | 0 | 0 | 1 | 4 | 3 | 0 | 2 | 3 | 0 | 0 | 0 | 0 | 0 | 2 | 2 | 1 | 0 | 1 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 2 | 3 | 1 | 2 | 1 | 1 | 0 | 1 | 1 | 2 | 1 | 0 | 1 | 1 | 2 | 2 | 2 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 3 | 2 | 2 | 3 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 2 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 2 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 2 | 1 | 2 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 2 | 0 | 2 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 2 | 2 | 0 | 0 | 2 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 4 | 313 | 30 | 28 | 37 | 34 | 43 | 43 | 35 | 32 | 44 | 70 | 63 | 89 | 101 | 119 | 154 | 186 | 214 | 267 | 320 | 370 | 438 | 583 | 624 | 723 | 828 | 897 | 1031 | 1134 | 1210 | 1288 | 1379 | 1526 | 1516 | 1519 | 1536 | 1542 | 1581 | 1566 | 1523 | 1553 | 1601 | 1591 | 1649 | 1625 | 1587 | 1593 | 1557 | 1565 | 1607 | 1603 | 1575 | 1628 | 1654 | 1662 | 1574 | 1574 | 1582 | 1545 | 1582 | 1596 | 1597 | 1597 | 1614 | 1660 | 1710 | 1585 | 1591 | 1549 | 1577 | 1569 | 1583 | 1581 | 1574 | 1583 | 1630 | 1655 | 1599 | 1631 | 1635 | 1665 | 1717 | 1731 | 1686 | 1681 | 1675 | 1699 | 1697 | 1631 | 1610 | 1620 | 1601 | 1600 | 1579 | 1562 | 1554 | 1564 | 1606 | 1603 | 1584 | 1583 | 1550 | 1528 | 1533 | 1535 | 1555 | 1569 | 1614 | 1600 | 1518 | 1538 | 1580 | 1553 | 1513 | 1487 | 1496 | 1556 | 1552 | 1544 | 1465 | 1436 | 1425 | 1376 | 1398 | 1404 | 1400 | 1416 | 1401 | 1382 | 1402 | 1411 | 1372 | 1374 | 1392 | 1376 | 1359 | 1380 | 1422 | 1424 | 1405 | 1390 | 1401 | 1388 | 1356 | 1395 | 1437 | 1432 | 1447 | 1441 | 1425 | 1450 | 1432 | 1442 | 1495 | 1507 | 1467 | 1492 | 1519 | 1491 | 1556 | 1560 | 1545 | 1587 | 1663 | 1677 | 1659 | 1774 | 1785 | 1829 | 1839 | 1852 | 1883 | 1932 | 1994 | 2039 | 2082 | 2096 | 2072 | 2126 | 2137 | 2176 | 2160 | 2208 | 2237 | 2225 | 2211 | 2187 | 2132 | 2148 | 2159 | 2105 | 2087 | 2068 | 2028 | 2029 | 2014 | 2002 | 1999 | 1968 | 1911 | 1882 | 1880 | 1824 | 1845 | 1847 | 1843 | 1825 | 1802 | 1802 | 1735 | 1677 | 1689 | 1716 | 1699 | 1620 | 1671 | 1718 | 1625 | 1576 | 1594 | 1607 | 1576 | 1584 | 1515 | 1530 | 1511 | 1498 | 1554 | 1480 | 1445 | 1449 | 1465 | 1448 | 1402 | 1362 | 1386 | 1412 | 1342 | 1332 | 1343 | 1338 | 1359 | 1370 | 1335 | 1341 | 1311 | 1286 | 1279 | 1272 | 1300 | 1284 | 1246 | 1242 | 1220 | 1210 | 1236 | 1225 | 1204 | 1148 | 1126 | 1173 | 1239 | 1195 | 1174 | 1166 | 1117 | 1158 | 1139 | 1123 | 1132 | 1132 | 1147 | 1092 | 1091 | 1117 | 1103 | 1079 | 1060 | 1079 | 1051 | 1033 | 987 | 1009 | 1010 | 1018 | 1005 | 1010 | 1024 | 1012 | 995 | 1004 | 969 | 962 | 965 | 957 | 953 | 950 | 960 | 974 | 960 | 936 | 939 | 944 | 928 | 909 | 914 | 926 | 904 | 878 | 875 | 895 | 911 | 883 | 876 | 866 | 855 | 866 | 851 | 846 | 842 | 849 | 840 | 829 | 830 | 803 | 755 | 804 | 818 | 797 | 819 | 770 | 740 | 761 | 764 | 746 | 725 | 706 | 714 | 711 | 685 | 669 | 676 | 686 | 692 | 682 | 661 | 650 | 664 | 647 | 610 | 625 | 662 | 660 | 632 | 608 | 596 | 600 | 606 | 595 | 603 | 614 | 604 | 591 | 581 | 554 | 537 | 542 | 572 | 608 | 584 | 562 | 548 | 561 | 581 | 597 | 566 | 567 | 571 | 543 | 542 | 524 | 561 | 557 | 553 | 537 | 505 | 513 | 524 | 531 | 564 | 554 | 526 | 511 | 516 | 530 | 544 | 532 | 532 | 529 | 525 | 530 | 523 | 512 | 500 | 492 | 494 | 501 | 506 | 522 | 526 | 482 | 470 | 485 | 492 | 487 | 498 | 517 | 518 | 494 | 490 | 503 | 471 | 510 | 505 | 517 | 507 | 497 | 463 | 514 | 540 | 536 | 502 | 535 | 525 | 542 | 519 | 517 | 530 | 529 | 536 | 503 | 579 | 553 | 582 | 614 | 632 | 624 | 665 | 691 | 729 | 760 | 823 | 874 | 909 | 980 | 1012 | 1093 | 1253 | 1323 | 1392 | 1548 | 1689 | 1864 | 2086 | 2209 | 2393 | 2643 | 2866 | 3174 | 3542 | 3748 | 4057 | 4310 | 4667 | 5079 | 5565 | 5942 | 6354 | 6546 | 7229 | 7534 | 7901 | 8571 | 8807 | 9212 | 9706 | 10021 | 10283 | 10859 | 10987 | 11112 | 11543 | 11804 | 12004 | 11982 | 12297 | 12303 | 12288 | 12269 | 12257 | 12137 | 11850 | 11629 | 11456 | 11121 | 10820 | 10428 | 9960 | 9550 | 9214 | 8819 | 8292 | 7734 | 7262 | 6844 | 6366 | 5923 | 5439 | 5041 | 4665 | 4284 | 3913 | 3577 | 3234 | 2888 | 2639 | 2368 | 2152 | 1999 | 1800 | 1653 | 1499 | 1368 | 1269 | 1184 | 1074 | 1003 | 916 | 874 | 843 | 811 | 779 | 750 | 737 | 715 | 682 | 660 | 651 | 651 | 650 | 680 | 646 | 616 | 635 | 652 | 638 | 636 | 657 | 639 | 611 | 627 | 644 | 650 | 640 | 625 | 638 | 659 | 647 | 623 | 613 | 619 | 642 | 643 | 633 | 623 | 610 | 611 | 606 | 610 | 601 | 588 | 599 | 596 | 593 | 604 | 592 | 613 | 597 | 573 | 613 | 612 | 585 | 569 | 567 | 558 | 564 | 564 | 581 | 605 | 596 | 589 | 573 | 558 | 549 | 549 | 551 | 552 | 567 | 574 | 574 | 562 | 545 | 553 | 564 | 559 | 557 | 557 | 540 | 531 | 550 | 544 | 519 | 543 | 561 | 542 | 544 | 535 | 519 | 521 | 508 | 540 | 553 | 539 | 549 | 569 | 569 | 577 | 536 | 523 | 549 | 586 | 570 | 545 | 541 | 539 | 544 | 566 | 557 | 548 | 571 | 586 | 575 | 560 | 552 | 586 | 606 | 600 | 591 | 649 | 602 | 591 | 604 | 631 | 640 | 649 | 661 | 660 | 692 | 709 | 725 | 747 | 709 | 699 | 768 | 741 | 729 | 738 | 704 | 806 | 689 | 761 | 731 | 739 | 744 | 704 | 751 | 764 | 734 | 718 | 711 | 697 | 695 | 707 | 711 | 759 | 714 | 697 | 671 | 697 | 727 | 666 | 689 | 674 | 671 | 674 | 639 | 639 | 637 | 638 | 631 | 610 | 612 | 603 | 581 | 597 | 554 | 557 | 577 | 582 | 522 | 542 | 537 | 519 | 494 | 502 | 514 | 521 | 526 | 512 | 487 | 472 | 469 | 487 | 495 | 448 | 451 | 449 | 457 | 471 | 458 | 428 | 438 | 422 | 426 | 425 | 408 | 414 | 407 | 386 | 388 | 400 | 404 | 406 | 394 | 385 | 399 | 390 | 387 | 350 | 344 | 362 | 376 | 356 | 349 | 332 | 319 | 324 | 338 | 350 | 350 | 340 | 323 | 314 | 322 | 311 | 300 | 300 | 319 | 325 | 311 | 311 | 306 | 294 | 294 | 293 | 285 | 288 | 283 | 273 | 281 | 297 | 289 | 271 | 269 | 277 | 282 | 277 | 267 | 275 | 284 | 267 | 242 | 236 | 245 | 247 | 236 | 234 | 241 | 243 | 244 | 240 | 230 | 229 | 240 | 239 | 228 | 239 | 233 | 216 | 214 | 218 | 207 | 203 | 211 | 207 | 197 | 185 | 188 | 203 | 204 | 202 | 198 | 196 | 203 | 192 | 173 | 180 | 184 | 185 | 179 | 167 | 170 | 168 | 169 | 170 | 166 | 156 | 150 | 158 | 165 | 159 | 156 | 158 | 154 | 160 | 148 | 164 | 162 | 150 | 148 | 142 | 144 | 137 | 137 | 139 | 137 | 141 | 137 | 127 | 134 | 142 | 134 | 119 | 123 | 127 | 126 | 137 | 127 | 127 | 147 | 134 | 121 | 116 | 114 | 116 | 123 | 121 | 111 | 110 | 121 | 127 | 131 | 110 | 110 | 117 | 115 | 120 | 126 | 119 | 116 | 120 | 107 | 123 | 122 | 128 | 127 | 112 | 109 | 116 | 117 | 107 | 107 | 121 | 102 | 118 | 118 | 121 | 97 | 107 | 114 | 112 | 100 | 99 | 114 | 112 | 111 | 112 | 113 | 122 | 113 | 100 | 119 | 132 | 126 | 118 | 127 | 132 | 114 | 118 | 132 | 135 | 121 | 108 | 124 | 146 | 147 | 142 | 146 | 147 | 157 | 164 | 147 | 164 | 187 | 198 | 203 | 211 | 222 | 241 | 259 | 275 | 277 | 281 | 288 | 312 | 335 | 341 | 356 | 382 | 397 | 413 | 435 | 454 | 492 | 510 | 542 | 586 | 593 | 660 | 652 | 687 | 723 | 762 | 787 | 818 | 850 | 894 | 917 | 924 | 960 | 1004 | 1031 | 1044 | 1073 | 1113 | 1156 | 1196 | 1184 | 1165 | 1174 | 1180 | 1192 | 1206 | 1212 | 1235 | 1266 | 1293 | 1281 | 1272 | 1291 | 1272 | 1243 | 1269 | 1269 | 1206 | 1221 | 1200 | 1125 | 1112 | 1088 | 1061 | 1005 | 992 | 983 | 942 | 909 | 878 | 828 | 787 | 763 | 734 | 690 | 652 | 620 | 580 | 545 | 518 | 499 | 472 | 446 | 422 | 377 | 336 | 323 | 306 | 276 | 247 | 227 | 225 | 227 | 205 | 183 | 176 | 158 | 145 | 135 | 105 | 99 | 93 | 79 | 77 | 63 | 51 | 50 | 51 | 51 | 48 | 43 | 44 | 44 | 45 | 50 | 40 | 28 | 29 | 21 | 19 | 27 | 21 | 21 | 20 | 22 | 24 | 23 | 24 | 25 | 22 | 23 | 25 | 20 | 18 | 17 | 17 | 18 | 14 | 25 | 18 | 14 | 19 | 16 | 13 | 17 | 18 | 18 | 22 | 19 | 16 | 17 | 15 | 18 | 22 | 17 | 15 | 19 | 12 | 18 | 12 | 17 | 10 | 13 | 20 | 17 | 15 | 15 | 17 | 16 | 15 | 16 | 14 | 18 | 17 | 14 | 20 | 20 | 8 | 10 | 15 | 13 | 13 | 11 | 13 | 13 | 21 | 24 | 19 | 13 | 13 | 11 | 13 | 17 | 17 | 13 | 17 | 10 | 13 | 15 | 14 | 15 | 12 | 14 | 13 | 16 | 15 | 11 | 12 | 16 | 14 | 12 | 15 | 13 | 12 | 13 | 15 | 11 | 8 | 10 | 12 | 15 | 13 | 18 | 18 | 14 | 16 | 15 | 16 | 16 | 15 | 14 | 15 | 17 | 16 | 17 | 14 | 16 | 17 | 16 | 17 | 16 | 13 | 14 | 15 | 17 | 17 | 17 | 16 | 13 | 12 | 13 | 14 | 14 | 12 | 10 | 11 | 13 | 16 | 11 | 11 | 11 | 10 | 11 | 12 | 10 | 12 | 11 | 8 | 9 | 11 | 9 | 9 | 11 | 10 | 12 | 14 | 16 | 16 | 11 | 11 | 12 | 9 | 7 | 7 | 8 | 10 | 10 | 9 | 8 | 10 | 13 | 11 | 8 | 7 | 6 | 6 | 11 | 12 | 10 | 10 | 12 | 11 | 10 | 11 | 11 | 9 | 10 | 11 | 11 | 9 | 6 | 6 | 6 | 3 | 3 | 5 | 8 | 5 | 4 | 6 | 5 | 6 | 8 | 6 | 4 | 4 | 5 | 6 | 6 | 7 | 6 | 6 | 7 | 4 | 3 | 4 | 7 | 12 | 11 | 6 | 3 | 9 | 11 | 9 | 7 | 7 | 7 | 8 | 9 | 7 | 7 | 6 | 5 | 4 | 7 | 6 | 6 | 7 | 4 | 8 | 5 | 7 | 6 | 6 | 6 | 6 | 4 | 5 | 10 | 9 | 6 | 9 | 7 | 6 | 8 | 5 | 8 | 7 | 5 | 4 | 6 | 4 | 4 | 11 | 13 | 7 | 5 | 8 | 7 | 5 | 7 | 5 | 4 | 3 | 7 | 5 | 4 | 2 | 4 | 3 | 6 | 7 | 8 | 5 | 3 | 5 | 6 | 6 | 6 | 5 | 5 | 5 | 5 | 5 | 4 | 7 | 8 | 4 | 4 | 4 | 7 | 5 | 5 | 5 | 7 | 7 | 4 | 3 | 6 | 6 | 4 | 8 | 4 | 6 | 5 | 5 | 6 | 7 | 4 | 7 | 5 | 7 | 4 | 4 | 5 | 6 | 4 | 4 | 4 | 3 | 5 | 6 | 4 | 6 | 4 | 5 | 3 | 4 | 5 | 7 | 6 | 5 | 4 | 7 | 4 | 3 | 5 | 8 | 7 | 5 | 4 | 4 | 7 | 7 | 3 | 2 | 5 | 8 | 8 | 5 | 4 | 4 | 3 | 5 | 7 | 8 | 8 | 4 | 6 | 6 | 5 | 2 | 5 | 7 | 6 | 7 | 8 | 5 | 5 | 6 | 5 | 5 | 6 | 5 | 6 | 7 | 6 | 4 | 3 | 6 | 8 | 8 | 5 | 5 | 6 | 7 | 5 | 5 | 4 | 5 | 9 | 10 | 8 | 9 | 10 | 6 | 3 | 6 | 7 | 7 | 10 | 8 | 10 | 10 | 8 | 7 | 11 | 10 | 7 | 6 | 9 | 8 | 9 | 6 | 6 | 7 | 9 | 9 | 9 | 6 | 2 | 5 | 7 | 8 | 5 | 6 | 7 | 8 | 9 | 8 | 8 | 6 | 3 | 4 | 9 | 7 | 6 | 7 | 6 | 3 | 7 | 6 | 9 | 7 | 6 | 3 | 4 | 4 | 4 | 4 | 5 | 3 | 1 | 1 | 3 | 1 | 3 | 3 | 3 | 2 | 7 | 5 | 4 | 2 | 2 | 3 | 4 | 2 | 1 | 4 | 2 | 1 | 1 | 2 | 1 | 3 | 2 | 1 | 1 | 0 | 6 | 2 | 3 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 2 | 1 | 1 | 3 | 0 | 2 | 3 | 2 | 4 | 4 | 0 | 1 | 1 | 0 | 1 | 1 | 2 | 1 | 2 | 5 | 1 | 3 | 0 | 1 | 0 | 2 | 0 | 1 | 2 | 2 | 1 | 0 | 3 | 0 | 1 | 0 | 0 | 1 | 0 | 2 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 2 | 0 | 1 | 1 | 2 | 0 | 1 | 0 | 1 | 4 | 2 | 1 | 2 | 0 | 2 | 2 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 2 | 1 | 1 | 2 | 1 | 0 | 1 | 2 | 2 | 0 | 1 | 1 | 0 | 1 | 2 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 2 | 0 | 1 | 2 | 1 | 0 | 2 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 2 | 0 | 0 | 2 | 2 | 0 | 1 | 1 | 0 | 1 | 2 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 2 | 2 | 1 | 2 | 1 | 0 | 1 | 3 | 1 | 0 | 0 | 0 | 2 | 2 | 2 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 2 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 2 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 2 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 2 | 2 | 0 | 1 | 0 | 2 | 2 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 2 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 5 | 336 | 34 | 34 | 40 | 39 | 29 | 31 | 35 | 63 | 58 | 54 | 77 | 95 | 115 | 134 | 153 | 161 | 216 | 282 | 325 | 371 | 454 | 550 | 576 | 696 | 799 | 927 | 1066 | 1185 | 1242 | 1339 | 1406 | 1469 | 1533 | 1489 | 1481 | 1540 | 1544 | 1528 | 1505 | 1503 | 1600 | 1601 | 1615 | 1595 | 1563 | 1575 | 1596 | 1614 | 1618 | 1580 | 1529 | 1570 | 1632 | 1667 | 1595 | 1641 | 1631 | 1588 | 1585 | 1585 | 1624 | 1628 | 1619 | 1605 | 1701 | 1603 | 1589 | 1608 | 1580 | 1579 | 1616 | 1592 | 1599 | 1623 | 1671 | 1719 | 1671 | 1707 | 1680 | 1659 | 1692 | 1697 | 1686 | 1662 | 1658 | 1738 | 1731 | 1664 | 1603 | 1593 | 1573 | 1575 | 1565 | 1557 | 1543 | 1573 | 1623 | 1583 | 1553 | 1530 | 1526 | 1534 | 1545 | 1534 | 1517 | 1482 | 1516 | 1567 | 1521 | 1475 | 1459 | 1461 | 1480 | 1489 | 1485 | 1482 | 1505 | 1529 | 1479 | 1403 | 1381 | 1386 | 1405 | 1407 | 1403 | 1409 | 1417 | 1429 | 1414 | 1381 | 1379 | 1407 | 1421 | 1414 | 1428 | 1432 | 1431 | 1425 | 1420 | 1403 | 1397 | 1424 | 1432 | 1439 | 1442 | 1431 | 1427 | 1439 | 1464 | 1461 | 1470 | 1480 | 1487 | 1487 | 1495 | 1503 | 1495 | 1497 | 1558 | 1561 | 1555 | 1578 | 1626 | 1656 | 1710 | 1767 | 1774 | 1796 | 1855 | 1838 | 1935 | 1966 | 1971 | 1989 | 2043 | 2147 | 2130 | 2140 | 2095 | 2156 | 2168 | 2210 | 2207 | 2172 | 2155 | 2144 | 2141 | 2148 | 2123 | 2088 | 2074 | 2079 | 2056 | 2047 | 2066 | 2019 | 1964 | 1921 | 1893 | 1905 | 1872 | 1866 | 1877 | 1850 | 1782 | 1742 | 1775 | 1782 | 1767 | 1706 | 1623 | 1661 | 1624 | 1645 | 1600 | 1633 | 1572 | 1602 | 1626 | 1595 | 1602 | 1538 | 1538 | 1504 | 1527 | 1507 | 1427 | 1464 | 1479 | 1505 | 1461 | 1411 | 1439 | 1372 | 1406 | 1429 | 1376 | 1324 | 1389 | 1360 | 1326 | 1342 | 1356 | 1327 | 1309 | 1354 | 1345 | 1301 | 1305 | 1260 | 1207 | 1180 | 1154 | 1221 | 1281 | 1256 | 1206 | 1178 | 1180 | 1172 | 1153 | 1171 | 1137 | 1146 | 1131 | 1125 | 1122 | 1155 | 1117 | 1131 | 1079 | 1126 | 1109 | 1126 | 1064 | 1051 | 1053 | 1066 | 1034 | 1055 | 1064 | 1019 | 1032 | 1043 | 1003 | 971 | 1024 | 1024 | 1018 | 1002 | 1006 | 1002 | 972 | 945 | 925 | 950 | 971 | 954 | 931 | 937 | 927 | 914 | 939 | 954 | 941 | 935 | 940 | 923 | 882 | 863 | 886 | 913 | 905 | 882 | 875 | 867 | 846 | 833 | 841 | 853 | 848 | 857 | 844 | 801 | 797 | 801 | 769 | 777 | 783 | 753 | 744 | 723 | 742 | 740 | 726 | 707 | 710 | 722 | 701 | 671 | 675 | 680 | 662 | 661 | 670 | 656 | 631 | 608 | 603 | 619 | 621 | 623 | 634 | 639 | 647 | 638 | 620 | 610 | 598 | 620 | 631 | 576 | 549 | 577 | 604 | 620 | 576 | 552 | 579 | 589 | 573 | 592 | 595 | 568 | 574 | 573 | 561 | 557 | 561 | 538 | 535 | 573 | 552 | 548 | 545 | 564 | 546 | 545 | 561 | 531 | 517 | 540 | 531 | 512 | 521 | 531 | 531 | 520 | 515 | 521 | 517 | 504 | 496 | 493 | 503 | 522 | 519 | 513 | 510 | 504 | 516 | 509 | 521 | 533 | 516 | 496 | 492 | 512 | 485 | 471 | 460 | 492 | 496 | 495 | 513 | 501 | 507 | 480 | 515 | 500 | 509 | 517 | 551 | 517 | 517 | 538 | 558 | 557 | 539 | 571 | 567 | 546 | 580 | 619 | 615 | 647 | 667 | 675 | 700 | 755 | 804 | 890 | 923 | 951 | 1039 | 1110 | 1193 | 1307 | 1416 | 1569 | 1715 | 1833 | 1989 | 2193 | 2423 | 2631 | 2832 | 3059 | 3334 | 3640 | 4000 | 4283 | 4686 | 4986 | 5444 | 5760 | 6153 | 6479 | 6993 | 7498 | 7950 | 8351 | 8732 | 9104 | 9563 | 9797 | 10229 | 10704 | 10879 | 11157 | 11512 | 11616 | 11990 | 12176 | 12216 | 12169 | 12321 | 12352 | 12307 | 12268 | 11960 | 11756 | 11504 | 11179 | 10906 | 10438 | 10051 | 9728 | 9226 | 8737 | 8296 | 7797 | 7343 | 6924 | 6421 | 5985 | 5556 | 5114 | 4653 | 4217 | 3878 | 3544 | 3261 | 2899 | 2582 | 2405 | 2182 | 1997 | 1754 | 1642 | 1481 | 1346 | 1234 | 1161 | 1061 | 984 | 910 | 894 | 855 | 780 | 737 | 729 | 713 | 697 | 683 | 658 | 671 | 678 | 675 | 673 | 643 | 649 | 660 | 642 | 636 | 646 | 629 | 608 | 625 | 647 | 642 | 641 | 642 | 630 | 616 | 615 | 636 | 644 | 633 | 629 | 618 | 612 | 628 | 621 | 593 | 598 | 625 | 620 | 589 | 562 | 558 | 588 | 596 | 573 | 575 | 587 | 582 | 607 | 614 | 606 | 591 | 597 | 645 | 618 | 571 | 573 | 564 | 562 | 557 | 541 | 545 | 567 | 560 | 548 | 547 | 563 | 569 | 558 | 550 | 545 | 548 | 534 | 520 | 541 | 546 | 545 | 561 | 558 | 546 | 513 | 505 | 527 | 530 | 557 | 551 | 509 | 542 | 559 | 537 | 525 | 529 | 541 | 517 | 513 | 534 | 558 | 552 | 531 | 497 | 495 | 511 | 523 | 530 | 539 | 539 | 548 | 540 | 562 | 593 | 581 | 575 | 589 | 622 | 597 | 573 | 614 | 626 | 608 | 610 | 597 | 619 | 636 | 655 | 654 | 662 | 677 | 672 | 655 | 670 | 687 | 687 | 708 | 691 | 705 | 731 | 658 | 713 | 762 | 725 | 678 | 751 | 757 | 741 | 741 | 729 | 736 | 729 | 722 | 751 | 743 | 682 | 696 | 723 | 735 | 716 | 713 | 696 | 721 | 709 | 669 | 677 | 675 | 657 | 664 | 635 | 637 | 644 | 627 | 605 | 592 | 590 | 575 | 556 | 561 | 552 | 584 | 583 | 586 | 542 | 542 | 528 | 529 | 537 | 526 | 495 | 485 | 476 | 472 | 486 | 490 | 487 | 491 | 471 | 475 | 463 | 431 | 446 | 456 | 454 | 436 | 407 | 397 | 412 | 403 | 408 | 416 | 406 | 397 | 393 | 388 | 383 | 375 | 375 | 371 | 356 | 361 | 376 | 383 | 369 | 353 | 347 | 352 | 350 | 337 | 332 | 326 | 321 | 323 | 334 | 344 | 351 | 342 | 317 | 310 | 308 | 322 | 320 | 316 | 310 | 290 | 301 | 299 | 282 | 289 | 303 | 301 | 282 | 272 | 274 | 278 | 278 | 272 | 276 | 277 | 273 | 279 | 297 | 275 | 260 | 261 | 242 | 234 | 229 | 245 | 261 | 250 | 234 | 235 | 241 | 237 | 245 | 244 | 234 | 236 | 232 | 227 | 240 | 250 | 232 | 210 | 206 | 212 | 217 | 218 | 199 | 177 | 184 | 203 | 207 | 197 | 187 | 185 | 177 | 159 | 167 | 193 | 194 | 192 | 198 | 178 | 170 | 179 | 177 | 170 | 162 | 164 | 158 | 147 | 158 | 158 | 153 | 152 | 153 | 156 | 161 | 144 | 152 | 144 | 149 | 150 | 136 | 147 | 147 | 146 | 141 | 140 | 148 | 145 | 133 | 126 | 123 | 128 | 133 | 139 | 145 | 132 | 129 | 127 | 120 | 133 | 116 | 127 | 134 | 140 | 132 | 128 | 130 | 119 | 116 | 129 | 110 | 113 | 105 | 118 | 119 | 117 | 123 | 99 | 106 | 116 | 116 | 105 | 101 | 109 | 109 | 101 | 111 | 126 | 118 | 116 | 120 | 130 | 137 | 105 | 94 | 109 | 120 | 105 | 121 | 113 | 119 | 111 | 111 | 111 | 112 | 113 | 117 | 116 | 113 | 111 | 95 | 107 | 106 | 117 | 111 | 122 | 119 | 120 | 120 | 134 | 129 | 153 | 147 | 147 | 147 | 155 | 148 | 142 | 164 | 166 | 166 | 185 | 192 | 188 | 189 | 204 | 220 | 231 | 229 | 256 | 257 | 285 | 288 | 295 | 311 | 344 | 362 | 372 | 402 | 418 | 430 | 461 | 483 | 509 | 538 | 606 | 608 | 623 | 622 | 645 | 694 | 734 | 779 | 809 | 849 | 864 | 871 | 913 | 951 | 961 | 987 | 1032 | 1063 | 1058 | 1122 | 1196 | 1174 | 1198 | 1233 | 1194 | 1196 | 1214 | 1236 | 1242 | 1263 | 1281 | 1279 | 1271 | 1265 | 1240 | 1206 | 1184 | 1171 | 1182 | 1178 | 1149 | 1105 | 1130 | 1102 | 1028 | 1024 | 1000 | 971 | 932 | 888 | 858 | 811 | 777 | 759 | 724 | 707 | 662 | 603 | 589 | 562 | 527 | 510 | 462 | 432 | 435 | 406 | 346 | 307 | 298 | 282 | 261 | 245 | 223 | 206 | 189 | 175 | 169 | 152 | 138 | 116 | 102 | 93 | 87 | 81 | 81 | 76 | 70 | 60 | 48 | 44 | 42 | 35 | 32 | 32 | 32 | 30 | 26 | 29 | 30 | 31 | 33 | 30 | 31 | 24 | 22 | 20 | 21 | 18 | 18 | 21 | 21 | 21 | 23 | 22 | 16 | 14 | 22 | 24 | 21 | 18 | 16 | 19 | 18 | 16 | 15 | 18 | 22 | 21 | 17 | 14 | 13 | 16 | 15 | 22 | 22 | 14 | 15 | 15 | 13 | 12 | 21 | 14 | 15 | 19 | 20 | 19 | 18 | 15 | 19 | 18 | 12 | 12 | 11 | 15 | 21 | 22 | 17 | 16 | 18 | 17 | 10 | 19 | 13 | 14 | 18 | 16 | 18 | 16 | 16 | 16 | 14 | 13 | 14 | 13 | 20 | 21 | 21 | 21 | 21 | 15 | 16 | 19 | 12 | 8 | 16 | 12 | 13 | 12 | 13 | 16 | 15 | 15 | 16 | 12 | 14 | 14 | 16 | 12 | 11 | 16 | 14 | 17 | 14 | 15 | 14 | 12 | 9 | 12 | 16 | 16 | 13 | 13 | 15 | 17 | 17 | 10 | 11 | 12 | 15 | 13 | 12 | 12 | 11 | 16 | 18 | 14 | 10 | 11 | 12 | 12 | 12 | 14 | 14 | 11 | 12 | 10 | 7 | 10 | 9 | 10 | 15 | 17 | 14 | 11 | 12 | 14 | 14 | 12 | 9 | 11 | 12 | 13 | 12 | 11 | 11 | 10 | 11 | 13 | 14 | 10 | 8 | 10 | 9 | 12 | 11 | 9 | 11 | 13 | 11 | 9 | 9 | 8 | 7 | 9 | 9 | 9 | 10 | 8 | 6 | 7 | 6 | 9 | 12 | 10 | 10 | 10 | 10 | 8 | 6 | 6 | 7 | 10 | 11 | 9 | 9 | 8 | 7 | 7 | 10 | 10 | 12 | 10 | 6 | 7 | 8 | 9 | 9 | 8 | 8 | 8 | 9 | 10 | 12 | 9 | 7 | 8 | 9 | 6 | 4 | 5 | 12 | 10 | 8 | 7 | 7 | 8 | 7 | 9 | 12 | 10 | 8 | 6 | 5 | 5 | 6 | 5 | 5 | 6 | 6 | 8 | 8 | 5 | 6 | 6 | 9 | 8 | 5 | 5 | 6 | 4 | 3 | 3 | 4 | 6 | 5 | 7 | 7 | 8 | 7 | 6 | 7 | 7 | 6 | 7 | 10 | 7 | 5 | 7 | 8 | 6 | 2 | 5 | 5 | 4 | 4 | 9 | 6 | 4 | 9 | 2 | 7 | 6 | 6 | 10 | 7 | 7 | 6 | 4 | 5 | 7 | 5 | 5 | 9 | 6 | 5 | 5 | 5 | 3 | 3 | 4 | 3 | 5 | 5 | 6 | 5 | 4 | 5 | 4 | 3 | 4 | 3 | 4 | 4 | 7 | 8 | 3 | 3 | 3 | 4 | 2 | 8 | 4 | 5 | 3 | 4 | 5 | 6 | 6 | 5 | 5 | 7 | 8 | 7 | 3 | 5 | 5 | 6 | 4 | 3 | 5 | 7 | 5 | 6 | 6 | 4 | 4 | 6 | 6 | 6 | 4 | 4 | 4 | 4 | 4 | 4 | 3 | 4 | 4 | 8 | 7 | 10 | 10 | 3 | 5 | 5 | 5 | 5 | 6 | 5 | 5 | 3 | 2 | 4 | 5 | 5 | 4 | 7 | 8 | 7 | 5 | 5 | 5 | 7 | 6 | 5 | 5 | 6 | 6 | 5 | 5 | 4 | 6 | 9 | 12 | 9 | 7 | 8 | 6 | 7 | 8 | 9 | 9 | 6 | 4 | 6 | 4 | 4 | 5 | 8 | 8 | 4 | 5 | 8 | 9 | 9 | 5 | 4 | 5 | 8 | 8 | 5 | 5 | 4 | 5 | 5 | 5 | 9 | 10 | 7 | 4 | 6 | 7 | 7 | 6 | 7 | 7 | 7 | 8 | 7 | 4 | 4 | 3 | 5 | 5 | 7 | 7 | 6 | 5 | 5 | 6 | 4 | 4 | 7 | 5 | 7 | 8 | 4 | 4 | 6 | 4 | 7 | 4 | 4 | 3 | 4 | 2 | 4 | 2 | 3 | 6 | 3 | 4 | 1 | 2 | 4 | 4 | 2 | 4 | 3 | 5 | 3 | 2 | 2 | 0 | 1 | 3 | 4 | 2 | 1 | 2 | 2 | 2 | 0 | 1 | 2 | 2 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 3 | 1 | 0 | 3 | 2 | 1 | 0 | 1 | 3 | 0 | 2 | 3 | 2 | 0 | 1 | 1 | 0 | 1 | 3 | 4 | 1 | 1 | 4 | 3 | 1 | 1 | 1 | 3 | 0 | 1 | 0 | 0 | 1 | 1 | 2 | 4 | 0 | 0 | 0 | 3 | 1 | 1 | 1 | 1 | 3 | 1 | 0 | 0 | 0 | 3 | 0 | 2 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 2 | 2 | 2 | 1 | 1 | 0 | 1 | 1 | 1 | 2 | 0 | 0 | 1 | 0 | 1 | 2 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 2 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 2 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 2 | 1 | 0 | 1 | 2 | 2 | 1 | 2 | 1 | 3 | 4 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 1 | 1 | 1 | 0 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 2 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 6 | 90 | 5 | 6 | 5 | 5 | 5 | 5 | 8 | 7 | 11 | 12 | 14 | 17 | 28 | 30 | 37 | 49 | 60 | 78 | 108 | 152 | 200 | 268 | 332 | 379 | 416 | 442 | 454 | 479 | 502 | 533 | 542 | 527 | 577 | 599 | 587 | 577 | 563 | 562 | 575 | 578 | 586 | 591 | 597 | 609 | 582 | 569 | 589 | 576 | 572 | 579 | 570 | 604 | 606 | 591 | 602 | 566 | 565 | 569 | 569 | 587 | 587 | 591 | 586 | 568 | 588 | 595 | 579 | 581 | 587 | 583 | 586 | 589 | 578 | 572 | 598 | 633 | 624 | 605 | 607 | 593 | 580 | 597 | 601 | 612 | 615 | 597 | 591 | 571 | 560 | 579 | 569 | 548 | 570 | 599 | 579 | 536 | 560 | 595 | 561 | 540 | 532 | 514 | 529 | 553 | 555 | 548 | 545 | 547 | 549 | 543 | 568 | 586 | 544 | 529 | 540 | 524 | 509 | 499 | 511 | 526 | 503 | 486 | 489 | 489 | 513 | 505 | 495 | 507 | 500 | 476 | 492 | 527 | 516 | 506 | 508 | 491 | 495 | 506 | 514 | 516 | 498 | 495 | 493 | 483 | 500 | 525 | 522 | 514 | 521 | 532 | 532 | 510 | 506 | 528 | 556 | 561 | 547 | 557 | 565 | 568 | 553 | 569 | 578 | 603 | 635 | 654 | 636 | 626 | 657 | 697 | 768 | 756 | 754 | 776 | 785 | 810 | 807 | 788 | 774 | 782 | 777 | 782 | 803 | 789 | 774 | 765 | 748 | 758 | 759 | 741 | 727 | 725 | 720 | 705 | 702 | 686 | 669 | 669 | 672 | 661 | 647 | 645 | 636 | 613 | 599 | 621 | 602 | 587 | 594 | 597 | 565 | 549 | 567 | 559 | 559 | 538 | 518 | 550 | 542 | 523 | 503 | 496 | 540 | 533 | 499 | 483 | 501 | 529 | 502 | 492 | 495 | 473 | 466 | 446 | 461 | 453 | 463 | 462 | 460 | 470 | 471 | 466 | 477 | 462 | 453 | 466 | 466 | 439 | 431 | 420 | 409 | 428 | 439 | 434 | 411 | 397 | 393 | 390 | 389 | 386 | 404 | 394 | 394 | 379 | 370 | 391 | 387 | 366 | 359 | 368 | 352 | 354 | 354 | 370 | 353 | 350 | 370 | 340 | 354 | 341 | 337 | 345 | 363 | 339 | 328 | 333 | 329 | 319 | 327 | 303 | 306 | 307 | 308 | 307 | 308 | 317 | 319 | 310 | 318 | 322 | 314 | 318 | 323 | 327 | 318 | 307 | 306 | 286 | 282 | 287 | 297 | 288 | 295 | 291 | 268 | 281 | 275 | 276 | 274 | 275 | 252 | 283 | 276 | 254 | 245 | 228 | 254 | 261 | 257 | 263 | 255 | 267 | 244 | 238 | 236 | 228 | 238 | 222 | 224 | 213 | 204 | 213 | 222 | 225 | 220 | 231 | 228 | 222 | 222 | 223 | 216 | 209 | 208 | 206 | 205 | 210 | 212 | 207 | 200 | 211 | 205 | 191 | 177 | 189 | 187 | 193 | 194 | 193 | 193 | 188 | 181 | 189 | 177 | 201 | 192 | 171 | 192 | 193 | 201 | 203 | 177 | 196 | 197 | 196 | 211 | 201 | 204 | 183 | 184 | 183 | 177 | 178 | 168 | 173 | 184 | 172 | 167 | 162 | 184 | 192 | 181 | 184 | 189 | 186 | 186 | 179 | 161 | 163 | 172 | 167 | 163 | 155 | 150 | 175 | 195 | 178 | 164 | 170 | 148 | 172 | 169 | 173 | 137 | 149 | 158 | 167 | 168 | 180 | 160 | 167 | 175 | 183 | 179 | 166 | 157 | 183 | 168 | 163 | 159 | 156 | 182 | 184 | 188 | 195 | 195 | 182 | 224 | 229 | 230 | 268 | 270 | 282 | 305 | 343 | 369 | 399 | 467 | 522 | 552 | 612 | 691 | 749 | 813 | 893 | 992 | 1118 | 1230 | 1331 | 1462 | 1581 | 1743 | 1906 | 2062 | 2169 | 2418 | 2558 | 2689 | 2840 | 3038 | 3207 | 3372 | 3489 | 3702 | 3990 | 4089 | 4275 | 4374 | 4481 | 4504 | 4553 | 4654 | 4725 | 4837 | 4869 | 4777 | 4754 | 4764 | 4673 | 4588 | 4558 | 4434 | 4250 | 4125 | 3955 | 3743 | 3589 | 3396 | 3216 | 3085 | 2868 | 2617 | 2460 | 2309 | 2098 | 1900 | 1737 | 1605 | 1480 | 1374 | 1284 | 1140 | 1032 | 921 | 820 | 758 | 683 | 589 | 572 | 507 | 495 | 486 | 432 | 401 | 377 | 363 | 369 | 362 | 347 | 329 | 346 | 319 | 306 | 288 | 310 | 315 | 300 | 265 | 266 | 288 | 280 | 284 | 289 | 268 | 263 | 275 | 268 | 271 | 279 | 290 | 288 | 278 | 283 | 298 | 296 | 273 | 265 | 280 | 290 | 289 | 288 | 273 | 282 | 276 | 271 | 265 | 261 | 273 | 283 | 282 | 272 | 270 | 261 | 249 | 241 | 263 | 256 | 254 | 273 | 251 | 254 | 260 | 254 | 251 | 257 | 251 | 265 | 256 | 257 | 248 | 249 | 241 | 233 | 234 | 250 | 265 | 245 | 232 | 240 | 241 | 239 | 231 | 236 | 240 | 240 | 251 | 245 | 229 | 236 | 251 | 244 | 229 | 242 | 251 | 251 | 240 | 220 | 222 | 235 | 244 | 240 | 240 | 246 | 233 | 241 | 249 | 246 | 256 | 256 | 243 | 259 | 257 | 262 | 257 | 253 | 247 | 240 | 241 | 238 | 235 | 248 | 259 | 258 | 251 | 247 | 254 | 260 | 271 | 276 | 285 | 306 | 296 | 272 | 273 | 283 | 294 | 305 | 306 | 293 | 292 | 298 | 306 | 311 | 316 | 311 | 304 | 309 | 316 | 338 | 327 | 322 | 339 | 322 | 309 | 300 | 309 | 315 | 321 | 316 | 317 | 319 | 312 | 309 | 311 | 307 | 303 | 300 | 294 | 294 | 295 | 288 | 285 | 271 | 261 | 269 | 273 | 273 | 271 | 261 | 258 | 258 | 256 | 256 | 249 | 249 | 244 | 236 | 243 | 234 | 226 | 227 | 225 | 225 | 220 | 216 | 208 | 192 | 187 | 206 | 222 | 217 | 204 | 193 | 194 | 196 | 191 | 179 | 178 | 192 | 198 | 187 | 179 | 177 | 168 | 172 | 179 | 177 | 178 | 175 | 164 | 159 | 165 | 174 | 166 | 155 | 157 | 151 | 158 | 165 | 162 | 163 | 151 | 143 | 147 | 141 | 151 | 154 | 139 | 132 | 132 | 134 | 136 | 136 | 135 | 138 | 143 | 139 | 139 | 126 | 118 | 125 | 119 | 109 | 108 | 108 | 109 | 114 | 114 | 115 | 120 | 120 | 120 | 118 | 108 | 111 | 107 | 93 | 101 | 106 | 92 | 90 | 104 | 114 | 114 | 98 | 91 | 92 | 95 | 97 | 98 | 98 | 98 | 101 | 94 | 88 | 89 | 93 | 88 | 87 | 87 | 92 | 86 | 82 | 85 | 82 | 77 | 75 | 78 | 82 | 81 | 70 | 63 | 66 | 69 | 66 | 66 | 65 | 68 | 64 | 70 | 78 | 70 | 59 | 62 | 60 | 64 | 63 | 64 | 69 | 71 | 72 | 69 | 64 | 66 | 66 | 63 | 61 | 65 | 65 | 65 | 69 | 51 | 46 | 50 | 61 | 60 | 53 | 50 | 52 | 54 | 53 | 48 | 43 | 45 | 51 | 51 | 51 | 50 | 55 | 47 | 51 | 56 | 40 | 43 | 48 | 43 | 43 | 48 | 54 | 47 | 39 | 45 | 45 | 37 | 32 | 32 | 40 | 48 | 49 | 33 | 44 | 41 | 43 | 43 | 40 | 48 | 48 | 49 | 45 | 38 | 38 | 44 | 43 | 37 | 36 | 38 | 43 | 36 | 38 | 42 | 45 | 46 | 46 | 43 | 38 | 40 | 38 | 35 | 38 | 37 | 40 | 49 | 48 | 45 | 52 | 44 | 51 | 51 | 49 | 55 | 42 | 54 | 60 | 53 | 53 | 64 | 70 | 68 | 66 | 70 | 69 | 70 | 75 | 80 | 94 | 93 | 88 | 98 | 106 | 121 | 135 | 153 | 137 | 156 | 157 | 172 | 180 | 179 | 178 | 199 | 231 | 249 | 262 | 275 | 291 | 292 | 339 | 331 | 342 | 396 | 376 | 407 | 439 | 454 | 449 | 460 | 466 | 473 | 498 | 522 | 545 | 570 | 570 | 578 | 612 | 597 | 634 | 617 | 659 | 658 | 669 | 668 | 613 | 651 | 683 | 658 | 656 | 665 | 660 | 673 | 652 | 606 | 593 | 584 | 615 | 566 | 534 | 506 | 578 | 566 | 467 | 452 | 436 | 458 | 427 | 398 | 400 | 380 | 354 | 327 | 303 | 287 | 272 | 259 | 247 | 228 | 218 | 212 | 169 | 176 | 161 | 146 | 129 | 123 | 108 | 110 | 95 | 93 | 85 | 72 | 70 | 64 | 52 | 48 | 48 | 37 | 41 | 41 | 29 | 27 | 26 | 22 | 13 | 13 | 14 | 16 | 17 | 11 | 9 | 8 | 9 | 11 | 11 | 8 | 5 | 5 | 6 | 5 | 8 | 4 | 7 | 6 | 6 | 5 | 4 | 4 | 6 | 6 | 5 | 4 | 4 | 2 | 4 | 3 | 6 | 6 | 2 | 1 | 5 | 5 | 3 | 7 | 5 | 5 | 5 | 6 | 3 | 1 | 4 | 5 | 5 | 5 | 5 | 4 | 6 | 5 | 5 | 4 | 7 | 6 | 6 | 2 | 3 | 4 | 4 | 2 | 2 | 4 | 6 | 5 | 2 | 2 | 3 | 4 | 4 | 3 | 4 | 3 | 2 | 5 | 6 | 3 | 2 | 2 | 1 | 3 | 4 | 5 | 4 | 2 | 3 | 3 | 4 | 3 | 1 | 0 | 1 | 3 | 3 | 4 | 1 | 2 | 4 | 3 | 2 | 4 | 3 | 3 | 3 | 2 | 4 | 3 | 1 | 2 | 2 | 2 | 4 | 5 | 2 | 4 | 1 | 3 | 2 | 3 | 3 | 3 | 2 | 2 | 2 | 1 | 3 | 3 | 2 | 2 | 3 | 1 | 2 | 4 | 4 | 4 | 4 | 4 | 5 | 4 | 4 | 5 | 4 | 3 | 3 | 4 | 4 | 5 | 3 | 3 | 3 | 1 | 1 | 1 | 0 | 1 | 2 | 2 | 2 | 1 | 2 | 3 | 2 | 2 | 2 | 1 | 1 | 1 | 1 | 3 | 3 | 2 | 1 | 2 | 2 | 1 | 1 | 1 | 1 | 2 | 1 | 3 | 4 | 3 | 2 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 2 | 2 | 2 | 3 | 1 | 3 | 4 | 1 | 2 | 3 | 2 | 2 | 2 | 1 | 2 | 3 | 2 | 2 | 1 | 1 | 2 | 3 | 3 | 2 | 3 | 3 | 2 | 1 | 1 | 0 | 0 | 2 | 3 | 2 | 3 | 3 | 1 | 0 | 0 | 0 | 2 | 2 | 1 | 2 | 1 | 0 | 1 | 2 | 2 | 1 | 1 | 1 | 2 | 2 | 1 | 2 | 1 | 1 | 2 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 2 | 2 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 2 | 1 | 0 | 1 | 2 | 2 | 2 | 4 | 2 | 1 | 1 | 0 | 1 | 2 | 1 | 0 | 2 | 1 | 1 | 1 | 1 | 1 | 0 | 2 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 2 | 2 | 3 | 4 | 3 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 2 | 2 | 1 | 0 | 1 | 4 | 2 | 2 | 2 | 2 | 2 | 1 | 2 | 2 | 2 | 2 | 2 | 3 | 2 | 0 | 1 | 2 | 2 | 3 | 2 | 2 | 2 | 1 | 1 | 1 | 1 | 1 | 1 | 2 | 1 | 2 | 3 | 2 | 1 | 1 | 0 | 1 | 2 | 1 | 1 | 2 | 1 | 0 | 2 | 2 | 1 | 0 | 1 | 3 | 1 | 1 | 1 | 1 | 1 | 1 | 3 | 1 | 1 | 1 | 1 | 1 | 1 | 2 | 1 | 1 | 2 | 1 | 2 | 2 | 2 | 2 | 3 | 2 | 1 | 2 | 1 | 0 | 1 | 1 | 1 | 1 | 2 | 1 | 0 | 1 | 2 | 1 | 2 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 3 | 1 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 4 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 1 | 2 | 2 | 1 | 2 | 1 | 0 | 2 | 2 | 1 | 2 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 7 | 79 | 6 | 4 | 5 | 5 | 7 | 5 | 6 | 8 | 4 | 9 | 13 | 16 | 18 | 27 | 40 | 51 | 60 | 65 | 90 | 132 | 181 | 236 | 288 | 329 | 367 | 414 | 442 | 465 | 486 | 510 | 523 | 524 | 538 | 558 | 545 | 529 | 541 | 557 | 572 | 575 | 566 | 563 | 576 | 594 | 575 | 551 | 543 | 545 | 564 | 570 | 555 | 544 | 546 | 549 | 544 | 530 | 537 | 527 | 505 | 525 | 551 | 569 | 576 | 546 | 544 | 575 | 560 | 555 | 550 | 566 | 571 | 566 | 553 | 555 | 566 | 575 | 572 | 556 | 584 | 600 | 606 | 586 | 570 | 567 | 544 | 542 | 573 | 576 | 560 | 567 | 591 | 580 | 556 | 553 | 548 | 539 | 556 | 579 | 549 | 528 | 525 | 506 | 513 | 537 | 535 | 517 | 521 | 553 | 565 | 551 | 533 | 534 | 549 | 520 | 512 | 552 | 545 | 541 | 561 | 531 | 484 | 448 | 447 | 474 | 481 | 470 | 472 | 481 | 487 | 497 | 503 | 493 | 482 | 474 | 480 | 497 | 501 | 492 | 476 | 464 | 466 | 468 | 467 | 477 | 487 | 488 | 498 | 513 | 506 | 490 | 494 | 502 | 499 | 496 | 513 | 527 | 526 | 504 | 513 | 550 | 569 | 576 | 580 | 590 | 605 | 612 | 640 | 633 | 668 | 708 | 662 | 680 | 730 | 743 | 761 | 756 | 753 | 750 | 740 | 728 | 769 | 788 | 761 | 757 | 753 | 731 | 726 | 722 | 710 | 698 | 692 | 701 | 701 | 684 | 673 | 665 | 663 | 669 | 655 | 623 | 614 | 623 | 609 | 607 | 617 | 592 | 586 | 594 | 571 | 552 | 548 | 548 | 559 | 555 | 532 | 538 | 525 | 494 | 516 | 514 | 497 | 472 | 516 | 521 | 517 | 518 | 527 | 519 | 486 | 474 | 450 | 445 | 465 | 467 | 447 | 445 | 437 | 433 | 436 | 442 | 452 | 434 | 427 | 434 | 430 | 435 | 442 | 423 | 407 | 406 | 389 | 369 | 369 | 379 | 409 | 412 | 382 | 375 | 375 | 369 | 360 | 366 | 376 | 369 | 372 | 387 | 379 | 365 | 340 | 351 | 358 | 333 | 350 | 346 | 346 | 350 | 347 | 351 | 345 | 333 | 324 | 345 | 333 | 334 | 321 | 312 | 325 | 318 | 325 | 336 | 323 | 296 | 285 | 287 | 297 | 314 | 324 | 312 | 281 | 266 | 279 | 287 | 284 | 270 | 260 | 271 | 278 | 288 | 292 | 284 | 274 | 293 | 283 | 275 | 260 | 257 | 256 | 277 | 286 | 269 | 254 | 272 | 256 | 236 | 230 | 236 | 234 | 252 | 243 | 244 | 232 | 240 | 228 | 224 | 249 | 218 | 217 | 217 | 201 | 231 | 225 | 218 | 206 | 202 | 213 | 218 | 196 | 207 | 216 | 204 | 204 | 210 | 210 | 216 | 217 | 209 | 208 | 217 | 207 | 188 | 209 | 211 | 196 | 184 | 182 | 187 | 197 | 206 | 200 | 195 | 192 | 177 | 216 | 201 | 188 | 191 | 182 | 188 | 177 | 196 | 197 | 180 | 198 | 187 | 175 | 176 | 176 | 177 | 185 | 163 | 172 | 175 | 159 | 171 | 184 | 172 | 183 | 195 | 181 | 170 | 168 | 170 | 168 | 164 | 158 | 169 | 185 | 176 | 166 | 170 | 163 | 150 | 159 | 163 | 170 | 162 | 150 | 150 | 159 | 181 | 161 | 159 | 154 | 160 | 166 | 170 | 176 | 147 | 159 | 145 | 167 | 145 | 181 | 152 | 134 | 166 | 169 | 169 | 165 | 164 | 173 | 195 | 187 | 174 | 196 | 205 | 210 | 216 | 226 | 243 | 254 | 281 | 305 | 325 | 328 | 373 | 412 | 442 | 522 | 569 | 624 | 706 | 761 | 853 | 938 | 1012 | 1105 | 1208 | 1286 | 1405 | 1616 | 1794 | 1909 | 2080 | 2195 | 2273 | 2463 | 2627 | 2878 | 3032 | 3193 | 3386 | 3573 | 3868 | 3796 | 3925 | 4180 | 4255 | 4250 | 4475 | 4479 | 4568 | 4630 | 4683 | 4663 | 4567 | 4505 | 4515 | 4388 | 4286 | 4266 | 4136 | 4012 | 3846 | 3705 | 3545 | 3366 | 3144 | 2969 | 2815 | 2546 | 2319 | 2218 | 2103 | 1931 | 1773 | 1655 | 1506 | 1335 | 1188 | 1099 | 996 | 893 | 832 | 759 | 679 | 612 | 585 | 554 | 485 | 439 | 410 | 425 | 359 | 356 | 312 | 340 | 321 | 318 | 285 | 297 | 285 | 279 | 265 | 268 | 269 | 288 | 284 | 269 | 270 | 280 | 279 | 257 | 253 | 273 | 279 | 275 | 274 | 273 | 264 | 265 | 284 | 276 | 266 | 264 | 260 | 281 | 298 | 288 | 277 | 280 | 272 | 264 | 257 | 267 | 260 | 264 | 270 | 246 | 254 | 256 | 242 | 254 | 263 | 236 | 262 | 252 | 260 | 251 | 231 | 238 | 254 | 254 | 253 | 242 | 227 | 229 | 225 | 254 | 263 | 245 | 231 | 228 | 237 | 235 | 243 | 252 | 235 | 225 | 233 | 237 | 231 | 231 | 244 | 239 | 216 | 213 | 235 | 239 | 225 | 227 | 229 | 230 | 228 | 219 | 219 | 224 | 230 | 232 | 222 | 221 | 219 | 239 | 235 | 222 | 225 | 221 | 246 | 257 | 255 | 252 | 229 | 220 | 236 | 234 | 243 | 238 | 234 | 243 | 252 | 245 | 238 | 237 | 231 | 228 | 236 | 251 | 257 | 267 | 272 | 268 | 267 | 265 | 273 | 274 | 284 | 276 | 257 | 269 | 290 | 296 | 287 | 284 | 296 | 306 | 307 | 313 | 321 | 301 | 306 | 331 | 298 | 286 | 294 | 323 | 335 | 315 | 296 | 290 | 283 | 293 | 303 | 295 | 284 | 274 | 276 | 283 | 292 | 314 | 301 | 289 | 262 | 264 | 267 | 258 | 263 | 260 | 257 | 263 | 266 | 254 | 242 | 237 | 239 | 233 | 217 | 226 | 234 | 221 | 220 | 215 | 215 | 224 | 216 | 210 | 195 | 194 | 196 | 198 | 192 | 186 | 185 | 185 | 184 | 184 | 184 | 174 | 162 | 167 | 187 | 193 | 187 | 180 | 172 | 169 | 163 | 157 | 159 | 163 | 154 | 143 | 144 | 149 | 155 | 157 | 153 | 150 | 144 | 140 | 137 | 149 | 153 | 148 | 148 | 140 | 137 | 135 | 132 | 130 | 132 | 128 | 119 | 123 | 133 | 123 | 121 | 129 | 119 | 110 | 110 | 115 | 104 | 95 | 106 | 107 | 103 | 108 | 115 | 114 | 108 | 110 | 119 | 122 | 121 | 112 | 96 | 98 | 107 | 105 | 96 | 85 | 91 | 98 | 91 | 81 | 90 | 99 | 90 | 87 | 96 | 97 | 90 | 90 | 89 | 89 | 84 | 78 | 78 | 76 | 89 | 94 | 90 | 96 | 85 | 68 | 68 | 74 | 72 | 73 | 77 | 75 | 77 | 84 | 77 | 75 | 76 | 75 | 63 | 64 | 70 | 64 | 60 | 68 | 61 | 58 | 56 | 60 | 67 | 63 | 56 | 59 | 60 | 51 | 52 | 62 | 59 | 52 | 56 | 54 | 56 | 50 | 46 | 58 | 59 | 53 | 56 | 54 | 54 | 56 | 50 | 42 | 42 | 45 | 45 | 44 | 40 | 47 | 52 | 50 | 47 | 48 | 40 | 44 | 46 | 48 | 50 | 48 | 47 | 46 | 46 | 45 | 44 | 42 | 38 | 43 | 43 | 44 | 44 | 47 | 47 | 50 | 38 | 29 | 30 | 40 | 45 | 44 | 46 | 44 | 45 | 47 | 40 | 36 | 33 | 41 | 39 | 39 | 36 | 34 | 44 | 36 | 45 | 45 | 40 | 37 | 43 | 47 | 44 | 39 | 33 | 29 | 32 | 44 | 47 | 46 | 45 | 49 | 52 | 50 | 45 | 52 | 55 | 52 | 60 | 61 | 56 | 56 | 56 | 59 | 65 | 73 | 80 | 84 | 93 | 99 | 89 | 111 | 112 | 109 | 124 | 129 | 121 | 138 | 154 | 166 | 179 | 185 | 195 | 199 | 201 | 228 | 258 | 268 | 274 | 288 | 292 | 325 | 319 | 372 | 360 | 386 | 394 | 412 | 447 | 467 | 465 | 475 | 502 | 506 | 494 | 506 | 536 | 560 | 576 | 606 | 577 | 571 | 618 | 641 | 678 | 655 | 654 | 651 | 644 | 664 | 664 | 643 | 654 | 658 | 635 | 609 | 585 | 581 | 575 | 566 | 553 | 532 | 519 | 522 | 489 | 467 | 442 | 441 | 430 | 427 | 393 | 358 | 354 | 338 | 305 | 291 | 288 | 277 | 246 | 216 | 207 | 210 | 203 | 183 | 158 | 134 | 139 | 130 | 117 | 115 | 107 | 89 | 73 | 66 | 61 | 60 | 58 | 51 | 54 | 43 | 29 | 32 | 28 | 21 | 24 | 17 | 20 | 25 | 17 | 16 | 17 | 14 | 12 | 9 | 8 | 8 | 8 | 7 | 9 | 9 | 8 | 6 | 6 | 7 | 8 | 6 | 3 | 5 | 3 | 2 | 3 | 4 | 3 | 3 | 5 | 5 | 4 | 1 | 2 | 4 | 9 | 5 | 2 | 3 | 2 | 4 | 6 | 6 | 3 | 4 | 4 | 3 | 5 | 4 | 3 | 4 | 5 | 5 | 3 | 3 | 6 | 7 | 1 | 4 | 4 | 5 | 4 | 5 | 2 | 2 | 1 | 2 | 2 | 4 | 4 | 3 | 3 | 1 | 1 | 2 | 3 | 5 | 2 | 3 | 3 | 6 | 7 | 3 | 3 | 3 | 2 | 2 | 1 | 2 | 2 | 2 | 2 | 2 | 4 | 4 | 2 | 2 | 2 | 3 | 2 | 1 | 2 | 2 | 1 | 5 | 3 | 2 | 5 | 4 | 4 | 4 | 5 | 6 | 3 | 4 | 2 | 3 | 3 | 3 | 2 | 3 | 3 | 3 | 3 | 3 | 2 | 2 | 1 | 1 | 2 | 2 | 2 | 2 | 2 | 1 | 3 | 5 | 3 | 5 | 4 | 2 | 3 | 4 | 2 | 2 | 2 | 1 | 3 | 5 | 3 | 1 | 1 | 1 | 4 | 6 | 1 | 2 | 4 | 3 | 5 | 4 | 3 | 3 | 1 | 1 | 2 | 2 | 3 | 1 | 2 | 3 | 3 | 4 | 0 | 1 | 2 | 4 | 3 | 2 | 0 | 1 | 2 | 0 | 1 | 1 | 1 | 3 | 1 | 3 | 2 | 1 | 2 | 1 | 1 | 0 | 1 | 1 | 2 | 3 | 5 | 1 | 1 | 1 | 1 | 1 | 2 | 2 | 3 | 2 | 1 | 2 | 3 | 2 | 2 | 2 | 1 | 1 | 1 | 3 | 1 | 0 | 1 | 1 | 1 | 2 | 3 | 2 | 1 | 2 | 2 | 1 | 0 | 1 | 1 | 1 | 2 | 2 | 1 | 2 | 1 | 1 | 0 | 1 | 1 | 2 | 2 | 1 | 1 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 1 | 2 | 1 | 1 | 3 | 1 | 0 | 1 | 2 | 2 | 1 | 2 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 2 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 2 | 1 | 0 | 1 | 2 | 1 | 1 | 1 | 2 | 1 | 2 | 2 | 1 | 1 | 0 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 2 | 1 | 1 | 1 | 2 | 1 | 1 | 3 | 2 | 2 | 2 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 3 | 2 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 2 | 2 | 1 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 1 | 2 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 2 | 2 | 1 | 0 | 0 | 1 | 1 | 3 | 3 | 4 | 3 | 2 | 1 | 1 | 1 | 0 | 1 | 2 | 1 | 1 | 0 | 1 | 2 | 1 | 1 | 2 | 1 | 1 | 0 | 2 | 2 | 1 | 2 | 3 | 1 | 1 | 2 | 1 | 3 | 3 | 1 | 3 | 2 | 1 | 0 | 0 | 1 | 2 | 3 | 3 | 2 | 3 | 1 | 1 | 1 | 2 | 2 | 1 | 2 | 2 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 2 | 3 | 2 | 1 | 1 | 0 | 0 | 1 | 2 | 3 | 2 | 1 | 2 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 8 | 86 | 4 | 2 | 4 | 4 | 4 | 9 | 5 | 4 | 6 | 7 | 9 | 15 | 23 | 23 | 30 | 43 | 54 | 68 | 92 | 128 | 165 | 215 | 262 | 297 | 348 | 415 | 432 | 451 | 473 | 486 | 519 | 530 | 543 | 558 | 525 | 515 | 516 | 526 | 519 | 524 | 529 | 524 | 534 | 564 | 550 | 537 | 564 | 554 | 534 | 532 | 535 | 537 | 539 | 557 | 551 | 534 | 543 | 537 | 539 | 534 | 524 | 524 | 548 | 572 | 567 | 544 | 532 | 519 | 532 | 556 | 552 | 545 | 540 | 551 | 558 | 578 | 594 | 552 | 575 | 594 | 575 | 558 | 545 | 551 | 571 | 569 | 564 | 561 | 553 | 546 | 542 | 552 | 546 | 524 | 520 | 532 | 529 | 533 | 550 | 560 | 544 | 536 | 537 | 510 | 497 | 511 | 527 | 526 | 521 | 522 | 534 | 530 | 521 | 529 | 536 | 533 | 516 | 505 | 514 | 498 | 449 | 433 | 467 | 480 | 489 | 495 | 477 | 487 | 493 | 482 | 479 | 494 | 493 | 470 | 468 | 480 | 488 | 498 | 504 | 492 | 479 | 477 | 477 | 475 | 476 | 470 | 467 | 484 | 498 | 484 | 469 | 479 | 496 | 504 | 517 | 534 | 539 | 516 | 513 | 518 | 533 | 549 | 565 | 600 | 613 | 602 | 618 | 617 | 633 | 650 | 683 | 688 | 707 | 702 | 713 | 726 | 720 | 744 | 761 | 764 | 747 | 707 | 708 | 747 | 746 | 705 | 693 | 722 | 726 | 695 | 689 | 695 | 690 | 683 | 655 | 642 | 653 | 651 | 634 | 611 | 603 | 607 | 611 | 596 | 585 | 579 | 579 | 586 | 574 | 556 | 558 | 560 | 544 | 505 | 491 | 515 | 543 | 509 | 508 | 485 | 504 | 522 | 506 | 488 | 469 | 487 | 466 | 453 | 478 | 479 | 466 | 475 | 470 | 445 | 445 | 440 | 448 | 452 | 449 | 437 | 427 | 407 | 411 | 435 | 431 | 411 | 404 | 403 | 386 | 383 | 391 | 385 | 394 | 400 | 388 | 379 | 366 | 365 | 375 | 389 | 390 | 389 | 374 | 360 | 365 | 362 | 353 | 345 | 362 | 351 | 347 | 359 | 345 | 345 | 350 | 338 | 326 | 340 | 333 | 330 | 327 | 309 | 309 | 320 | 325 | 297 | 292 | 310 | 322 | 306 | 289 | 304 | 308 | 297 | 291 | 291 | 296 | 304 | 302 | 294 | 298 | 295 | 282 | 277 | 276 | 284 | 285 | 277 | 282 | 291 | 292 | 279 | 267 | 264 | 267 | 283 | 273 | 239 | 240 | 248 | 260 | 256 | 236 | 259 | 244 | 277 | 264 | 248 | 250 | 241 | 249 | 233 | 229 | 235 | 210 | 239 | 224 | 225 | 209 | 222 | 218 | 204 | 212 | 213 | 215 | 222 | 209 | 201 | 197 | 198 | 210 | 208 | 202 | 201 | 207 | 200 | 197 | 208 | 189 | 173 | 170 | 181 | 186 | 179 | 181 | 191 | 186 | 192 | 182 | 182 | 196 | 170 | 177 | 184 | 194 | 184 | 187 | 191 | 185 | 226 | 190 | 195 | 186 | 187 | 186 | 195 | 191 | 180 | 167 | 169 | 181 | 159 | 175 | 157 | 169 | 161 | 170 | 175 | 169 | 166 | 164 | 159 | 164 | 174 | 169 | 155 | 148 | 161 | 171 | 166 | 163 | 155 | 148 | 157 | 154 | 155 | 178 | 170 | 173 | 158 | 147 | 159 | 156 | 174 | 158 | 142 | 130 | 174 | 145 | 183 | 139 | 140 | 176 | 171 | 159 | 161 | 154 | 171 | 162 | 159 | 183 | 164 | 167 | 169 | 176 | 193 | 197 | 224 | 234 | 251 | 254 | 262 | 277 | 317 | 338 | 358 | 389 | 450 | 517 | 549 | 609 | 704 | 767 | 845 | 965 | 1065 | 1111 | 1182 | 1304 | 1392 | 1556 | 1658 | 1832 | 2039 | 2201 | 2227 | 2494 | 2628 | 2807 | 2909 | 3151 | 3237 | 3457 | 3731 | 3740 | 3977 | 4099 | 4218 | 4281 | 4252 | 4440 | 4589 | 4506 | 4463 | 4598 | 4494 | 4502 | 4522 | 4490 | 4290 | 4218 | 4060 | 3871 | 3719 | 3593 | 3460 | 3313 | 3153 | 2939 | 2688 | 2477 | 2300 | 2152 | 2040 | 1888 | 1713 | 1582 | 1463 | 1309 | 1187 | 1081 | 953 | 877 | 806 | 790 | 681 | 625 | 575 | 541 | 482 | 465 | 421 | 377 | 353 | 347 | 341 | 305 | 336 | 330 | 304 | 293 | 296 | 282 | 292 | 272 | 297 | 276 | 277 | 289 | 278 | 276 | 288 | 275 | 257 | 276 | 283 | 276 | 272 | 263 | 265 | 280 | 287 | 273 | 264 | 261 | 259 | 267 | 262 | 259 | 276 | 268 | 251 | 257 | 253 | 234 | 235 | 255 | 261 | 259 | 241 | 222 | 243 | 264 | 247 | 240 | 270 | 259 | 257 | 264 | 261 | 256 | 235 | 239 | 241 | 213 | 228 | 247 | 236 | 250 | 255 | 239 | 248 | 252 | 247 | 243 | 252 | 266 | 248 | 231 | 225 | 232 | 229 | 223 | 237 | 237 | 225 | 224 | 231 | 239 | 237 | 236 | 226 | 221 | 228 | 227 | 225 | 224 | 215 | 196 | 201 | 228 | 253 | 241 | 229 | 236 | 234 | 231 | 230 | 223 | 240 | 234 | 238 | 230 | 243 | 250 | 251 | 255 | 261 | 247 | 234 | 233 | 244 | 243 | 241 | 255 | 258 | 254 | 248 | 268 | 292 | 280 | 267 | 262 | 280 | 281 | 281 | 289 | 284 | 286 | 287 | 289 | 301 | 313 | 308 | 297 | 309 | 324 | 305 | 297 | 311 | 309 | 294 | 282 | 279 | 290 | 296 | 305 | 306 | 299 | 289 | 289 | 291 | 292 | 291 | 289 | 275 | 255 | 269 | 274 | 289 | 291 | 262 | 260 | 267 | 266 | 261 | 252 | 258 | 263 | 254 | 238 | 233 | 236 | 226 | 212 | 224 | 233 | 216 | 204 | 211 | 219 | 211 | 203 | 203 | 197 | 184 | 184 | 188 | 184 | 181 | 184 | 183 | 181 | 188 | 189 | 180 | 173 | 171 | 165 | 172 | 186 | 188 | 175 | 165 | 174 | 167 | 160 | 166 | 166 | 161 | 151 | 143 | 145 | 146 | 145 | 143 | 148 | 151 | 148 | 152 | 138 | 126 | 136 | 142 | 139 | 133 | 129 | 133 | 134 | 124 | 119 | 128 | 129 | 123 | 123 | 122 | 118 | 120 | 119 | 117 | 118 | 112 | 111 | 112 | 112 | 111 | 111 | 112 | 109 | 107 | 110 | 109 | 97 | 96 | 98 | 94 | 95 | 99 | 104 | 113 | 100 | 94 | 93 | 84 | 88 | 90 | 85 | 86 | 98 | 104 | 97 | 96 | 96 | 94 | 87 | 75 | 76 | 84 | 83 | 86 | 88 | 86 | 81 | 73 | 70 | 81 | 82 | 75 | 77 | 79 | 78 | 74 | 72 | 80 | 83 | 74 | 63 | 58 | 52 | 59 | 62 | 61 | 54 | 56 | 63 | 63 | 56 | 52 | 56 | 62 | 66 | 68 | 67 | 60 | 61 | 58 | 53 | 62 | 69 | 59 | 57 | 53 | 47 | 48 | 54 | 60 | 56 | 51 | 52 | 55 | 54 | 48 | 43 | 43 | 48 | 52 | 55 | 50 | 51 | 47 | 39 | 42 | 46 | 52 | 49 | 51 | 45 | 40 | 47 | 50 | 44 | 40 | 38 | 44 | 50 | 51 | 43 | 41 | 45 | 48 | 43 | 46 | 43 | 42 | 40 | 40 | 39 | 43 | 47 | 43 | 39 | 38 | 39 | 41 | 42 | 41 | 35 | 32 | 41 | 41 | 36 | 31 | 39 | 43 | 39 | 36 | 40 | 42 | 41 | 41 | 37 | 35 | 37 | 40 | 42 | 41 | 38 | 42 | 39 | 48 | 51 | 50 | 56 | 54 | 55 | 56 | 61 | 62 | 70 | 74 | 65 | 64 | 76 | 84 | 94 | 98 | 105 | 106 | 115 | 140 | 127 | 122 | 146 | 167 | 177 | 179 | 171 | 182 | 208 | 226 | 237 | 242 | 251 | 265 | 287 | 321 | 299 | 321 | 346 | 372 | 381 | 392 | 421 | 428 | 447 | 487 | 492 | 511 | 531 | 524 | 536 | 569 | 573 | 571 | 586 | 566 | 594 | 607 | 633 | 592 | 624 | 646 | 637 | 601 | 623 | 629 | 626 | 588 | 585 | 605 | 602 | 610 | 603 | 604 | 576 | 566 | 519 | 503 | 497 | 510 | 464 | 419 | 425 | 429 | 378 | 370 | 359 | 343 | 342 | 311 | 275 | 281 | 278 | 267 | 243 | 228 | 193 | 171 | 183 | 173 | 156 | 115 | 133 | 107 | 98 | 88 | 89 | 84 | 70 | 61 | 60 | 60 | 50 | 40 | 44 | 38 | 32 | 30 | 24 | 24 | 17 | 18 | 17 | 18 | 15 | 11 | 13 | 15 | 11 | 6 | 7 | 7 | 4 | 4 | 8 | 4 | 5 | 7 | 5 | 6 | 4 | 4 | 8 | 5 | 4 | 6 | 5 | 5 | 5 | 5 | 3 | 3 | 4 | 2 | 4 | 4 | 5 | 4 | 5 | 1 | 6 | 6 | 2 | 2 | 3 | 5 | 3 | 3 | 3 | 2 | 4 | 5 | 3 | 3 | 5 | 3 | 5 | 5 | 4 | 3 | 4 | 2 | 0 | 3 | 3 | 2 | 3 | 5 | 2 | 2 | 2 | 4 | 6 | 2 | 5 | 5 | 3 | 6 | 2 | 2 | 1 | 3 | 3 | 1 | 3 | 3 | 4 | 3 | 5 | 5 | 2 | 2 | 3 | 1 | 2 | 3 | 4 | 4 | 2 | 3 | 4 | 2 | 3 | 5 | 5 | 4 | 4 | 5 | 4 | 5 | 4 | 2 | 4 | 0 | 3 | 4 | 4 | 5 | 2 | 1 | 4 | 1 | 5 | 4 | 2 | 4 | 4 | 4 | 2 | 2 | 2 | 4 | 3 | 1 | 2 | 2 | 4 | 3 | 2 | 4 | 2 | 1 | 1 | 4 | 2 | 3 | 4 | 2 | 2 | 2 | 3 | 3 | 1 | 0 | 1 | 3 | 1 | 2 | 0 | 3 | 5 | 4 | 2 | 2 | 5 | 6 | 4 | 6 | 4 | 2 | 1 | 0 | 1 | 2 | 2 | 3 | 1 | 2 | 1 | 3 | 2 | 2 | 2 | 2 | 2 | 1 | 3 | 1 | 1 | 2 | 2 | 3 | 1 | 1 | 2 | 2 | 3 | 1 | 1 | 2 | 1 | 2 | 2 | 1 | 1 | 3 | 2 | 2 | 2 | 1 | 2 | 2 | 1 | 2 | 2 | 2 | 1 | 3 | 1 | 2 | 2 | 1 | 1 | 1 | 1 | 1 | 1 | 3 | 3 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 2 | 2 | 2 | 2 | 1 | 2 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 3 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 2 | 2 | 2 | 1 | 1 | 2 | 2 | 2 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 2 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 2 | 1 | 2 | 2 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 2 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 2 | 2 | 1 | 1 | 0 | 1 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 2 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 2 | 3 | 0 | 1 | 2 | 1 | 0 | 1 | 2 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 2 | 1 | 0 | 2 | 2 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 2 | 1 | 1 | 1 | 0 | 1 | 2 | 1 | 1 | 3 | 3 | 1 | 2 | 3 | 2 | 3 | 4 | 2 | 2 | 2 | 2 | 2 | 1 | 2 | 3 | 1 | 0 | 1 | 1 | 1 | 2 | 1 | 2 | 1 | 0 | 0 | 2 | 2 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 0 | 1 | 1 | 2 | 1 | 0 | 1 | 0 | 0 | 2 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 3 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 9 | 76 | 4 | 3 | 4 | 4 | 3 | 3 | 4 | 7 | 11 | 9 | 13 | 11 | 12 | 19 | 26 | 34 | 50 | 67 | 85 | 116 | 149 | 191 | 250 | 299 | 358 | 406 | 420 | 449 | 482 | 495 | 488 | 496 | 513 | 517 | 530 | 545 | 569 | 586 | 556 | 540 | 531 | 523 | 546 | 572 | 580 | 565 | 572 | 550 | 561 | 549 | 526 | 531 | 541 | 562 | 546 | 521 | 541 | 539 | 532 | 544 | 541 | 548 | 548 | 545 | 555 | 554 | 534 | 544 | 539 | 534 | 544 | 544 | 550 | 531 | 536 | 588 | 604 | 573 | 566 | 561 | 557 | 555 | 547 | 535 | 538 | 549 | 555 | 551 | 541 | 547 | 556 | 569 | 560 | 532 | 521 | 542 | 556 | 521 | 504 | 513 | 510 | 517 | 533 | 534 | 517 | 512 | 527 | 536 | 516 | 515 | 515 | 493 | 496 | 514 | 521 | 513 | 506 | 522 | 521 | 488 | 465 | 470 | 468 | 454 | 462 | 454 | 456 | 483 | 472 | 451 | 460 | 483 | 476 | 459 | 468 | 482 | 492 | 491 | 478 | 474 | 476 | 464 | 449 | 447 | 444 | 449 | 474 | 484 | 488 | 491 | 500 | 518 | 517 | 491 | 488 | 499 | 490 | 498 | 521 | 535 | 550 | 542 | 552 | 563 | 555 | 589 | 605 | 611 | 667 | 670 | 662 | 682 | 690 | 691 | 699 | 694 | 700 | 743 | 748 | 721 | 700 | 707 | 722 | 726 | 725 | 688 | 674 | 695 | 724 | 721 | 694 | 682 | 674 | 662 | 649 | 630 | 614 | 600 | 596 | 605 | 600 | 583 | 587 | 593 | 587 | 583 | 569 | 555 | 555 | 569 | 577 | 574 | 556 | 513 | 497 | 510 | 518 | 504 | 503 | 495 | 489 | 484 | 512 | 501 | 470 | 472 | 469 | 493 | 463 | 429 | 422 | 439 | 446 | 457 | 450 | 447 | 447 | 429 | 412 | 411 | 413 | 414 | 405 | 424 | 446 | 423 | 411 | 402 | 398 | 408 | 398 | 375 | 380 | 404 | 394 | 378 | 370 | 384 | 396 | 392 | 368 | 365 | 365 | 379 | 367 | 355 | 353 | 353 | 350 | 360 | 350 | 346 | 345 | 338 | 352 | 361 | 349 | 336 | 327 | 340 | 336 | 332 | 326 | 320 | 342 | 337 | 330 | 316 | 299 | 294 | 302 | 308 | 312 | 299 | 293 | 287 | 280 | 275 | 280 | 292 | 282 | 274 | 286 | 289 | 285 | 288 | 295 | 282 | 279 | 275 | 285 | 272 | 267 | 282 | 278 | 263 | 244 | 246 | 275 | 270 | 259 | 228 | 249 | 242 | 255 | 252 | 238 | 253 | 234 | 234 | 211 | 238 | 229 | 223 | 214 | 235 | 228 | 230 | 210 | 209 | 198 | 195 | 200 | 222 | 210 | 213 | 207 | 195 | 207 | 208 | 206 | 204 | 188 | 183 | 198 | 191 | 178 | 184 | 188 | 205 | 196 | 191 | 188 | 170 | 165 | 187 | 194 | 184 | 198 | 201 | 201 | 195 | 185 | 180 | 185 | 177 | 175 | 176 | 163 | 159 | 162 | 165 | 196 | 174 | 185 | 150 | 163 | 163 | 161 | 152 | 176 | 165 | 154 | 157 | 162 | 164 | 155 | 169 | 173 | 160 | 170 | 174 | 171 | 152 | 157 | 171 | 159 | 151 | 160 | 164 | 152 | 149 | 151 | 160 | 155 | 144 | 137 | 149 | 143 | 159 | 159 | 151 | 169 | 168 | 146 | 160 | 170 | 168 | 150 | 154 | 143 | 141 | 170 | 170 | 151 | 156 | 162 | 163 | 146 | 177 | 173 | 167 | 164 | 190 | 186 | 172 | 207 | 218 | 223 | 224 | 250 | 249 | 268 | 300 | 311 | 350 | 394 | 429 | 486 | 554 | 609 | 667 | 717 | 771 | 872 | 965 | 1087 | 1175 | 1275 | 1405 | 1496 | 1612 | 1729 | 1960 | 2160 | 2314 | 2411 | 2512 | 2642 | 2866 | 3036 | 3164 | 3352 | 3590 | 3836 | 3902 | 3999 | 4075 | 4100 | 4301 | 4316 | 4456 | 4554 | 4423 | 4525 | 4473 | 4537 | 4443 | 4324 | 4191 | 4163 | 4060 | 3929 | 3818 | 3679 | 3545 | 3383 | 3180 | 2969 | 2733 | 2560 | 2443 | 2289 | 2097 | 1911 | 1783 | 1630 | 1480 | 1376 | 1230 | 1107 | 1020 | 903 | 832 | 788 | 708 | 602 | 572 | 516 | 498 | 446 | 443 | 420 | 353 | 349 | 329 | 329 | 302 | 269 | 332 | 298 | 272 | 274 | 282 | 261 | 281 | 284 | 297 | 271 | 282 | 282 | 277 | 259 | 274 | 265 | 261 | 276 | 283 | 282 | 268 | 256 | 260 | 268 | 284 | 288 | 282 | 272 | 246 | 248 | 259 | 256 | 239 | 242 | 259 | 254 | 250 | 259 | 241 | 257 | 262 | 250 | 268 | 256 | 255 | 245 | 236 | 266 | 273 | 251 | 256 | 255 | 236 | 240 | 251 | 242 | 230 | 218 | 241 | 238 | 240 | 218 | 210 | 213 | 232 | 232 | 232 | 229 | 216 | 212 | 223 | 226 | 236 | 234 | 228 | 231 | 230 | 225 | 227 | 235 | 228 | 219 | 227 | 228 | 231 | 234 | 222 | 225 | 227 | 228 | 229 | 225 | 223 | 220 | 220 | 213 | 221 | 222 | 226 | 238 | 247 | 215 | 227 | 243 | 222 | 228 | 241 | 236 | 230 | 234 | 247 | 245 | 238 | 238 | 246 | 254 | 248 | 257 | 250 | 269 | 266 | 260 | 263 | 264 | 265 | 284 | 289 | 263 | 272 | 290 | 285 | 267 | 275 | 295 | 294 | 290 | 288 | 281 | 270 | 270 | 278 | 321 | 306 | 291 | 259 | 276 | 296 | 290 | 306 | 311 | 310 | 310 | 301 | 301 | 312 | 307 | 287 | 276 | 289 | 298 | 297 | 289 | 254 | 263 | 261 | 248 | 244 | 238 | 238 | 242 | 247 | 254 | 248 | 238 | 238 | 233 | 223 | 216 | 214 | 211 | 201 | 215 | 216 | 204 | 214 | 212 | 194 | 189 | 195 | 198 | 190 | 187 | 178 | 172 | 173 | 177 | 179 | 174 | 171 | 177 | 177 | 169 | 170 | 160 | 171 | 173 | 159 | 156 | 160 | 158 | 148 | 138 | 137 | 146 | 153 | 149 | 134 | 134 | 147 | 145 | 151 | 149 | 133 | 137 | 139 | 136 | 139 | 139 | 128 | 121 | 128 | 138 | 132 | 120 | 122 | 120 | 115 | 117 | 119 | 117 | 114 | 110 | 110 | 118 | 117 | 108 | 105 | 106 | 105 | 111 | 117 | 110 | 102 | 104 | 102 | 99 | 106 | 106 | 101 | 102 | 105 | 106 | 103 | 105 | 102 | 96 | 94 | 98 | 103 | 98 | 92 | 94 | 97 | 93 | 93 | 96 | 96 | 87 | 84 | 88 | 89 | 88 | 80 | 74 | 75 | 79 | 82 | 76 | 73 | 68 | 61 | 64 | 68 | 67 | 65 | 68 | 71 | 62 | 67 | 80 | 71 | 63 | 63 | 54 | 55 | 65 | 62 | 63 | 66 | 62 | 64 | 62 | 54 | 60 | 70 | 68 | 59 | 50 | 62 | 59 | 53 | 52 | 51 | 55 | 59 | 51 | 46 | 52 | 56 | 56 | 57 | 53 | 44 | 38 | 48 | 56 | 58 | 55 | 48 | 45 | 52 | 42 | 44 | 48 | 39 | 38 | 45 | 51 | 49 | 43 | 38 | 36 | 41 | 43 | 48 | 52 | 45 | 37 | 36 | 44 | 48 | 40 | 33 | 36 | 41 | 44 | 45 | 40 | 42 | 48 | 46 | 41 | 38 | 34 | 35 | 35 | 33 | 42 | 43 | 37 | 48 | 37 | 33 | 35 | 33 | 36 | 34 | 31 | 37 | 40 | 39 | 43 | 43 | 43 | 37 | 40 | 36 | 49 | 51 | 50 | 40 | 41 | 45 | 50 | 62 | 58 | 50 | 52 | 60 | 72 | 77 | 74 | 75 | 80 | 85 | 94 | 105 | 99 | 97 | 111 | 128 | 134 | 130 | 127 | 148 | 158 | 166 | 182 | 190 | 202 | 212 | 222 | 244 | 239 | 238 | 259 | 287 | 294 | 301 | 352 | 340 | 356 | 379 | 403 | 424 | 455 | 454 | 470 | 468 | 476 | 511 | 545 | 574 | 583 | 577 | 572 | 595 | 603 | 595 | 605 | 595 | 632 | 619 | 631 | 613 | 587 | 593 | 612 | 612 | 603 | 600 | 605 | 608 | 588 | 563 | 578 | 541 | 517 | 502 | 523 | 458 | 447 | 460 | 479 | 438 | 399 | 382 | 368 | 352 | 318 | 297 | 307 | 310 | 291 | 250 | 243 | 215 | 195 | 183 | 170 | 147 | 134 | 130 | 131 | 122 | 119 | 106 | 92 | 75 | 71 | 69 | 56 | 52 | 51 | 47 | 36 | 39 | 38 | 32 | 26 | 20 | 19 | 15 | 18 | 17 | 15 | 13 | 11 | 6 | 8 | 12 | 11 | 9 | 8 | 6 | 9 | 5 | 4 | 6 | 7 | 6 | 5 | 4 | 4 | 5 | 4 | 5 | 5 | 1 | 2 | 5 | 4 | 2 | 5 | 7 | 2 | 4 | 2 | 2 | 2 | 6 | 5 | 6 | 6 | 6 | 3 | 3 | 6 | 6 | 4 | 4 | 6 | 3 | 0 | 2 | 4 | 5 | 4 | 1 | 3 | 4 | 4 | 5 | 5 | 6 | 4 | 4 | 4 | 2 | 1 | 3 | 3 | 5 | 3 | 2 | 3 | 2 | 3 | 4 | 3 | 2 | 1 | 6 | 7 | 3 | 4 | 3 | 1 | 3 | 1 | 3 | 3 | 4 | 5 | 4 | 2 | 2 | 0 | 2 | 1 | 5 | 2 | 2 | 2 | 4 | 3 | 3 | 2 | 3 | 3 | 3 | 5 | 5 | 4 | 3 | 2 | 1 | 4 | 3 | 2 | 4 | 1 | 2 | 2 | 1 | 1 | 4 | 4 | 1 | 1 | 2 | 2 | 2 | 1 | 0 | 2 | 2 | 1 | 4 | 2 | 2 | 2 | 0 | 2 | 2 | 3 | 2 | 2 | 4 | 4 | 3 | 2 | 2 | 2 | 2 | 1 | 2 | 1 | 2 | 3 | 2 | 0 | 1 | 1 | 1 | 1 | 2 | 3 | 3 | 1 | 2 | 2 | 1 | 1 | 3 | 3 | 1 | 0 | 0 | 1 | 0 | 1 | 2 | 2 | 2 | 2 | 0 | 1 | 1 | 0 | 2 | 1 | 0 | 1 | 3 | 3 | 2 | 3 | 4 | 1 | 2 | 3 | 1 | 0 | 2 | 3 | 1 | 1 | 2 | 2 | 0 | 1 | 1 | 0 | 2 | 1 | 1 | 2 | 1 | 3 | 1 | 0 | 0 | 1 | 1 | 3 | 2 | 1 | 1 | 3 | 1 | 1 | 1 | 1 | 1 | 2 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 2 | 1 | 0 | 1 | 1 | 2 | 1 | 1 | 0 | 0 | 1 | 1 | 2 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 2 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 2 | 2 | 3 | 4 | 3 | 2 | 1 | 1 | 2 | 1 | 1 | 2 | 2 | 2 | 2 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 2 | 2 | 1 | 1 | 1 | 2 | 1 | 0 | 1 | 2 | 3 | 2 | 2 | 2 | 2 | 2 | 2 | 1 | 0 | 0 | 1 | 2 | 1 | 2 | 1 | 2 | 2 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 2 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 2 | 1 | 0 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 2 | 1 | 1 | 0 | 0 | 1 | 2 | 1 | 1 | 1 | 1 | 0 | 1 | 3 | 2 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 2 | 1 | 0 | 2 | 1 | 1 | 2 | 1 | 2 | 1 | 0 | 0 | 1 | 1 | 2 | 1 | 1 | 1 | 0 | 1 | 2 | 1 | 2 | 1 | 0 | 1 | 1 | 1 | 1 | 2 | 1 | 1 | 2 | 2 | 1 | 1 | 2 | 2 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 2 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 3 | 1 | 1 | 1 | 2 | 1 | 2 | 2 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 2 | 1 | 2 | 0 | 0 | 1 | 0 | 1 | 2 | 2 | 2 | 2 | 0 | 2 | 2 | 4 | 2 | 2 | 4 | 1 | 1 | 2 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 2 | 2 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 2 | 1 | 1 | 1 | 0 | 1 | 2 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 2 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 3 | 0 | 1 | 3 | 1 | 1 | 2 | 1 | 2 | 2 | 0 | 1 | 0 | 3 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 2 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 2 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 2 | 0 | 0 | 1 | 2 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0